## Slide 1
Prof. Geoff Smith & Yowwares Jeeraruangrattana
Leicester School of Pharmacy, De Montfort University,
United Kingdom
@LyoGroupDMU

## Slide 2: Outline
Through Vial Impedance Spectroscopy
  Description of measurement system
  Dielectric loss mechanisms
TVIS Applications
  Overview
  TVIS parameters for different facets of the freezing process (ice nucleation to solidification end point)
  Determination of in-vial glass transitions
  Determination of Collapse Phenomena
Acknowledgements
Yowwares Jeeraruangrattana
GPO Thailand
TVIS system connected on Lab Scale Freeze Dryer LyoBeta15, NIBSC, UK
TVIS pass through on
GEA Lyophil dryer, Hurth, Cologne

## Slide 3: Introduction to the TVIS System
Impedance spectroscopy characterizes the ability of materials to conduct electricity under an applied an oscillating voltage (of varying frequency)
Impedance measurements across a vial rather than within the vial
Hence “Through Vial Impedance Spectroscopy”
Features
  Single vial “non-product invasive”
  Both freezing and drying characterised in a single technique
  Non-perturbing to the packing of vials
  Stopper mechanism unaffected
TVIS

## Slide 4: TVIS Measurement System
The designs of various vials that have been modified with copper foil electrodes (10 mm in height and 3 mm from the base of each container
  20 mm crimp-neck vial with 10 ml nominal capacity
  20 mm crimp-neck vial with 5 ml nominal capacity
  screw-neck vial with 5 ml nominal capacity
The different styles of a bespoke pass-through for TVIS systems
  Connected via the manifold hose on the outside of the dryer
  Connected via the port on top left side of the door on the dryer
  Connected to a port on the top of the drying chamber
Adelphi
VC010-20C
Adelphi
VC005-20C
Adelphi
VCD005
A
B
C
Notes: Fig. 2. Examples of pass-throughs for the TVIS instrument that have been deployed on a range of freeze-dryers at various institutions and companies:
(A) Virtis Advantage Plus XL freeze-dryer (De Montfort University) in which the pass-through is connected via the manifold hose on the outside of the dryer;
(B) Telstar LyoBeta freeze-dryer (Image courtesy of the National Institute of Biological Standards and Control) in which the pass-through is connected via the port on top left side of the door on the dryer;
(C) GEA LYOVAC™ FCM2 LN2 pilot scale dryer (Image courtesy of GEA Lyophil GmbH) where the pass-through is connected to a port on the top of the drying chamber; Each pass-through has a minimum of six TE-connectivity-AMP SMA-jack to SMA jack bulk-head adapters (1054874-1): One adapter is used for the stimulating signal and the other five for the receiving signal. All pass-throughs were manufactured by GEA Pharma Systems (Eastleigh, UK).

## Slide 5
Through Vial Impedance Spectroscopy (TVIS)
Description of Measurement System

## Slide 6
Freeze drying chamber
Stimulating voltage
Resultant current
LyoDEATM measurement software
Junction
box
TVIS system
(I to V convertor)
Pass-through
TVIS
measurement vial
LyoViewTM analysis software

## Slide 7
TVIS Applications
Overview

## Slide 8: Through Vial Impedance Spectroscopy (TVIS)
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034
Time 0.3 h | 0.27396 | 0.24912 | 0.22663 | 0.2063 | 0.18797 | 0.17147 | 0.15668 | 0.14349 | 0.13183 | 0.12162 | 0.11284 | 0.10548 | 0.099574 | 0.095186 | 0.092431 | 0.091474 | 0.092551 | 0.095973 | 0.10215 | 0.11162 | 0.12501 | 0.14319 | 0.16715 | 0.198 | 0.23699 | 0.28544 | 0.34395 | 0.41214 | 0.48865 | 0.56683 | 0.63782 | 0.69048 | 0.71191 | 0.69725 | 0.65001 | 0.57934 | 0.49898 | 0.41726 | 0.34526 | 0.28116 | 0.22886 | 0.17991 | 0.14643 | 0.1181 | 0.089415 | 0.072054
Time 4 h | 0.099081 | 0.096153 | 0.095569 | 0.097591 | 0.10258 | 0.11101 | 0.12349 | 0.14075 | 0.16366 | 0.19315 | 0.23007 | 0.27492 | 0.32735 | 0.38533 | 0.44432 | 0.49685 | 0.53348 | 0.54588 | 0.53067 | 0.49125 | 0.43601 | 0.37424 | 0.31336 | 0.25794 | 0.20992 | 0.16948 | 0.13623 | 0.10926 | 0.087318 | 0.069838 | 0.055881 | 0.044566 | 0.035658 | 0.0285 | 0.02283 | 0.018246 | 0.014616 | 0.011669 | 0.0094101 | 0.0075633 | 0.0061328 | 0.0048349
Time 0.3 h | 0.26705 | 0.24478 | 0.21978 | 0.1979 | 0.17863 | 0.15615 | 0.14276 | 0.12639 | 0.11756 | 0.11024 | 0.10127 | 0.1004 | 0.098391 | 0.10089 | 0.10904 | 0.11009 | 0.11451 | 0.12496 | 0.13439 | 0.14683 | 0.16234 | 0.18119 | 0.20238 | 0.22944 | 0.26256 | 0.30331 | 0.35113 | 0.4089 | 0.47307 | 0.54005 | 0.60366 | 0.65431 | 0.67974 | 0.67269 | 0.63391 | 0.56955 | 0.49367 | 0.41413 | 0.34323 | 0.27998 | 0.22777
Time 4 h | 0.093429 | 0.092843 | 0.096342 | 0.10075 | 0.11232 | 0.12322 | 0.1435 | 0.1596 | 0.19684 | 0.23027 | 0.27214 | 0.31647 | 0.36246 | 0.41025 | 0.45421 | 0.48896 | 0.51147 | 0.51695 | 0.50285 | 0.46926 | 0.42098 | 0.36431 | 0.30625 | 0.25218 | 0.20511 | 0.16565 | 0.13355 | 0.10767 | 0.086605 | 0.069689 | 0.056026 | 0.045028 | 0.036377 | 0.029122 | 0.023205 | 0.018299 | 0.014688 | 0.011778 | 0.0095536 | 0.0077519 | 0.0063112
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034
Time 0.3 h | 0.27396 | 0.24912 | 0.22663 | 0.2063 | 0.18797 | 0.17147 | 0.15668 | 0.14349 | 0.13183 | 0.12162 | 0.11284 | 0.10548 | 0.099574 | 0.095186 | 0.092431 | 0.091474 | 0.092551 | 0.095973 | 0.10215 | 0.11162 | 0.12501 | 0.14319 | 0.16715 | 0.198 | 0.23699 | 0.28544 | 0.34395 | 0.41214 | 0.48865 | 0.56683 | 0.63782 | 0.69048 | 0.71191 | 0.69725 | 0.65001 | 0.57934 | 0.49898 | 0.41726 | 0.34526 | 0.28116 | 0.22886 | 0.17991 | 0.14643 | 0.1181 | 0.089415 | 0.072054
Time 4 h | 0.099081 | 0.096153 | 0.095569 | 0.097591 | 0.10258 | 0.11101 | 0.12349 | 0.14075 | 0.16366 | 0.19315 | 0.23007 | 0.27492 | 0.32735 | 0.38533 | 0.44432 | 0.49685 | 0.53348 | 0.54588 | 0.53067 | 0.49125 | 0.43601 | 0.37424 | 0.31336 | 0.25794 | 0.20992 | 0.16948 | 0.13623 | 0.10926 | 0.087318 | 0.069838 | 0.055881 | 0.044566 | 0.035658 | 0.0285 | 0.02283 | 0.018246 | 0.014616 | 0.011669 | 0.0094101 | 0.0075633 | 0.0061328 | 0.0048349
Time 0.3 h | 0.26705 | 0.24478 | 0.21978 | 0.1979 | 0.17863 | 0.15615 | 0.14276 | 0.12639 | 0.11756 | 0.11024 | 0.10127 | 0.1004 | 0.098391 | 0.10089 | 0.10904 | 0.11009 | 0.11451 | 0.12496 | 0.13439 | 0.14683 | 0.16234 | 0.18119 | 0.20238 | 0.22944 | 0.26256 | 0.30331 | 0.35113 | 0.4089 | 0.47307 | 0.54005 | 0.60366 | 0.65431 | 0.67974 | 0.67269 | 0.63391 | 0.56955 | 0.49367 | 0.41413 | 0.34323 | 0.27998 | 0.22777
Time 4 h | 0.093429 | 0.092843 | 0.096342 | 0.10075 | 0.11232 | 0.12322 | 0.1435 | 0.1596 | 0.19684 | 0.23027 | 0.27214 | 0.31647 | 0.36246 | 0.41025 | 0.45421 | 0.48896 | 0.51147 | 0.51695 | 0.50285 | 0.46926 | 0.42098 | 0.36431 | 0.30625 | 0.25218 | 0.20511 | 0.16565 | 0.13355 | 0.10767 | 0.086605 | 0.069689 | 0.056026 | 0.045028 | 0.036377 | 0.029122 | 0.023205 | 0.018299 | 0.014688 | 0.011778 | 0.0095536 | 0.0077519 | 0.0063112
FPEAK
FPEAK
Liquid State
Solid State
Monitoring Phase Behaviour (ice nucleation temperature and solidification end points by using
Notes: Eg temperature cycling below to calibrate Fpeak for temperature (in order to predict product temperatures in primary drying) and the compensate cpeak for temperature so that it can used to monitor ice content and hence drying rate
Eg 2 left graph: monitoring phase behaviour (ice nucleation temperatures and solidufication end points)

## Slide 9: Through Vial Impedance Spectroscopy (TVIS)
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034
Time 0.3 h | 0.27396 | 0.24912 | 0.22663 | 0.2063 | 0.18797 | 0.17147 | 0.15668 | 0.14349 | 0.13183 | 0.12162 | 0.11284 | 0.10548 | 0.099574 | 0.095186 | 0.092431 | 0.091474 | 0.092551 | 0.095973 | 0.10215 | 0.11162 | 0.12501 | 0.14319 | 0.16715 | 0.198 | 0.23699 | 0.28544 | 0.34395 | 0.41214 | 0.48865 | 0.56683 | 0.63782 | 0.69048 | 0.71191 | 0.69725 | 0.65001 | 0.57934 | 0.49898 | 0.41726 | 0.34526 | 0.28116 | 0.22886 | 0.17991 | 0.14643 | 0.1181 | 0.089415 | 0.072054
Time 4 h | 0.099081 | 0.096153 | 0.095569 | 0.097591 | 0.10258 | 0.11101 | 0.12349 | 0.14075 | 0.16366 | 0.19315 | 0.23007 | 0.27492 | 0.32735 | 0.38533 | 0.44432 | 0.49685 | 0.53348 | 0.54588 | 0.53067 | 0.49125 | 0.43601 | 0.37424 | 0.31336 | 0.25794 | 0.20992 | 0.16948 | 0.13623 | 0.10926 | 0.087318 | 0.069838 | 0.055881 | 0.044566 | 0.035658 | 0.0285 | 0.02283 | 0.018246 | 0.014616 | 0.011669 | 0.0094101 | 0.0075633 | 0.0061328 | 0.0048349
Time 0.3 h | 0.26705 | 0.24478 | 0.21978 | 0.1979 | 0.17863 | 0.15615 | 0.14276 | 0.12639 | 0.11756 | 0.11024 | 0.10127 | 0.1004 | 0.098391 | 0.10089 | 0.10904 | 0.11009 | 0.11451 | 0.12496 | 0.13439 | 0.14683 | 0.16234 | 0.18119 | 0.20238 | 0.22944 | 0.26256 | 0.30331 | 0.35113 | 0.4089 | 0.47307 | 0.54005 | 0.60366 | 0.65431 | 0.67974 | 0.67269 | 0.63391 | 0.56955 | 0.49367 | 0.41413 | 0.34323 | 0.27998 | 0.22777
Time 4 h | 0.093429 | 0.092843 | 0.096342 | 0.10075 | 0.11232 | 0.12322 | 0.1435 | 0.1596 | 0.19684 | 0.23027 | 0.27214 | 0.31647 | 0.36246 | 0.41025 | 0.45421 | 0.48896 | 0.51147 | 0.51695 | 0.50285 | 0.46926 | 0.42098 | 0.36431 | 0.30625 | 0.25218 | 0.20511 | 0.16565 | 0.13355 | 0.10767 | 0.086605 | 0.069689 | 0.056026 | 0.045028 | 0.036377 | 0.029122 | 0.023205 | 0.018299 | 0.014688 | 0.011778 | 0.0095536 | 0.0077519 | 0.0063112
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034
Time 0.3 h | 0.27396 | 0.24912 | 0.22663 | 0.2063 | 0.18797 | 0.17147 | 0.15668 | 0.14349 | 0.13183 | 0.12162 | 0.11284 | 0.10548 | 0.099574 | 0.095186 | 0.092431 | 0.091474 | 0.092551 | 0.095973 | 0.10215 | 0.11162 | 0.12501 | 0.14319 | 0.16715 | 0.198 | 0.23699 | 0.28544 | 0.34395 | 0.41214 | 0.48865 | 0.56683 | 0.63782 | 0.69048 | 0.71191 | 0.69725 | 0.65001 | 0.57934 | 0.49898 | 0.41726 | 0.34526 | 0.28116 | 0.22886 | 0.17991 | 0.14643 | 0.1181 | 0.089415 | 0.072054
Time 4 h | 0.099081 | 0.096153 | 0.095569 | 0.097591 | 0.10258 | 0.11101 | 0.12349 | 0.14075 | 0.16366 | 0.19315 | 0.23007 | 0.27492 | 0.32735 | 0.38533 | 0.44432 | 0.49685 | 0.53348 | 0.54588 | 0.53067 | 0.49125 | 0.43601 | 0.37424 | 0.31336 | 0.25794 | 0.20992 | 0.16948 | 0.13623 | 0.10926 | 0.087318 | 0.069838 | 0.055881 | 0.044566 | 0.035658 | 0.0285 | 0.02283 | 0.018246 | 0.014616 | 0.011669 | 0.0094101 | 0.0075633 | 0.0061328 | 0.0048349
Time 0.3 h | 0.26705 | 0.24478 | 0.21978 | 0.1979 | 0.17863 | 0.15615 | 0.14276 | 0.12639 | 0.11756 | 0.11024 | 0.10127 | 0.1004 | 0.098391 | 0.10089 | 0.10904 | 0.11009 | 0.11451 | 0.12496 | 0.13439 | 0.14683 | 0.16234 | 0.18119 | 0.20238 | 0.22944 | 0.26256 | 0.30331 | 0.35113 | 0.4089 | 0.47307 | 0.54005 | 0.60366 | 0.65431 | 0.67974 | 0.67269 | 0.63391 | 0.56955 | 0.49367 | 0.41413 | 0.34323 | 0.27998 | 0.22777
Time 4 h | 0.093429 | 0.092843 | 0.096342 | 0.10075 | 0.11232 | 0.12322 | 0.1435 | 0.1596 | 0.19684 | 0.23027 | 0.27214 | 0.31647 | 0.36246 | 0.41025 | 0.45421 | 0.48896 | 0.51147 | 0.51695 | 0.50285 | 0.46926 | 0.42098 | 0.36431 | 0.30625 | 0.25218 | 0.20511 | 0.16565 | 0.13355 | 0.10767 | 0.086605 | 0.069689 | 0.056026 | 0.045028 | 0.036377 | 0.029122 | 0.023205 | 0.018299 | 0.014688 | 0.011778 | 0.0095536 | 0.0077519 | 0.0063112
FPEAK
FPEAK
Liquid State
Solid State
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 4.4 h | 0.003361 | 0.0040465 | 0.0025946 | 0.0026214 | 0.0044553 | 0.0026323 | 0.0029087 | 0.0037366 | 0.0047149 | 0.0057564 | 0.0073092 | 0.0089323 | 0.011077 | 0.013703 | 0.017072 | 0.021236 | 0.026498 | 0.03301 | 0.041177 | 0.051293 | 0.064055 | 0.079683 | 0.099225 | 0.12384 | 0.15437 | 0.19231 | 0.23844 | 0.29224 | 0.35137 | 0.41125 | 0.46468 | 0.50437 | 0.52462 | 0.52356 | 0.50251 | 0.46708 | 0.42347 | 0.37253 | 0.32299 | 0.27537 | 0.23266 | 0.1978 | 0.16633 | 0.14424 | 0.12252 | 0.10889 | 0.097697 | 0.093226 | 0.09166 | 0.092155 | 0.093069
Time 4.5 h | 0.0034249 | 0.0041377 | 0.0027283 | 0.0027814 | 0.0046887 | 0.0029259 | 0.0032733 | 0.0042363 | 0.0053358 | 0.0065296 | 0.0082746 | 0.010141 | 0.012601 | 0.015587 | 0.01947 | 0.024233 | 0.030264 | 0.037737 | 0.047171 | 0.058901 | 0.073903 | 0.092482 | 0.11581 | 0.14538 | 0.18208 | 0.2274 | 0.28124 | 0.3416 | 0.40418 | 0.46243 | 0.50811 | 0.53414 | 0.53731 | 0.51809 | 0.48197 | 0.43577 | 0.38293 | 0.3311 | 0.28177 | 0.23748 | 0.20077 | 0.16802 | 0.14416 | 0.11848 | 0.11171 | 0.10187 | 0.09506 | 0.092791 | 0.094429 | 0.097904 | 0.10371
Time 4.6 h | 0.0036091 | 0.0043103 | 0.0029276 | 0.0030069 | 0.0049863 | 0.0032919 | 0.0037411 | 0.0048548 | 0.006126 | 0.0075084 | 0.0095269 | 0.011725 | 0.014579 | 0.018062 | 0.02259 | 0.028182 | 0.035264 | 0.044078 | 0.055262 | 0.069282 | 0.08729 | 0.10998 | 0.13862 | 0.17503 | 0.21981 | 0.27381 | 0.33549 | 0.40061 | 0.46249 | 0.51282 | 0.5432 | 0.54939 | 0.53148 | 0.49439 | 0.44517 | 0.39079 | 0.33577 | 0.28667 | 0.23998 | 0.20372 | 0.17181 | 0.14592 | 0.12721 | 0.11818 | 0.10489 | 0.098824 | 0.097908 | 0.099283 | 0.10394 | 0.11202 | 0.12205
Time 4.7 h | 0.0040254 | 0.0023197 | 0.0060162 | 0.0032852 | 0.0053707 | 0.003773 | 0.0043423 | 0.005672 | 0.0071637 | 0.0088006 | 0.011181 | 0.013771 | 0.017191 | 0.021329 | 0.026752 | 0.033478 | 0.04203 | 0.052722 | 0.066401 | 0.083738 | 0.10613 | 0.13436 | 0.17037 | 0.21573 | 0.27035 | 0.33347 | 0.40107 | 0.46603 | 0.51937 | 0.55253 | 0.55996 | 0.54171 | 0.50311 | 0.45097 | 0.39455 | 0.33876 | 0.28887 | 0.24104 | 0.20387 | 0.17339 | 0.14886 | 0.13225 | 0.11753 | 0.11 | 0.10482 | 0.10438 | 0.1074 | 0.1138 | 0.12408 | 0.13639 | 0.14993
Time 4.8 h | 0.0039527 | 0.0047899 | 0.0034912 | 0.0036734 | 0.0059298 | 0.004463 | 0.0052209 | 0.0068624 | 0.0086692 | 0.010687 | 0.013626 | 0.016821 | 0.021041 | 0.026213 | 0.03302 | 0.041432 | 0.052261 | 0.065911 | 0.083539 | 0.10606 | 0.13545 | 0.17252 | 0.21928 | 0.27665 | 0.34237 | 0.41271 | 0.47977 | 0.53353 | 0.56486 | 0.56874 | 0.5461 | 0.50317 | 0.44813 | 0.38872 | 0.33224 | 0.28045 | 0.23905 | 0.20172 | 0.17254 | 0.14998 | 0.13098 | 0.12232 | 0.11386 | 0.11155 | 0.11412 | 0.11812 | 0.12658 | 0.13848 | 0.15361 | 0.17099 | 0.18866
Time 4.9 h | 0.0043188 | 0.0052277 | 0.0039868 | 0.0042288 | 0.0066529 | 0.0053479 | 0.0063485 | 0.0083796 | 0.010604 | 0.013107 | 0.016765 | 0.020805 | 0.02609 | 0.032614 | 0.041224 | 0.052023 | 0.065957 | 0.083695 | 0.10683 | 0.13651 | 0.17523 | 0.22353 | 0.28261 | 0.35128 | 0.42376 | 0.49222 | 0.54583 | 0.5749 | 0.57491 | 0.54814 | 0.50162 | 0.4443 | 0.38384 | 0.32634 | 0.27564 | 0.23272 | 0.19908 | 0.1717 | 0.1495 | 0.13558 | 0.12502 | 0.11938 | 0.11838 | 0.12437 | 0.13062 | 0.13992 | 0.15492 | 0.1723 | 0.19073 | 0.21306 | 0.23377
Time 5 h | 0.0049133 | 0.0058042 | 0.0046338 | 0.0049712 | 0.0076664 | 0.0066031 | 0.0079479 | 0.010533 | 0.013358 | 0.01657 | 0.021298 | 0.026516 | 0.033397 | 0.041936 | 0.053268 | 0.067599 | 0.086277 | 0.11024 | 0.14161 | 0.18183 | 0.2334 | 0.29549 | 0.36662 | 0.44167 | 0.50969 | 0.5605 | 0.58424 | 0.57776 | 0.54458 | 0.49349 | 0.43374 | 0.3731 | 0.31662 | 0.26729 | 0.22557 | 0.19336 | 0.16657 | 0.14984 | 0.13717 | 0.1291 | 0.12498 | 0.12759 | 0.13258 | 0.1444 | 0.15601 | 0.17194 | 0.19151 | 0.21398 | 0.23437 | 0.26066 | 0.28203
Time 5.1 h | 0.0051615 | 0.0065394 | 0.0054847 | 0.0059393 | 0.0090001 | 0.0082333 | 0.010001 | 0.013324 | 0.016914 | 0.021051 | 0.027147 | 0.03396 | 0.042946 | 0.054219 | 0.069319 | 0.088593 | 0.11394 | 0.14657 | 0.18909 | 0.24251 | 0.30812 | 0.38167 | 0.45767 | 0.52576 | 0.57299 | 0.59132 | 0.57824 | 0.5397 | 0.48481 | 0.42344 | 0.36292 | 0.3071 | 0.25963 | 0.22024 | 0.18846 | 0.16573 | 0.14947 | 0.13846 | 0.13032 | 0.13113 | 0.1334 | 0.1422 | 0.15348 | 0.16641 | 0.18887 | 0.2096 | 0.23238 | 0.25699 | 0.28353 | 0.3063 | 0.32649
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 4.4 h | 0.003361 | 0.0040465 | 0.0025946 | 0.0026214 | 0.0044553 | 0.0026323 | 0.0029087 | 0.0037366 | 0.0047149 | 0.0057564 | 0.0073092 | 0.0089323 | 0.011077 | 0.013703 | 0.017072 | 0.021236 | 0.026498 | 0.03301 | 0.041177 | 0.051293 | 0.064055 | 0.079683 | 0.099225 | 0.12384 | 0.15437 | 0.19231 | 0.23844 | 0.29224 | 0.35137 | 0.41125 | 0.46468 | 0.50437 | 0.52462 | 0.52356 | 0.50251 | 0.46708 | 0.42347 | 0.37253 | 0.32299 | 0.27537 | 0.23266 | 0.1978 | 0.16633 | 0.14424 | 0.12252 | 0.10889 | 0.097697 | 0.093226 | 0.09166 | 0.092155 | 0.093069
Time 4.5 h | 0.0034249 | 0.0041377 | 0.0027283 | 0.0027814 | 0.0046887 | 0.0029259 | 0.0032733 | 0.0042363 | 0.0053358 | 0.0065296 | 0.0082746 | 0.010141 | 0.012601 | 0.015587 | 0.01947 | 0.024233 | 0.030264 | 0.037737 | 0.047171 | 0.058901 | 0.073903 | 0.092482 | 0.11581 | 0.14538 | 0.18208 | 0.2274 | 0.28124 | 0.3416 | 0.40418 | 0.46243 | 0.50811 | 0.53414 | 0.53731 | 0.51809 | 0.48197 | 0.43577 | 0.38293 | 0.3311 | 0.28177 | 0.23748 | 0.20077 | 0.16802 | 0.14416 | 0.11848 | 0.11171 | 0.10187 | 0.09506 | 0.092791 | 0.094429 | 0.097904 | 0.10371
Time 4.6 h | 0.0036091 | 0.0043103 | 0.0029276 | 0.0030069 | 0.0049863 | 0.0032919 | 0.0037411 | 0.0048548 | 0.006126 | 0.0075084 | 0.0095269 | 0.011725 | 0.014579 | 0.018062 | 0.02259 | 0.028182 | 0.035264 | 0.044078 | 0.055262 | 0.069282 | 0.08729 | 0.10998 | 0.13862 | 0.17503 | 0.21981 | 0.27381 | 0.33549 | 0.40061 | 0.46249 | 0.51282 | 0.5432 | 0.54939 | 0.53148 | 0.49439 | 0.44517 | 0.39079 | 0.33577 | 0.28667 | 0.23998 | 0.20372 | 0.17181 | 0.14592 | 0.12721 | 0.11818 | 0.10489 | 0.098824 | 0.097908 | 0.099283 | 0.10394 | 0.11202 | 0.12205
Time 4.7 h | 0.0040254 | 0.0023197 | 0.0060162 | 0.0032852 | 0.0053707 | 0.003773 | 0.0043423 | 0.005672 | 0.0071637 | 0.0088006 | 0.011181 | 0.013771 | 0.017191 | 0.021329 | 0.026752 | 0.033478 | 0.04203 | 0.052722 | 0.066401 | 0.083738 | 0.10613 | 0.13436 | 0.17037 | 0.21573 | 0.27035 | 0.33347 | 0.40107 | 0.46603 | 0.51937 | 0.55253 | 0.55996 | 0.54171 | 0.50311 | 0.45097 | 0.39455 | 0.33876 | 0.28887 | 0.24104 | 0.20387 | 0.17339 | 0.14886 | 0.13225 | 0.11753 | 0.11 | 0.10482 | 0.10438 | 0.1074 | 0.1138 | 0.12408 | 0.13639 | 0.14993
Time 4.8 h | 0.0039527 | 0.0047899 | 0.0034912 | 0.0036734 | 0.0059298 | 0.004463 | 0.0052209 | 0.0068624 | 0.0086692 | 0.010687 | 0.013626 | 0.016821 | 0.021041 | 0.026213 | 0.03302 | 0.041432 | 0.052261 | 0.065911 | 0.083539 | 0.10606 | 0.13545 | 0.17252 | 0.21928 | 0.27665 | 0.34237 | 0.41271 | 0.47977 | 0.53353 | 0.56486 | 0.56874 | 0.5461 | 0.50317 | 0.44813 | 0.38872 | 0.33224 | 0.28045 | 0.23905 | 0.20172 | 0.17254 | 0.14998 | 0.13098 | 0.12232 | 0.11386 | 0.11155 | 0.11412 | 0.11812 | 0.12658 | 0.13848 | 0.15361 | 0.17099 | 0.18866
Time 4.9 h | 0.0043188 | 0.0052277 | 0.0039868 | 0.0042288 | 0.0066529 | 0.0053479 | 0.0063485 | 0.0083796 | 0.010604 | 0.013107 | 0.016765 | 0.020805 | 0.02609 | 0.032614 | 0.041224 | 0.052023 | 0.065957 | 0.083695 | 0.10683 | 0.13651 | 0.17523 | 0.22353 | 0.28261 | 0.35128 | 0.42376 | 0.49222 | 0.54583 | 0.5749 | 0.57491 | 0.54814 | 0.50162 | 0.4443 | 0.38384 | 0.32634 | 0.27564 | 0.23272 | 0.19908 | 0.1717 | 0.1495 | 0.13558 | 0.12502 | 0.11938 | 0.11838 | 0.12437 | 0.13062 | 0.13992 | 0.15492 | 0.1723 | 0.19073 | 0.21306 | 0.23377
Time 5 h | 0.0049133 | 0.0058042 | 0.0046338 | 0.0049712 | 0.0076664 | 0.0066031 | 0.0079479 | 0.010533 | 0.013358 | 0.01657 | 0.021298 | 0.026516 | 0.033397 | 0.041936 | 0.053268 | 0.067599 | 0.086277 | 0.11024 | 0.14161 | 0.18183 | 0.2334 | 0.29549 | 0.36662 | 0.44167 | 0.50969 | 0.5605 | 0.58424 | 0.57776 | 0.54458 | 0.49349 | 0.43374 | 0.3731 | 0.31662 | 0.26729 | 0.22557 | 0.19336 | 0.16657 | 0.14984 | 0.13717 | 0.1291 | 0.12498 | 0.12759 | 0.13258 | 0.1444 | 0.15601 | 0.17194 | 0.19151 | 0.21398 | 0.23437 | 0.26066 | 0.28203
Time 5.1 h | 0.0051615 | 0.0065394 | 0.0054847 | 0.0059393 | 0.0090001 | 0.0082333 | 0.010001 | 0.013324 | 0.016914 | 0.021051 | 0.027147 | 0.03396 | 0.042946 | 0.054219 | 0.069319 | 0.088593 | 0.11394 | 0.14657 | 0.18909 | 0.24251 | 0.30812 | 0.38167 | 0.45767 | 0.52576 | 0.57299 | 0.59132 | 0.57824 | 0.5397 | 0.48481 | 0.42344 | 0.36292 | 0.3071 | 0.25963 | 0.22024 | 0.18846 | 0.16573 | 0.14947 | 0.13846 | 0.13032 | 0.13113 | 0.1334 | 0.1422 | 0.15348 | 0.16641 | 0.18887 | 0.2096 | 0.23238 | 0.25699 | 0.28353 | 0.3063 | 0.32649
-38 oC
-18 oC
Monitoring Phase Behaviour (ice nucleation temperature and solidification end points by using
temperature calibration for predicting temperature of the product in primary drying
Notes: Eg temperature cycling below to calibrate Fpeak for temperature (in order to predict product temperatures in primary drying) and the compensate cpeak for temperature so that it can used to monitor ice content and hence drying rate
Eg 2 left graph: monitoring phase behaviour (ice nucleation temperatures and solidufication end points)

## Slide 10: Through Vial Impedance Spectroscopy (TVIS)
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034
Time 0.3 h | 0.27396 | 0.24912 | 0.22663 | 0.2063 | 0.18797 | 0.17147 | 0.15668 | 0.14349 | 0.13183 | 0.12162 | 0.11284 | 0.10548 | 0.099574 | 0.095186 | 0.092431 | 0.091474 | 0.092551 | 0.095973 | 0.10215 | 0.11162 | 0.12501 | 0.14319 | 0.16715 | 0.198 | 0.23699 | 0.28544 | 0.34395 | 0.41214 | 0.48865 | 0.56683 | 0.63782 | 0.69048 | 0.71191 | 0.69725 | 0.65001 | 0.57934 | 0.49898 | 0.41726 | 0.34526 | 0.28116 | 0.22886 | 0.17991 | 0.14643 | 0.1181 | 0.089415 | 0.072054
Time 4 h | 0.099081 | 0.096153 | 0.095569 | 0.097591 | 0.10258 | 0.11101 | 0.12349 | 0.14075 | 0.16366 | 0.19315 | 0.23007 | 0.27492 | 0.32735 | 0.38533 | 0.44432 | 0.49685 | 0.53348 | 0.54588 | 0.53067 | 0.49125 | 0.43601 | 0.37424 | 0.31336 | 0.25794 | 0.20992 | 0.16948 | 0.13623 | 0.10926 | 0.087318 | 0.069838 | 0.055881 | 0.044566 | 0.035658 | 0.0285 | 0.02283 | 0.018246 | 0.014616 | 0.011669 | 0.0094101 | 0.0075633 | 0.0061328 | 0.0048349
Time 0.3 h | 0.26705 | 0.24478 | 0.21978 | 0.1979 | 0.17863 | 0.15615 | 0.14276 | 0.12639 | 0.11756 | 0.11024 | 0.10127 | 0.1004 | 0.098391 | 0.10089 | 0.10904 | 0.11009 | 0.11451 | 0.12496 | 0.13439 | 0.14683 | 0.16234 | 0.18119 | 0.20238 | 0.22944 | 0.26256 | 0.30331 | 0.35113 | 0.4089 | 0.47307 | 0.54005 | 0.60366 | 0.65431 | 0.67974 | 0.67269 | 0.63391 | 0.56955 | 0.49367 | 0.41413 | 0.34323 | 0.27998 | 0.22777
Time 4 h | 0.093429 | 0.092843 | 0.096342 | 0.10075 | 0.11232 | 0.12322 | 0.1435 | 0.1596 | 0.19684 | 0.23027 | 0.27214 | 0.31647 | 0.36246 | 0.41025 | 0.45421 | 0.48896 | 0.51147 | 0.51695 | 0.50285 | 0.46926 | 0.42098 | 0.36431 | 0.30625 | 0.25218 | 0.20511 | 0.16565 | 0.13355 | 0.10767 | 0.086605 | 0.069689 | 0.056026 | 0.045028 | 0.036377 | 0.029122 | 0.023205 | 0.018299 | 0.014688 | 0.011778 | 0.0095536 | 0.0077519 | 0.0063112
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034
Time 0.3 h | 0.27396 | 0.24912 | 0.22663 | 0.2063 | 0.18797 | 0.17147 | 0.15668 | 0.14349 | 0.13183 | 0.12162 | 0.11284 | 0.10548 | 0.099574 | 0.095186 | 0.092431 | 0.091474 | 0.092551 | 0.095973 | 0.10215 | 0.11162 | 0.12501 | 0.14319 | 0.16715 | 0.198 | 0.23699 | 0.28544 | 0.34395 | 0.41214 | 0.48865 | 0.56683 | 0.63782 | 0.69048 | 0.71191 | 0.69725 | 0.65001 | 0.57934 | 0.49898 | 0.41726 | 0.34526 | 0.28116 | 0.22886 | 0.17991 | 0.14643 | 0.1181 | 0.089415 | 0.072054
Time 4 h | 0.099081 | 0.096153 | 0.095569 | 0.097591 | 0.10258 | 0.11101 | 0.12349 | 0.14075 | 0.16366 | 0.19315 | 0.23007 | 0.27492 | 0.32735 | 0.38533 | 0.44432 | 0.49685 | 0.53348 | 0.54588 | 0.53067 | 0.49125 | 0.43601 | 0.37424 | 0.31336 | 0.25794 | 0.20992 | 0.16948 | 0.13623 | 0.10926 | 0.087318 | 0.069838 | 0.055881 | 0.044566 | 0.035658 | 0.0285 | 0.02283 | 0.018246 | 0.014616 | 0.011669 | 0.0094101 | 0.0075633 | 0.0061328 | 0.0048349
Time 0.3 h | 0.26705 | 0.24478 | 0.21978 | 0.1979 | 0.17863 | 0.15615 | 0.14276 | 0.12639 | 0.11756 | 0.11024 | 0.10127 | 0.1004 | 0.098391 | 0.10089 | 0.10904 | 0.11009 | 0.11451 | 0.12496 | 0.13439 | 0.14683 | 0.16234 | 0.18119 | 0.20238 | 0.22944 | 0.26256 | 0.30331 | 0.35113 | 0.4089 | 0.47307 | 0.54005 | 0.60366 | 0.65431 | 0.67974 | 0.67269 | 0.63391 | 0.56955 | 0.49367 | 0.41413 | 0.34323 | 0.27998 | 0.22777
Time 4 h | 0.093429 | 0.092843 | 0.096342 | 0.10075 | 0.11232 | 0.12322 | 0.1435 | 0.1596 | 0.19684 | 0.23027 | 0.27214 | 0.31647 | 0.36246 | 0.41025 | 0.45421 | 0.48896 | 0.51147 | 0.51695 | 0.50285 | 0.46926 | 0.42098 | 0.36431 | 0.30625 | 0.25218 | 0.20511 | 0.16565 | 0.13355 | 0.10767 | 0.086605 | 0.069689 | 0.056026 | 0.045028 | 0.036377 | 0.029122 | 0.023205 | 0.018299 | 0.014688 | 0.011778 | 0.0095536 | 0.0077519 | 0.0063112
FPEAK
FPEAK
Liquid State
Solid State
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625 | 5.7569630370485534 | 5.8239109124113 | 5.9030899869919438 | 6
Time 13.1 h | 0.10929 | 0.1004 | 0.092502 | 0.087998 | 0.086049 | 0.088115 | 0.0938 | 0.10112 | 0.11429 | 0.13106 | 0.15364 | 0.18166 | 0.21583 | 0.25512 | 0.29985 | 0.34774 | 0.3951 | 0.43795 | 0.46981 | 0.48496 | 0.48002 | 0.45469 | 0.41262 | 0.36037 | 0.30465 | 0.25101 | 0.20331 | 0.16311 | 0.13007 | 0.10386 | 0.083113 | 0.06641 | 0.053298 | 0.042748 | 0.034381 | 0.027604 | 0.022219 | 0.017839 | 0.014449 | 0.011634 | 0.0093969 | 0.0073773 | 0.0059706 | 0.0046834 | 0.003604 | 0.003119 | 0.0043457 | 0.0022557 | 0.0032184 | 0.0030577 | 0.0023234
Time 16.1 h | 0.11007 | 0.10093 | 0.090991 | 0.082047 | 0.079543 | 0.078788 | 0.081723 | 0.087173 | 0.096992 | 0.11005 | 0.12829 | 0.15003 | 0.17689 | 0.20803 | 0.24102 | 0.27482 | 0.30518 | 0.32877 | 0.34107 | 0.33971 | 0.32434 | 0.2969 | 0.26113 | 0.22177 | 0.1829 | 0.1475 | 0.11733 | 0.092752 | 0.073068 | 0.057802 | 0.045912 | 0.036469 | 0.029123 | 0.02329 | 0.018706 | 0.015032 | 0.012128 | 0.0097734 | 0.0079623 | 0.0064562 | 0.0052363 | 0.0041244 | 0.0033647 | 0.0027806 | 0.0022635 | 0.0029162 | 0.0030408 | 0.0014411 | 0.0021872 | 0.0031647 | 0.0015537
Time 19.1 h | 0.096631 | 0.084665 | 0.074113 | 0.066591 | 0.062691 | 0.058864 | 0.058511 | 0.060476 | 0.065071 | 0.072096 | 0.081916 | 0.094578 | 0.10999 | 0.12716 | 0.1449 | 0.16219 | 0.17646 | 0.18582 | 0.1881 | 0.18242 | 0.16923 | 0.15032 | 0.12827 | 0.10587 | 0.085142 | 0.067241 | 0.052563 | 0.040958 | 0.031878 | 0.024951 | 0.019669 | 0.015531 | 0.012366 | 0.0098704 | 0.0079292 | 0.0063816 | 0.0051741 | 0.0042062 | 0.0034663 | 0.0028605 | 0.0023538 | 0.0018919 | 0.0016185 | 0.0014417 | 0.0019173 | 0.0019251 | 0.0021233 | 0.0015948 | 0.0015646 | 0.0023139 | 0.00072884
 | 0.11394 | 0.093481 | 0.077657 | 0.065696 | 0.057013 | 0.051186 | 0.047934 | 0.047103 | 0.048651 | 0.05264 | 0.059221 | 0.068608 | 0.08103 | 0.096621 | 0.11525 | 0.13624 | 0.15804 | 0.17801 | 0.19272 | 0.1989 | 0.19496 | 0.18175 | 0.16215 | 0.13967 | 0.1172 | 0.096485 | 0.078464 | 0.063316 | 0.050716 | 0.040549 | 0.032372 | 0.025721 | 0.020479 | 0.01627 | 0.012943 | 0.010261 | 0.008146 | 0.0064366 | 0.0051337 | 0.0040747 | 0.00326 | 0.0025268 | 0.0020378 | 0.0016303
 | 0.12912 | 0.10754 | 0.091277 | 0.079538 | 0.071764 | 0.067588 | 0.066822 | 0.06944 | 0.075575 | 0.085504 | 0.099635 | 0.11846 | 0.1425 | 0.17204 | 0.20691 | 0.24592 | 0.28627 | 0.32317 | 0.35038 | 0.36195 | 0.35498 | 0.33106 | 0.29542 | 0.25449 | 0.21357 | 0.17583 | 0.143 | 0.11539 | 0.092429 | 0.0739 | 0.058998 | 0.046874 | 0.037321 | 0.02965 | 0.023587 | 0.018699 | 0.014845 | 0.01173 | 0.0093551 | 0.0074254 | 0.0059407 | 0.0046044 | 0.0037133 | 0.0029707
 | 0.1177 | 0.10593 | 0.097045 | 0.090946 | 0.087643 | 0.087239 | 0.089948 | 0.096102 | 0.10617 | 0.12074 | 0.14056 | 0.16644 | 0.1992 | 0.2394 | 0.28694 | 0.3404 | 0.39622 | 0.44803 | 0.48732 | 0.50564 | 0.49835 | 0.46702 | 0.41851 | 0.36179 | 0.30443 | 0.25118 | 0.20462 | 0.16535 | 0.13261 | 0.10615 | 0.084841 | 0.067487 | 0.053798 | 0.042795 | 0.034092 | 0.027068 | 0.021523 | 0.017036 | 0.013612 | 0.010826 | 0.00868 | 0.0067451 | 0.0054527 | 0.0043739
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625 | 5.7569630370485534 | 5.8239109124113 | 5.9030899869919438 | 6
Time 13.1 h | 0.10929 | 0.1004 | 0.092502 | 0.087998 | 0.086049 | 0.088115 | 0.0938 | 0.10112 | 0.11429 | 0.13106 | 0.15364 | 0.18166 | 0.21583 | 0.25512 | 0.29985 | 0.34774 | 0.3951 | 0.43795 | 0.46981 | 0.48496 | 0.48002 | 0.45469 | 0.41262 | 0.36037 | 0.30465 | 0.25101 | 0.20331 | 0.16311 | 0.13007 | 0.10386 | 0.083113 | 0.06641 | 0.053298 | 0.042748 | 0.034381 | 0.027604 | 0.022219 | 0.017839 | 0.014449 | 0.011634 | 0.0093969 | 0.0073773 | 0.0059706 | 0.0046834 | 0.003604 | 0.003119 | 0.0043457 | 0.0022557 | 0.0032184 | 0.0030577 | 0.0023234
Time 16.1 h | 0.11007 | 0.10093 | 0.090991 | 0.082047 | 0.079543 | 0.078788 | 0.081723 | 0.087173 | 0.096992 | 0.11005 | 0.12829 | 0.15003 | 0.17689 | 0.20803 | 0.24102 | 0.27482 | 0.30518 | 0.32877 | 0.34107 | 0.33971 | 0.32434 | 0.2969 | 0.26113 | 0.22177 | 0.1829 | 0.1475 | 0.11733 | 0.092752 | 0.073068 | 0.057802 | 0.045912 | 0.036469 | 0.029123 | 0.02329 | 0.018706 | 0.015032 | 0.012128 | 0.0097734 | 0.0079623 | 0.0064562 | 0.0052363 | 0.0041244 | 0.0033647 | 0.0027806 | 0.0022635 | 0.0029162 | 0.0030408 | 0.0014411 | 0.0021872 | 0.0031647 | 0.0015537
Time 19.1 h | 0.096631 | 0.084665 | 0.074113 | 0.066591 | 0.062691 | 0.058864 | 0.058511 | 0.060476 | 0.065071 | 0.072096 | 0.081916 | 0.094578 | 0.10999 | 0.12716 | 0.1449 | 0.16219 | 0.17646 | 0.18582 | 0.1881 | 0.18242 | 0.16923 | 0.15032 | 0.12827 | 0.10587 | 0.085142 | 0.067241 | 0.052563 | 0.040958 | 0.031878 | 0.024951 | 0.019669 | 0.015531 | 0.012366 | 0.0098704 | 0.0079292 | 0.0063816 | 0.0051741 | 0.0042062 | 0.0034663 | 0.0028605 | 0.0023538 | 0.0018919 | 0.0016185 | 0.0014417 | 0.0019173 | 0.0019251 | 0.0021233 | 0.0015948 | 0.0015646 | 0.0023139 | 0.00072884
 | 0.11394 | 0.093481 | 0.077657 | 0.065696 | 0.057013 | 0.051186 | 0.047934 | 0.047103 | 0.048651 | 0.05264 | 0.059221 | 0.068608 | 0.08103 | 0.096621 | 0.11525 | 0.13624 | 0.15804 | 0.17801 | 0.19272 | 0.1989 | 0.19496 | 0.18175 | 0.16215 | 0.13967 | 0.1172 | 0.096485 | 0.078464 | 0.063316 | 0.050716 | 0.040549 | 0.032372 | 0.025721 | 0.020479 | 0.01627 | 0.012943 | 0.010261 | 0.008146 | 0.0064366 | 0.0051337 | 0.0040747 | 0.00326 | 0.0025268 | 0.0020378 | 0.0016303
 | 0.12912 | 0.10754 | 0.091277 | 0.079538 | 0.071764 | 0.067588 | 0.066822 | 0.06944 | 0.075575 | 0.085504 | 0.099635 | 0.11846 | 0.1425 | 0.17204 | 0.20691 | 0.24592 | 0.28627 | 0.32317 | 0.35038 | 0.36195 | 0.35498 | 0.33106 | 0.29542 | 0.25449 | 0.21357 | 0.17583 | 0.143 | 0.11539 | 0.092429 | 0.0739 | 0.058998 | 0.046874 | 0.037321 | 0.02965 | 0.023587 | 0.018699 | 0.014845 | 0.01173 | 0.0093551 | 0.0074254 | 0.0059407 | 0.0046044 | 0.0037133 | 0.0029707
 | 0.1177 | 0.10593 | 0.097045 | 0.090946 | 0.087643 | 0.087239 | 0.089948 | 0.096102 | 0.10617 | 0.12074 | 0.14056 | 0.16644 | 0.1992 | 0.2394 | 0.28694 | 0.3404 | 0.39622 | 0.44803 | 0.48732 | 0.50564 | 0.49835 | 0.46702 | 0.41851 | 0.36179 | 0.30443 | 0.25118 | 0.20462 | 0.16535 | 0.13261 | 0.10615 | 0.084841 | 0.067487 | 0.053798 | 0.042795 | 0.034092 | 0.027068 | 0.021523 | 0.017036 | 0.013612 | 0.010826 | 0.00868 | 0.0067451 | 0.0054527 | 0.0043739
increase
Drying time
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 4.4 h | 0.003361 | 0.0040465 | 0.0025946 | 0.0026214 | 0.0044553 | 0.0026323 | 0.0029087 | 0.0037366 | 0.0047149 | 0.0057564 | 0.0073092 | 0.0089323 | 0.011077 | 0.013703 | 0.017072 | 0.021236 | 0.026498 | 0.03301 | 0.041177 | 0.051293 | 0.064055 | 0.079683 | 0.099225 | 0.12384 | 0.15437 | 0.19231 | 0.23844 | 0.29224 | 0.35137 | 0.41125 | 0.46468 | 0.50437 | 0.52462 | 0.52356 | 0.50251 | 0.46708 | 0.42347 | 0.37253 | 0.32299 | 0.27537 | 0.23266 | 0.1978 | 0.16633 | 0.14424 | 0.12252 | 0.10889 | 0.097697 | 0.093226 | 0.09166 | 0.092155 | 0.093069
Time 4.5 h | 0.0034249 | 0.0041377 | 0.0027283 | 0.0027814 | 0.0046887 | 0.0029259 | 0.0032733 | 0.0042363 | 0.0053358 | 0.0065296 | 0.0082746 | 0.010141 | 0.012601 | 0.015587 | 0.01947 | 0.024233 | 0.030264 | 0.037737 | 0.047171 | 0.058901 | 0.073903 | 0.092482 | 0.11581 | 0.14538 | 0.18208 | 0.2274 | 0.28124 | 0.3416 | 0.40418 | 0.46243 | 0.50811 | 0.53414 | 0.53731 | 0.51809 | 0.48197 | 0.43577 | 0.38293 | 0.3311 | 0.28177 | 0.23748 | 0.20077 | 0.16802 | 0.14416 | 0.11848 | 0.11171 | 0.10187 | 0.09506 | 0.092791 | 0.094429 | 0.097904 | 0.10371
Time 4.6 h | 0.0036091 | 0.0043103 | 0.0029276 | 0.0030069 | 0.0049863 | 0.0032919 | 0.0037411 | 0.0048548 | 0.006126 | 0.0075084 | 0.0095269 | 0.011725 | 0.014579 | 0.018062 | 0.02259 | 0.028182 | 0.035264 | 0.044078 | 0.055262 | 0.069282 | 0.08729 | 0.10998 | 0.13862 | 0.17503 | 0.21981 | 0.27381 | 0.33549 | 0.40061 | 0.46249 | 0.51282 | 0.5432 | 0.54939 | 0.53148 | 0.49439 | 0.44517 | 0.39079 | 0.33577 | 0.28667 | 0.23998 | 0.20372 | 0.17181 | 0.14592 | 0.12721 | 0.11818 | 0.10489 | 0.098824 | 0.097908 | 0.099283 | 0.10394 | 0.11202 | 0.12205
Time 4.7 h | 0.0040254 | 0.0023197 | 0.0060162 | 0.0032852 | 0.0053707 | 0.003773 | 0.0043423 | 0.005672 | 0.0071637 | 0.0088006 | 0.011181 | 0.013771 | 0.017191 | 0.021329 | 0.026752 | 0.033478 | 0.04203 | 0.052722 | 0.066401 | 0.083738 | 0.10613 | 0.13436 | 0.17037 | 0.21573 | 0.27035 | 0.33347 | 0.40107 | 0.46603 | 0.51937 | 0.55253 | 0.55996 | 0.54171 | 0.50311 | 0.45097 | 0.39455 | 0.33876 | 0.28887 | 0.24104 | 0.20387 | 0.17339 | 0.14886 | 0.13225 | 0.11753 | 0.11 | 0.10482 | 0.10438 | 0.1074 | 0.1138 | 0.12408 | 0.13639 | 0.14993
Time 4.8 h | 0.0039527 | 0.0047899 | 0.0034912 | 0.0036734 | 0.0059298 | 0.004463 | 0.0052209 | 0.0068624 | 0.0086692 | 0.010687 | 0.013626 | 0.016821 | 0.021041 | 0.026213 | 0.03302 | 0.041432 | 0.052261 | 0.065911 | 0.083539 | 0.10606 | 0.13545 | 0.17252 | 0.21928 | 0.27665 | 0.34237 | 0.41271 | 0.47977 | 0.53353 | 0.56486 | 0.56874 | 0.5461 | 0.50317 | 0.44813 | 0.38872 | 0.33224 | 0.28045 | 0.23905 | 0.20172 | 0.17254 | 0.14998 | 0.13098 | 0.12232 | 0.11386 | 0.11155 | 0.11412 | 0.11812 | 0.12658 | 0.13848 | 0.15361 | 0.17099 | 0.18866
Time 4.9 h | 0.0043188 | 0.0052277 | 0.0039868 | 0.0042288 | 0.0066529 | 0.0053479 | 0.0063485 | 0.0083796 | 0.010604 | 0.013107 | 0.016765 | 0.020805 | 0.02609 | 0.032614 | 0.041224 | 0.052023 | 0.065957 | 0.083695 | 0.10683 | 0.13651 | 0.17523 | 0.22353 | 0.28261 | 0.35128 | 0.42376 | 0.49222 | 0.54583 | 0.5749 | 0.57491 | 0.54814 | 0.50162 | 0.4443 | 0.38384 | 0.32634 | 0.27564 | 0.23272 | 0.19908 | 0.1717 | 0.1495 | 0.13558 | 0.12502 | 0.11938 | 0.11838 | 0.12437 | 0.13062 | 0.13992 | 0.15492 | 0.1723 | 0.19073 | 0.21306 | 0.23377
Time 5 h | 0.0049133 | 0.0058042 | 0.0046338 | 0.0049712 | 0.0076664 | 0.0066031 | 0.0079479 | 0.010533 | 0.013358 | 0.01657 | 0.021298 | 0.026516 | 0.033397 | 0.041936 | 0.053268 | 0.067599 | 0.086277 | 0.11024 | 0.14161 | 0.18183 | 0.2334 | 0.29549 | 0.36662 | 0.44167 | 0.50969 | 0.5605 | 0.58424 | 0.57776 | 0.54458 | 0.49349 | 0.43374 | 0.3731 | 0.31662 | 0.26729 | 0.22557 | 0.19336 | 0.16657 | 0.14984 | 0.13717 | 0.1291 | 0.12498 | 0.12759 | 0.13258 | 0.1444 | 0.15601 | 0.17194 | 0.19151 | 0.21398 | 0.23437 | 0.26066 | 0.28203
Time 5.1 h | 0.0051615 | 0.0065394 | 0.0054847 | 0.0059393 | 0.0090001 | 0.0082333 | 0.010001 | 0.013324 | 0.016914 | 0.021051 | 0.027147 | 0.03396 | 0.042946 | 0.054219 | 0.069319 | 0.088593 | 0.11394 | 0.14657 | 0.18909 | 0.24251 | 0.30812 | 0.38167 | 0.45767 | 0.52576 | 0.57299 | 0.59132 | 0.57824 | 0.5397 | 0.48481 | 0.42344 | 0.36292 | 0.3071 | 0.25963 | 0.22024 | 0.18846 | 0.16573 | 0.14947 | 0.13846 | 0.13032 | 0.13113 | 0.1334 | 0.1422 | 0.15348 | 0.16641 | 0.18887 | 0.2096 | 0.23238 | 0.25699 | 0.28353 | 0.3063 | 0.32649
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 4.4 h | 0.003361 | 0.0040465 | 0.0025946 | 0.0026214 | 0.0044553 | 0.0026323 | 0.0029087 | 0.0037366 | 0.0047149 | 0.0057564 | 0.0073092 | 0.0089323 | 0.011077 | 0.013703 | 0.017072 | 0.021236 | 0.026498 | 0.03301 | 0.041177 | 0.051293 | 0.064055 | 0.079683 | 0.099225 | 0.12384 | 0.15437 | 0.19231 | 0.23844 | 0.29224 | 0.35137 | 0.41125 | 0.46468 | 0.50437 | 0.52462 | 0.52356 | 0.50251 | 0.46708 | 0.42347 | 0.37253 | 0.32299 | 0.27537 | 0.23266 | 0.1978 | 0.16633 | 0.14424 | 0.12252 | 0.10889 | 0.097697 | 0.093226 | 0.09166 | 0.092155 | 0.093069
Time 4.5 h | 0.0034249 | 0.0041377 | 0.0027283 | 0.0027814 | 0.0046887 | 0.0029259 | 0.0032733 | 0.0042363 | 0.0053358 | 0.0065296 | 0.0082746 | 0.010141 | 0.012601 | 0.015587 | 0.01947 | 0.024233 | 0.030264 | 0.037737 | 0.047171 | 0.058901 | 0.073903 | 0.092482 | 0.11581 | 0.14538 | 0.18208 | 0.2274 | 0.28124 | 0.3416 | 0.40418 | 0.46243 | 0.50811 | 0.53414 | 0.53731 | 0.51809 | 0.48197 | 0.43577 | 0.38293 | 0.3311 | 0.28177 | 0.23748 | 0.20077 | 0.16802 | 0.14416 | 0.11848 | 0.11171 | 0.10187 | 0.09506 | 0.092791 | 0.094429 | 0.097904 | 0.10371
Time 4.6 h | 0.0036091 | 0.0043103 | 0.0029276 | 0.0030069 | 0.0049863 | 0.0032919 | 0.0037411 | 0.0048548 | 0.006126 | 0.0075084 | 0.0095269 | 0.011725 | 0.014579 | 0.018062 | 0.02259 | 0.028182 | 0.035264 | 0.044078 | 0.055262 | 0.069282 | 0.08729 | 0.10998 | 0.13862 | 0.17503 | 0.21981 | 0.27381 | 0.33549 | 0.40061 | 0.46249 | 0.51282 | 0.5432 | 0.54939 | 0.53148 | 0.49439 | 0.44517 | 0.39079 | 0.33577 | 0.28667 | 0.23998 | 0.20372 | 0.17181 | 0.14592 | 0.12721 | 0.11818 | 0.10489 | 0.098824 | 0.097908 | 0.099283 | 0.10394 | 0.11202 | 0.12205
Time 4.7 h | 0.0040254 | 0.0023197 | 0.0060162 | 0.0032852 | 0.0053707 | 0.003773 | 0.0043423 | 0.005672 | 0.0071637 | 0.0088006 | 0.011181 | 0.013771 | 0.017191 | 0.021329 | 0.026752 | 0.033478 | 0.04203 | 0.052722 | 0.066401 | 0.083738 | 0.10613 | 0.13436 | 0.17037 | 0.21573 | 0.27035 | 0.33347 | 0.40107 | 0.46603 | 0.51937 | 0.55253 | 0.55996 | 0.54171 | 0.50311 | 0.45097 | 0.39455 | 0.33876 | 0.28887 | 0.24104 | 0.20387 | 0.17339 | 0.14886 | 0.13225 | 0.11753 | 0.11 | 0.10482 | 0.10438 | 0.1074 | 0.1138 | 0.12408 | 0.13639 | 0.14993
Time 4.8 h | 0.0039527 | 0.0047899 | 0.0034912 | 0.0036734 | 0.0059298 | 0.004463 | 0.0052209 | 0.0068624 | 0.0086692 | 0.010687 | 0.013626 | 0.016821 | 0.021041 | 0.026213 | 0.03302 | 0.041432 | 0.052261 | 0.065911 | 0.083539 | 0.10606 | 0.13545 | 0.17252 | 0.21928 | 0.27665 | 0.34237 | 0.41271 | 0.47977 | 0.53353 | 0.56486 | 0.56874 | 0.5461 | 0.50317 | 0.44813 | 0.38872 | 0.33224 | 0.28045 | 0.23905 | 0.20172 | 0.17254 | 0.14998 | 0.13098 | 0.12232 | 0.11386 | 0.11155 | 0.11412 | 0.11812 | 0.12658 | 0.13848 | 0.15361 | 0.17099 | 0.18866
Time 4.9 h | 0.0043188 | 0.0052277 | 0.0039868 | 0.0042288 | 0.0066529 | 0.0053479 | 0.0063485 | 0.0083796 | 0.010604 | 0.013107 | 0.016765 | 0.020805 | 0.02609 | 0.032614 | 0.041224 | 0.052023 | 0.065957 | 0.083695 | 0.10683 | 0.13651 | 0.17523 | 0.22353 | 0.28261 | 0.35128 | 0.42376 | 0.49222 | 0.54583 | 0.5749 | 0.57491 | 0.54814 | 0.50162 | 0.4443 | 0.38384 | 0.32634 | 0.27564 | 0.23272 | 0.19908 | 0.1717 | 0.1495 | 0.13558 | 0.12502 | 0.11938 | 0.11838 | 0.12437 | 0.13062 | 0.13992 | 0.15492 | 0.1723 | 0.19073 | 0.21306 | 0.23377
Time 5 h | 0.0049133 | 0.0058042 | 0.0046338 | 0.0049712 | 0.0076664 | 0.0066031 | 0.0079479 | 0.010533 | 0.013358 | 0.01657 | 0.021298 | 0.026516 | 0.033397 | 0.041936 | 0.053268 | 0.067599 | 0.086277 | 0.11024 | 0.14161 | 0.18183 | 0.2334 | 0.29549 | 0.36662 | 0.44167 | 0.50969 | 0.5605 | 0.58424 | 0.57776 | 0.54458 | 0.49349 | 0.43374 | 0.3731 | 0.31662 | 0.26729 | 0.22557 | 0.19336 | 0.16657 | 0.14984 | 0.13717 | 0.1291 | 0.12498 | 0.12759 | 0.13258 | 0.1444 | 0.15601 | 0.17194 | 0.19151 | 0.21398 | 0.23437 | 0.26066 | 0.28203
Time 5.1 h | 0.0051615 | 0.0065394 | 0.0054847 | 0.0059393 | 0.0090001 | 0.0082333 | 0.010001 | 0.013324 | 0.016914 | 0.021051 | 0.027147 | 0.03396 | 0.042946 | 0.054219 | 0.069319 | 0.088593 | 0.11394 | 0.14657 | 0.18909 | 0.24251 | 0.30812 | 0.38167 | 0.45767 | 0.52576 | 0.57299 | 0.59132 | 0.57824 | 0.5397 | 0.48481 | 0.42344 | 0.36292 | 0.3071 | 0.25963 | 0.22024 | 0.18846 | 0.16573 | 0.14947 | 0.13846 | 0.13032 | 0.13113 | 0.1334 | 0.1422 | 0.15348 | 0.16641 | 0.18887 | 0.2096 | 0.23238 | 0.25699 | 0.28353 | 0.3063 | 0.32649
-38 oC
-18 oC
Monitoring Phase Behaviour (ice nucleation temperature and solidification end points by using
temperature calibration for predicting temperature of the product in primary drying
Surrogate drying rate (from )
Notes: Eg temperature cycling below to calibrate Fpeak for temperature (in order to predict product temperatures in primary drying) and the compensate cpeak for temperature so that it can used to monitor ice content and hence drying rate
Eg 2 left graph: monitoring phase behaviour (ice nucleation temperatures and solidufication end points)

## Slide 11: Through Vial Impedance Spectroscopy (TVIS)
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034
Time 0.3 h | 0.27396 | 0.24912 | 0.22663 | 0.2063 | 0.18797 | 0.17147 | 0.15668 | 0.14349 | 0.13183 | 0.12162 | 0.11284 | 0.10548 | 0.099574 | 0.095186 | 0.092431 | 0.091474 | 0.092551 | 0.095973 | 0.10215 | 0.11162 | 0.12501 | 0.14319 | 0.16715 | 0.198 | 0.23699 | 0.28544 | 0.34395 | 0.41214 | 0.48865 | 0.56683 | 0.63782 | 0.69048 | 0.71191 | 0.69725 | 0.65001 | 0.57934 | 0.49898 | 0.41726 | 0.34526 | 0.28116 | 0.22886 | 0.17991 | 0.14643 | 0.1181 | 0.089415 | 0.072054
Time 4 h | 0.099081 | 0.096153 | 0.095569 | 0.097591 | 0.10258 | 0.11101 | 0.12349 | 0.14075 | 0.16366 | 0.19315 | 0.23007 | 0.27492 | 0.32735 | 0.38533 | 0.44432 | 0.49685 | 0.53348 | 0.54588 | 0.53067 | 0.49125 | 0.43601 | 0.37424 | 0.31336 | 0.25794 | 0.20992 | 0.16948 | 0.13623 | 0.10926 | 0.087318 | 0.069838 | 0.055881 | 0.044566 | 0.035658 | 0.0285 | 0.02283 | 0.018246 | 0.014616 | 0.011669 | 0.0094101 | 0.0075633 | 0.0061328 | 0.0048349
Time 0.3 h | 0.26705 | 0.24478 | 0.21978 | 0.1979 | 0.17863 | 0.15615 | 0.14276 | 0.12639 | 0.11756 | 0.11024 | 0.10127 | 0.1004 | 0.098391 | 0.10089 | 0.10904 | 0.11009 | 0.11451 | 0.12496 | 0.13439 | 0.14683 | 0.16234 | 0.18119 | 0.20238 | 0.22944 | 0.26256 | 0.30331 | 0.35113 | 0.4089 | 0.47307 | 0.54005 | 0.60366 | 0.65431 | 0.67974 | 0.67269 | 0.63391 | 0.56955 | 0.49367 | 0.41413 | 0.34323 | 0.27998 | 0.22777
Time 4 h | 0.093429 | 0.092843 | 0.096342 | 0.10075 | 0.11232 | 0.12322 | 0.1435 | 0.1596 | 0.19684 | 0.23027 | 0.27214 | 0.31647 | 0.36246 | 0.41025 | 0.45421 | 0.48896 | 0.51147 | 0.51695 | 0.50285 | 0.46926 | 0.42098 | 0.36431 | 0.30625 | 0.25218 | 0.20511 | 0.16565 | 0.13355 | 0.10767 | 0.086605 | 0.069689 | 0.056026 | 0.045028 | 0.036377 | 0.029122 | 0.023205 | 0.018299 | 0.014688 | 0.011778 | 0.0095536 | 0.0077519 | 0.0063112
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034
Time 0.3 h | 0.27396 | 0.24912 | 0.22663 | 0.2063 | 0.18797 | 0.17147 | 0.15668 | 0.14349 | 0.13183 | 0.12162 | 0.11284 | 0.10548 | 0.099574 | 0.095186 | 0.092431 | 0.091474 | 0.092551 | 0.095973 | 0.10215 | 0.11162 | 0.12501 | 0.14319 | 0.16715 | 0.198 | 0.23699 | 0.28544 | 0.34395 | 0.41214 | 0.48865 | 0.56683 | 0.63782 | 0.69048 | 0.71191 | 0.69725 | 0.65001 | 0.57934 | 0.49898 | 0.41726 | 0.34526 | 0.28116 | 0.22886 | 0.17991 | 0.14643 | 0.1181 | 0.089415 | 0.072054
Time 4 h | 0.099081 | 0.096153 | 0.095569 | 0.097591 | 0.10258 | 0.11101 | 0.12349 | 0.14075 | 0.16366 | 0.19315 | 0.23007 | 0.27492 | 0.32735 | 0.38533 | 0.44432 | 0.49685 | 0.53348 | 0.54588 | 0.53067 | 0.49125 | 0.43601 | 0.37424 | 0.31336 | 0.25794 | 0.20992 | 0.16948 | 0.13623 | 0.10926 | 0.087318 | 0.069838 | 0.055881 | 0.044566 | 0.035658 | 0.0285 | 0.02283 | 0.018246 | 0.014616 | 0.011669 | 0.0094101 | 0.0075633 | 0.0061328 | 0.0048349
Time 0.3 h | 0.26705 | 0.24478 | 0.21978 | 0.1979 | 0.17863 | 0.15615 | 0.14276 | 0.12639 | 0.11756 | 0.11024 | 0.10127 | 0.1004 | 0.098391 | 0.10089 | 0.10904 | 0.11009 | 0.11451 | 0.12496 | 0.13439 | 0.14683 | 0.16234 | 0.18119 | 0.20238 | 0.22944 | 0.26256 | 0.30331 | 0.35113 | 0.4089 | 0.47307 | 0.54005 | 0.60366 | 0.65431 | 0.67974 | 0.67269 | 0.63391 | 0.56955 | 0.49367 | 0.41413 | 0.34323 | 0.27998 | 0.22777
Time 4 h | 0.093429 | 0.092843 | 0.096342 | 0.10075 | 0.11232 | 0.12322 | 0.1435 | 0.1596 | 0.19684 | 0.23027 | 0.27214 | 0.31647 | 0.36246 | 0.41025 | 0.45421 | 0.48896 | 0.51147 | 0.51695 | 0.50285 | 0.46926 | 0.42098 | 0.36431 | 0.30625 | 0.25218 | 0.20511 | 0.16565 | 0.13355 | 0.10767 | 0.086605 | 0.069689 | 0.056026 | 0.045028 | 0.036377 | 0.029122 | 0.023205 | 0.018299 | 0.014688 | 0.011778 | 0.0095536 | 0.0077519 | 0.0063112
FPEAK
FPEAK
Liquid State
Solid State
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625 | 5.7569630370485534 | 5.8239109124113 | 5.9030899869919438 | 6
Time 13.1 h | 0.10929 | 0.1004 | 0.092502 | 0.087998 | 0.086049 | 0.088115 | 0.0938 | 0.10112 | 0.11429 | 0.13106 | 0.15364 | 0.18166 | 0.21583 | 0.25512 | 0.29985 | 0.34774 | 0.3951 | 0.43795 | 0.46981 | 0.48496 | 0.48002 | 0.45469 | 0.41262 | 0.36037 | 0.30465 | 0.25101 | 0.20331 | 0.16311 | 0.13007 | 0.10386 | 0.083113 | 0.06641 | 0.053298 | 0.042748 | 0.034381 | 0.027604 | 0.022219 | 0.017839 | 0.014449 | 0.011634 | 0.0093969 | 0.0073773 | 0.0059706 | 0.0046834 | 0.003604 | 0.003119 | 0.0043457 | 0.0022557 | 0.0032184 | 0.0030577 | 0.0023234
Time 16.1 h | 0.11007 | 0.10093 | 0.090991 | 0.082047 | 0.079543 | 0.078788 | 0.081723 | 0.087173 | 0.096992 | 0.11005 | 0.12829 | 0.15003 | 0.17689 | 0.20803 | 0.24102 | 0.27482 | 0.30518 | 0.32877 | 0.34107 | 0.33971 | 0.32434 | 0.2969 | 0.26113 | 0.22177 | 0.1829 | 0.1475 | 0.11733 | 0.092752 | 0.073068 | 0.057802 | 0.045912 | 0.036469 | 0.029123 | 0.02329 | 0.018706 | 0.015032 | 0.012128 | 0.0097734 | 0.0079623 | 0.0064562 | 0.0052363 | 0.0041244 | 0.0033647 | 0.0027806 | 0.0022635 | 0.0029162 | 0.0030408 | 0.0014411 | 0.0021872 | 0.0031647 | 0.0015537
Time 19.1 h | 0.096631 | 0.084665 | 0.074113 | 0.066591 | 0.062691 | 0.058864 | 0.058511 | 0.060476 | 0.065071 | 0.072096 | 0.081916 | 0.094578 | 0.10999 | 0.12716 | 0.1449 | 0.16219 | 0.17646 | 0.18582 | 0.1881 | 0.18242 | 0.16923 | 0.15032 | 0.12827 | 0.10587 | 0.085142 | 0.067241 | 0.052563 | 0.040958 | 0.031878 | 0.024951 | 0.019669 | 0.015531 | 0.012366 | 0.0098704 | 0.0079292 | 0.0063816 | 0.0051741 | 0.0042062 | 0.0034663 | 0.0028605 | 0.0023538 | 0.0018919 | 0.0016185 | 0.0014417 | 0.0019173 | 0.0019251 | 0.0021233 | 0.0015948 | 0.0015646 | 0.0023139 | 0.00072884
 | 0.11394 | 0.093481 | 0.077657 | 0.065696 | 0.057013 | 0.051186 | 0.047934 | 0.047103 | 0.048651 | 0.05264 | 0.059221 | 0.068608 | 0.08103 | 0.096621 | 0.11525 | 0.13624 | 0.15804 | 0.17801 | 0.19272 | 0.1989 | 0.19496 | 0.18175 | 0.16215 | 0.13967 | 0.1172 | 0.096485 | 0.078464 | 0.063316 | 0.050716 | 0.040549 | 0.032372 | 0.025721 | 0.020479 | 0.01627 | 0.012943 | 0.010261 | 0.008146 | 0.0064366 | 0.0051337 | 0.0040747 | 0.00326 | 0.0025268 | 0.0020378 | 0.0016303
 | 0.12912 | 0.10754 | 0.091277 | 0.079538 | 0.071764 | 0.067588 | 0.066822 | 0.06944 | 0.075575 | 0.085504 | 0.099635 | 0.11846 | 0.1425 | 0.17204 | 0.20691 | 0.24592 | 0.28627 | 0.32317 | 0.35038 | 0.36195 | 0.35498 | 0.33106 | 0.29542 | 0.25449 | 0.21357 | 0.17583 | 0.143 | 0.11539 | 0.092429 | 0.0739 | 0.058998 | 0.046874 | 0.037321 | 0.02965 | 0.023587 | 0.018699 | 0.014845 | 0.01173 | 0.0093551 | 0.0074254 | 0.0059407 | 0.0046044 | 0.0037133 | 0.0029707
 | 0.1177 | 0.10593 | 0.097045 | 0.090946 | 0.087643 | 0.087239 | 0.089948 | 0.096102 | 0.10617 | 0.12074 | 0.14056 | 0.16644 | 0.1992 | 0.2394 | 0.28694 | 0.3404 | 0.39622 | 0.44803 | 0.48732 | 0.50564 | 0.49835 | 0.46702 | 0.41851 | 0.36179 | 0.30443 | 0.25118 | 0.20462 | 0.16535 | 0.13261 | 0.10615 | 0.084841 | 0.067487 | 0.053798 | 0.042795 | 0.034092 | 0.027068 | 0.021523 | 0.017036 | 0.013612 | 0.010826 | 0.00868 | 0.0067451 | 0.0054527 | 0.0043739
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625 | 5.7569630370485534 | 5.8239109124113 | 5.9030899869919438 | 6
Time 13.1 h | 0.10929 | 0.1004 | 0.092502 | 0.087998 | 0.086049 | 0.088115 | 0.0938 | 0.10112 | 0.11429 | 0.13106 | 0.15364 | 0.18166 | 0.21583 | 0.25512 | 0.29985 | 0.34774 | 0.3951 | 0.43795 | 0.46981 | 0.48496 | 0.48002 | 0.45469 | 0.41262 | 0.36037 | 0.30465 | 0.25101 | 0.20331 | 0.16311 | 0.13007 | 0.10386 | 0.083113 | 0.06641 | 0.053298 | 0.042748 | 0.034381 | 0.027604 | 0.022219 | 0.017839 | 0.014449 | 0.011634 | 0.0093969 | 0.0073773 | 0.0059706 | 0.0046834 | 0.003604 | 0.003119 | 0.0043457 | 0.0022557 | 0.0032184 | 0.0030577 | 0.0023234
Time 16.1 h | 0.11007 | 0.10093 | 0.090991 | 0.082047 | 0.079543 | 0.078788 | 0.081723 | 0.087173 | 0.096992 | 0.11005 | 0.12829 | 0.15003 | 0.17689 | 0.20803 | 0.24102 | 0.27482 | 0.30518 | 0.32877 | 0.34107 | 0.33971 | 0.32434 | 0.2969 | 0.26113 | 0.22177 | 0.1829 | 0.1475 | 0.11733 | 0.092752 | 0.073068 | 0.057802 | 0.045912 | 0.036469 | 0.029123 | 0.02329 | 0.018706 | 0.015032 | 0.012128 | 0.0097734 | 0.0079623 | 0.0064562 | 0.0052363 | 0.0041244 | 0.0033647 | 0.0027806 | 0.0022635 | 0.0029162 | 0.0030408 | 0.0014411 | 0.0021872 | 0.0031647 | 0.0015537
Time 19.1 h | 0.096631 | 0.084665 | 0.074113 | 0.066591 | 0.062691 | 0.058864 | 0.058511 | 0.060476 | 0.065071 | 0.072096 | 0.081916 | 0.094578 | 0.10999 | 0.12716 | 0.1449 | 0.16219 | 0.17646 | 0.18582 | 0.1881 | 0.18242 | 0.16923 | 0.15032 | 0.12827 | 0.10587 | 0.085142 | 0.067241 | 0.052563 | 0.040958 | 0.031878 | 0.024951 | 0.019669 | 0.015531 | 0.012366 | 0.0098704 | 0.0079292 | 0.0063816 | 0.0051741 | 0.0042062 | 0.0034663 | 0.0028605 | 0.0023538 | 0.0018919 | 0.0016185 | 0.0014417 | 0.0019173 | 0.0019251 | 0.0021233 | 0.0015948 | 0.0015646 | 0.0023139 | 0.00072884
 | 0.11394 | 0.093481 | 0.077657 | 0.065696 | 0.057013 | 0.051186 | 0.047934 | 0.047103 | 0.048651 | 0.05264 | 0.059221 | 0.068608 | 0.08103 | 0.096621 | 0.11525 | 0.13624 | 0.15804 | 0.17801 | 0.19272 | 0.1989 | 0.19496 | 0.18175 | 0.16215 | 0.13967 | 0.1172 | 0.096485 | 0.078464 | 0.063316 | 0.050716 | 0.040549 | 0.032372 | 0.025721 | 0.020479 | 0.01627 | 0.012943 | 0.010261 | 0.008146 | 0.0064366 | 0.0051337 | 0.0040747 | 0.00326 | 0.0025268 | 0.0020378 | 0.0016303
 | 0.12912 | 0.10754 | 0.091277 | 0.079538 | 0.071764 | 0.067588 | 0.066822 | 0.06944 | 0.075575 | 0.085504 | 0.099635 | 0.11846 | 0.1425 | 0.17204 | 0.20691 | 0.24592 | 0.28627 | 0.32317 | 0.35038 | 0.36195 | 0.35498 | 0.33106 | 0.29542 | 0.25449 | 0.21357 | 0.17583 | 0.143 | 0.11539 | 0.092429 | 0.0739 | 0.058998 | 0.046874 | 0.037321 | 0.02965 | 0.023587 | 0.018699 | 0.014845 | 0.01173 | 0.0093551 | 0.0074254 | 0.0059407 | 0.0046044 | 0.0037133 | 0.0029707
 | 0.1177 | 0.10593 | 0.097045 | 0.090946 | 0.087643 | 0.087239 | 0.089948 | 0.096102 | 0.10617 | 0.12074 | 0.14056 | 0.16644 | 0.1992 | 0.2394 | 0.28694 | 0.3404 | 0.39622 | 0.44803 | 0.48732 | 0.50564 | 0.49835 | 0.46702 | 0.41851 | 0.36179 | 0.30443 | 0.25118 | 0.20462 | 0.16535 | 0.13261 | 0.10615 | 0.084841 | 0.067487 | 0.053798 | 0.042795 | 0.034092 | 0.027068 | 0.021523 | 0.017036 | 0.013612 | 0.010826 | 0.00868 | 0.0067451 | 0.0054527 | 0.0043739
increase
Drying time
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 4.4 h | 0.003361 | 0.0040465 | 0.0025946 | 0.0026214 | 0.0044553 | 0.0026323 | 0.0029087 | 0.0037366 | 0.0047149 | 0.0057564 | 0.0073092 | 0.0089323 | 0.011077 | 0.013703 | 0.017072 | 0.021236 | 0.026498 | 0.03301 | 0.041177 | 0.051293 | 0.064055 | 0.079683 | 0.099225 | 0.12384 | 0.15437 | 0.19231 | 0.23844 | 0.29224 | 0.35137 | 0.41125 | 0.46468 | 0.50437 | 0.52462 | 0.52356 | 0.50251 | 0.46708 | 0.42347 | 0.37253 | 0.32299 | 0.27537 | 0.23266 | 0.1978 | 0.16633 | 0.14424 | 0.12252 | 0.10889 | 0.097697 | 0.093226 | 0.09166 | 0.092155 | 0.093069
Time 4.5 h | 0.0034249 | 0.0041377 | 0.0027283 | 0.0027814 | 0.0046887 | 0.0029259 | 0.0032733 | 0.0042363 | 0.0053358 | 0.0065296 | 0.0082746 | 0.010141 | 0.012601 | 0.015587 | 0.01947 | 0.024233 | 0.030264 | 0.037737 | 0.047171 | 0.058901 | 0.073903 | 0.092482 | 0.11581 | 0.14538 | 0.18208 | 0.2274 | 0.28124 | 0.3416 | 0.40418 | 0.46243 | 0.50811 | 0.53414 | 0.53731 | 0.51809 | 0.48197 | 0.43577 | 0.38293 | 0.3311 | 0.28177 | 0.23748 | 0.20077 | 0.16802 | 0.14416 | 0.11848 | 0.11171 | 0.10187 | 0.09506 | 0.092791 | 0.094429 | 0.097904 | 0.10371
Time 4.6 h | 0.0036091 | 0.0043103 | 0.0029276 | 0.0030069 | 0.0049863 | 0.0032919 | 0.0037411 | 0.0048548 | 0.006126 | 0.0075084 | 0.0095269 | 0.011725 | 0.014579 | 0.018062 | 0.02259 | 0.028182 | 0.035264 | 0.044078 | 0.055262 | 0.069282 | 0.08729 | 0.10998 | 0.13862 | 0.17503 | 0.21981 | 0.27381 | 0.33549 | 0.40061 | 0.46249 | 0.51282 | 0.5432 | 0.54939 | 0.53148 | 0.49439 | 0.44517 | 0.39079 | 0.33577 | 0.28667 | 0.23998 | 0.20372 | 0.17181 | 0.14592 | 0.12721 | 0.11818 | 0.10489 | 0.098824 | 0.097908 | 0.099283 | 0.10394 | 0.11202 | 0.12205
Time 4.7 h | 0.0040254 | 0.0023197 | 0.0060162 | 0.0032852 | 0.0053707 | 0.003773 | 0.0043423 | 0.005672 | 0.0071637 | 0.0088006 | 0.011181 | 0.013771 | 0.017191 | 0.021329 | 0.026752 | 0.033478 | 0.04203 | 0.052722 | 0.066401 | 0.083738 | 0.10613 | 0.13436 | 0.17037 | 0.21573 | 0.27035 | 0.33347 | 0.40107 | 0.46603 | 0.51937 | 0.55253 | 0.55996 | 0.54171 | 0.50311 | 0.45097 | 0.39455 | 0.33876 | 0.28887 | 0.24104 | 0.20387 | 0.17339 | 0.14886 | 0.13225 | 0.11753 | 0.11 | 0.10482 | 0.10438 | 0.1074 | 0.1138 | 0.12408 | 0.13639 | 0.14993
Time 4.8 h | 0.0039527 | 0.0047899 | 0.0034912 | 0.0036734 | 0.0059298 | 0.004463 | 0.0052209 | 0.0068624 | 0.0086692 | 0.010687 | 0.013626 | 0.016821 | 0.021041 | 0.026213 | 0.03302 | 0.041432 | 0.052261 | 0.065911 | 0.083539 | 0.10606 | 0.13545 | 0.17252 | 0.21928 | 0.27665 | 0.34237 | 0.41271 | 0.47977 | 0.53353 | 0.56486 | 0.56874 | 0.5461 | 0.50317 | 0.44813 | 0.38872 | 0.33224 | 0.28045 | 0.23905 | 0.20172 | 0.17254 | 0.14998 | 0.13098 | 0.12232 | 0.11386 | 0.11155 | 0.11412 | 0.11812 | 0.12658 | 0.13848 | 0.15361 | 0.17099 | 0.18866
Time 4.9 h | 0.0043188 | 0.0052277 | 0.0039868 | 0.0042288 | 0.0066529 | 0.0053479 | 0.0063485 | 0.0083796 | 0.010604 | 0.013107 | 0.016765 | 0.020805 | 0.02609 | 0.032614 | 0.041224 | 0.052023 | 0.065957 | 0.083695 | 0.10683 | 0.13651 | 0.17523 | 0.22353 | 0.28261 | 0.35128 | 0.42376 | 0.49222 | 0.54583 | 0.5749 | 0.57491 | 0.54814 | 0.50162 | 0.4443 | 0.38384 | 0.32634 | 0.27564 | 0.23272 | 0.19908 | 0.1717 | 0.1495 | 0.13558 | 0.12502 | 0.11938 | 0.11838 | 0.12437 | 0.13062 | 0.13992 | 0.15492 | 0.1723 | 0.19073 | 0.21306 | 0.23377
Time 5 h | 0.0049133 | 0.0058042 | 0.0046338 | 0.0049712 | 0.0076664 | 0.0066031 | 0.0079479 | 0.010533 | 0.013358 | 0.01657 | 0.021298 | 0.026516 | 0.033397 | 0.041936 | 0.053268 | 0.067599 | 0.086277 | 0.11024 | 0.14161 | 0.18183 | 0.2334 | 0.29549 | 0.36662 | 0.44167 | 0.50969 | 0.5605 | 0.58424 | 0.57776 | 0.54458 | 0.49349 | 0.43374 | 0.3731 | 0.31662 | 0.26729 | 0.22557 | 0.19336 | 0.16657 | 0.14984 | 0.13717 | 0.1291 | 0.12498 | 0.12759 | 0.13258 | 0.1444 | 0.15601 | 0.17194 | 0.19151 | 0.21398 | 0.23437 | 0.26066 | 0.28203
Time 5.1 h | 0.0051615 | 0.0065394 | 0.0054847 | 0.0059393 | 0.0090001 | 0.0082333 | 0.010001 | 0.013324 | 0.016914 | 0.021051 | 0.027147 | 0.03396 | 0.042946 | 0.054219 | 0.069319 | 0.088593 | 0.11394 | 0.14657 | 0.18909 | 0.24251 | 0.30812 | 0.38167 | 0.45767 | 0.52576 | 0.57299 | 0.59132 | 0.57824 | 0.5397 | 0.48481 | 0.42344 | 0.36292 | 0.3071 | 0.25963 | 0.22024 | 0.18846 | 0.16573 | 0.14947 | 0.13846 | 0.13032 | 0.13113 | 0.1334 | 0.1422 | 0.15348 | 0.16641 | 0.18887 | 0.2096 | 0.23238 | 0.25699 | 0.28353 | 0.3063 | 0.32649
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 4.4 h | 0.003361 | 0.0040465 | 0.0025946 | 0.0026214 | 0.0044553 | 0.0026323 | 0.0029087 | 0.0037366 | 0.0047149 | 0.0057564 | 0.0073092 | 0.0089323 | 0.011077 | 0.013703 | 0.017072 | 0.021236 | 0.026498 | 0.03301 | 0.041177 | 0.051293 | 0.064055 | 0.079683 | 0.099225 | 0.12384 | 0.15437 | 0.19231 | 0.23844 | 0.29224 | 0.35137 | 0.41125 | 0.46468 | 0.50437 | 0.52462 | 0.52356 | 0.50251 | 0.46708 | 0.42347 | 0.37253 | 0.32299 | 0.27537 | 0.23266 | 0.1978 | 0.16633 | 0.14424 | 0.12252 | 0.10889 | 0.097697 | 0.093226 | 0.09166 | 0.092155 | 0.093069
Time 4.5 h | 0.0034249 | 0.0041377 | 0.0027283 | 0.0027814 | 0.0046887 | 0.0029259 | 0.0032733 | 0.0042363 | 0.0053358 | 0.0065296 | 0.0082746 | 0.010141 | 0.012601 | 0.015587 | 0.01947 | 0.024233 | 0.030264 | 0.037737 | 0.047171 | 0.058901 | 0.073903 | 0.092482 | 0.11581 | 0.14538 | 0.18208 | 0.2274 | 0.28124 | 0.3416 | 0.40418 | 0.46243 | 0.50811 | 0.53414 | 0.53731 | 0.51809 | 0.48197 | 0.43577 | 0.38293 | 0.3311 | 0.28177 | 0.23748 | 0.20077 | 0.16802 | 0.14416 | 0.11848 | 0.11171 | 0.10187 | 0.09506 | 0.092791 | 0.094429 | 0.097904 | 0.10371
Time 4.6 h | 0.0036091 | 0.0043103 | 0.0029276 | 0.0030069 | 0.0049863 | 0.0032919 | 0.0037411 | 0.0048548 | 0.006126 | 0.0075084 | 0.0095269 | 0.011725 | 0.014579 | 0.018062 | 0.02259 | 0.028182 | 0.035264 | 0.044078 | 0.055262 | 0.069282 | 0.08729 | 0.10998 | 0.13862 | 0.17503 | 0.21981 | 0.27381 | 0.33549 | 0.40061 | 0.46249 | 0.51282 | 0.5432 | 0.54939 | 0.53148 | 0.49439 | 0.44517 | 0.39079 | 0.33577 | 0.28667 | 0.23998 | 0.20372 | 0.17181 | 0.14592 | 0.12721 | 0.11818 | 0.10489 | 0.098824 | 0.097908 | 0.099283 | 0.10394 | 0.11202 | 0.12205
Time 4.7 h | 0.0040254 | 0.0023197 | 0.0060162 | 0.0032852 | 0.0053707 | 0.003773 | 0.0043423 | 0.005672 | 0.0071637 | 0.0088006 | 0.011181 | 0.013771 | 0.017191 | 0.021329 | 0.026752 | 0.033478 | 0.04203 | 0.052722 | 0.066401 | 0.083738 | 0.10613 | 0.13436 | 0.17037 | 0.21573 | 0.27035 | 0.33347 | 0.40107 | 0.46603 | 0.51937 | 0.55253 | 0.55996 | 0.54171 | 0.50311 | 0.45097 | 0.39455 | 0.33876 | 0.28887 | 0.24104 | 0.20387 | 0.17339 | 0.14886 | 0.13225 | 0.11753 | 0.11 | 0.10482 | 0.10438 | 0.1074 | 0.1138 | 0.12408 | 0.13639 | 0.14993
Time 4.8 h | 0.0039527 | 0.0047899 | 0.0034912 | 0.0036734 | 0.0059298 | 0.004463 | 0.0052209 | 0.0068624 | 0.0086692 | 0.010687 | 0.013626 | 0.016821 | 0.021041 | 0.026213 | 0.03302 | 0.041432 | 0.052261 | 0.065911 | 0.083539 | 0.10606 | 0.13545 | 0.17252 | 0.21928 | 0.27665 | 0.34237 | 0.41271 | 0.47977 | 0.53353 | 0.56486 | 0.56874 | 0.5461 | 0.50317 | 0.44813 | 0.38872 | 0.33224 | 0.28045 | 0.23905 | 0.20172 | 0.17254 | 0.14998 | 0.13098 | 0.12232 | 0.11386 | 0.11155 | 0.11412 | 0.11812 | 0.12658 | 0.13848 | 0.15361 | 0.17099 | 0.18866
Time 4.9 h | 0.0043188 | 0.0052277 | 0.0039868 | 0.0042288 | 0.0066529 | 0.0053479 | 0.0063485 | 0.0083796 | 0.010604 | 0.013107 | 0.016765 | 0.020805 | 0.02609 | 0.032614 | 0.041224 | 0.052023 | 0.065957 | 0.083695 | 0.10683 | 0.13651 | 0.17523 | 0.22353 | 0.28261 | 0.35128 | 0.42376 | 0.49222 | 0.54583 | 0.5749 | 0.57491 | 0.54814 | 0.50162 | 0.4443 | 0.38384 | 0.32634 | 0.27564 | 0.23272 | 0.19908 | 0.1717 | 0.1495 | 0.13558 | 0.12502 | 0.11938 | 0.11838 | 0.12437 | 0.13062 | 0.13992 | 0.15492 | 0.1723 | 0.19073 | 0.21306 | 0.23377
Time 5 h | 0.0049133 | 0.0058042 | 0.0046338 | 0.0049712 | 0.0076664 | 0.0066031 | 0.0079479 | 0.010533 | 0.013358 | 0.01657 | 0.021298 | 0.026516 | 0.033397 | 0.041936 | 0.053268 | 0.067599 | 0.086277 | 0.11024 | 0.14161 | 0.18183 | 0.2334 | 0.29549 | 0.36662 | 0.44167 | 0.50969 | 0.5605 | 0.58424 | 0.57776 | 0.54458 | 0.49349 | 0.43374 | 0.3731 | 0.31662 | 0.26729 | 0.22557 | 0.19336 | 0.16657 | 0.14984 | 0.13717 | 0.1291 | 0.12498 | 0.12759 | 0.13258 | 0.1444 | 0.15601 | 0.17194 | 0.19151 | 0.21398 | 0.23437 | 0.26066 | 0.28203
Time 5.1 h | 0.0051615 | 0.0065394 | 0.0054847 | 0.0059393 | 0.0090001 | 0.0082333 | 0.010001 | 0.013324 | 0.016914 | 0.021051 | 0.027147 | 0.03396 | 0.042946 | 0.054219 | 0.069319 | 0.088593 | 0.11394 | 0.14657 | 0.18909 | 0.24251 | 0.30812 | 0.38167 | 0.45767 | 0.52576 | 0.57299 | 0.59132 | 0.57824 | 0.5397 | 0.48481 | 0.42344 | 0.36292 | 0.3071 | 0.25963 | 0.22024 | 0.18846 | 0.16573 | 0.14947 | 0.13846 | 0.13032 | 0.13113 | 0.1334 | 0.1422 | 0.15348 | 0.16641 | 0.18887 | 0.2096 | 0.23238 | 0.25699 | 0.28353 | 0.3063 | 0.32649
-38 oC
-18 oC
Monitoring Phase Behaviour (ice nucleation temperature and solidification end points by using
temperature calibration for predicting temperature of the product in primary drying
Surrogate drying rate (from )
  is highly sensitive to low ice volumes; therefore it could be used for determination end point of primary drying
Notes: Eg temperature cycling below to calibrate Fpeak for temperature (in order to predict product temperatures in primary drying) and the compensate cpeak for temperature so that it can used to monitor ice content and hence drying rate
Eg 2 left graph: monitoring phase behaviour (ice nucleation temperatures and solidufication end points)

## Slide 12
Through Vial Impedance Spectroscopy (TVIS)
Dielectric Loss Mechanisms

## Slide 13: Dielectric Loss Mechanisms
The polarization of the water dipole in liquid water at 20 ˚C, with a dielectric loss peak frequency of ~ 18 GHz
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 0.23 h, 20.3 celcius | 1.0162 | 1.0162 | 1.0162 | 1.0162 | 1.0176 | 1.024 | 1.0293 | 1.0371 | 1.0465 | 1.0575 | 1.0742 | 1.0971 | 1.131 | 1.18 | 1.2538 | 1.3532 | 1.4803 | 1.6281 | 1.7827 | 1.9288 | 2.0567 | 2.1587 | 2.2374 | 2.2971 | 2.3414 | 2.3756 | 2.4038 | 2.4272 | 2.4476 | 2.4662 | 2.4852 | 2.5003 | 2.5152 | 2.5286 | 2.5419 | 2.5531 | 2.5615 | 2.5717 | 2.5802 | 2.589 | 2.5992 | 2.6061 | 2.6201 | 2.625 | 2.6478 | 2.6645 | 2.686 | 2.708 | 2.7359 | 2.7645 | 2.7969
Time 5.03 h, -20 celcius | 0.084287 | 0.084287 | 0.084287 | 0.084287 | 0.083003 | 0.084127 | 0.084469 | 0.085543 | 0.086743 | 0.087572 | 0.088047 | 0.088435 | 0.088561 | 0.088505 | 0.088289 | 0.08811 | 0.08822 | 0.08935 | 0.092855 | 0.10126 | 0.11875 | 0.15045 | 0.2025 | 0.28088 | 0.38519 | 0.51085 | 0.64628 | 0.77735 | 0.89303 | 0.98847 | 1.0619 | 1.1159 | 1.155 | 1.183 | 1.203 | 1.2167 | 1.2267 | 1.2373 | 1.2444 | 1.2521 | 1.2603 | 1.2702 | 1.2796 | 1.2944 | 1.3089 | 1.3285 | 1.3531 | 1.3812 | 1.4143 | 1.4539 | 1.499
Time 8.5 h, -40 celcius | 0.083695 | 0.083695 | 0.083695 | 0.083695 | 0.083695 | 0.08324 | 0.083622 | 0.083847 | 0.084834 | 0.086128 | 0.086714 | 0.087381 | 0.087886 | 0.088558 | 0.089009 | 0.089429 | 0.089952 | 0.090576 | 0.09152 | 0.092833 | 0.094685 | 0.097293 | 0.10083 | 0.10562 | 0.11226 | 0.12153 | 0.13502 | 0.15512 | 0.18502 | 0.22832 | 0.28819 | 0.36571 | 0.45912 | 0.56375 | 0.67242 | 0.77798 | 0.87404 | 0.95702 | 1.0284 | 1.0845 | 1.1268 | 1.1614 | 1.1873 | 1.2052 | 1.2167 | 1.2287 | 1.2359 | 1.2394 | 1.2468 | 1.2512 | 1.2552 | 1.2597
Std. : Water (1:0)_ 5 g_17C | 2.3304 | 2.4387 | 2.4947 | 2.5334 | 2.5734 | 2.595 | 2.6101 | 2.6237 | 2.6363 | 2.6458 | 2.6525 | 2.6605 | 2.6675 | 2.6754 | 2.6841 | 2.6936 | 2.7028 | 2.7121 | 2.7221 | 2.732 | 2.7416 | 2.7519 | 2.7627 | 2.773 | 2.7842 | 2.7949 | 2.8043 | 2.8148 | 2.8246 | 2.8354 | 2.8466 | 2.8598 | 2.8739 | 2.8899 | 2.908 | 2.9285 | 2.9493 | 2.972 | 2.9971 | 3.0212 | 3.0504 | 3.076 | 3.1069 | 3.1381 | 3.1714 | 3.2072 | 3.2459 | 3.2864
Std. : Water (1:3)_5 g_22.3C | 1.4671 | 1.6855 | 1.744 | 1.8824 | 1.9929 | 2.1292 | 2.3085 | 2.4363 | 2.5381 | 2.631 | 2.6959 | 2.7426 | 2.7767 | 2.8037 | 2.8219 | 2.8373 | 2.8512 | 2.8646 | 2.8761 | 2.8875 | 2.8997 | 2.9118 | 2.9235 | 2.9355 | 2.9478 | 2.9606 | 2.9734 | 2.9866 | 2.9991 | 3.0123 | 3.0249 | 3.039 | 3.0531 | 3.0685 | 3.0846 | 3.1036 | 3.1223 | 3.1442 | 3.1674 | 3.1919 | 3.2199 | 3.2461 | 3.2797 | 3.3073 | 3.3426 | 3.3775 | 3.4158 | 3.4569 | 3.5017 | 3.5489
Fitting-Time 0.23 h (20.3 C) Liquid | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.4119 | 2.4119 | 2.4119 | 2.4119 | 2.4118 | 2.4117 | 2.4115 | 2.4112 | 2.4108 | 2.4101 | 2.409 | 2.4073 | 2.4046 | 2.4004 | 2.3937 | 2.3832 | 2.367 | 2.342 | 2.3043 | 2.2484 | 2.1697 | 2.0642 | 1.9301 | 1.7781 | 1.6192 | 1.4715 | 1.3445 | 1.2461 | 1.1731 | 1.1243 | 1.0907 | 1.0692 | 1.0535 | 1.0451 | 1.0395 | 1.035 | 1.0329 | 1.0318 | 1.0305 | 1.0302 | 1.0299 | 1.0297
Fitting-Time 5.03 h (-20 C) Frozen | 1.2389 | 1.2389 | 1.2389 | 1.2389 | 1.2388 | 1.2387 | 1.2386 | 1.2384 | 1.2381 | 1.2377 | 1.2369 | 1.2358 | 1.2339 | 1.231 | 1.2264 | 1.2192 | 1.208 | 1.1906 | 1.1642 | 1.1248 | 1.0679 | 0.98926 | 0.88724 | 0.76529 | 0.63274 | 0.50242 | 0.38716 | 0.29408 | 0.22349 | 0.17372 | 0.13979 | 0.11692 | 0.10222 | 0.092682 | 0.086624 | 0.08272 | 0.080267 | 0.078693 | 0.077739 | 0.077121 | 0.076743 | 0.076474 | 0.076332 | 0.076237 | 0.076164 | 0.07613 | 0.076111 | 0.076089 | 0.076084 | 0.076079 | 0.076076
18 GHz
TVIS range
~ 80
~ 3
https://en.wikipedia.org/wiki/Permittivity
Real part Capacitance
Imaginary part Capacitance
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 0.23 h, 20.3 celcius | 0.021874 | 0.027746 | 0.035469 | 0.045629 | 0.058992 | 0.069416 | 0.087012 | 0.11672 | 0.14702 | 0.18167 | 0.23092 | 0.28348 | 0.3471 | 0.41787 | 0.49639 | 0.57052 | 0.6315 | 0.66609 | 0.66868 | 0.64015 | 0.58783 | 0.52385 | 0.45806 | 0.3962 | 0.3406 | 0.2936 | 0.25428 | 0.22269 | 0.19622 | 0.17597 | 0.15804 | 0.14279 | 0.13021 | 0.12053 | 0.11142 | 0.10593 | 0.099251 | 0.10112 | 0.096409 | 0.099903 | 0.099379 | 0.11049 | 0.12227 | 0.13167 | 0.14829 | 0.16415 | 0.18118 | 0.20571 | 0.22768 | 0.25157 | 0.27443
Time 5.03 h, -20 celcius | 0.0046903 | 0.0060276 | 0.0048957 | 0.0052689 | 0.0080706 | 0.0070995 | 0.0085621 | 0.011385 | 0.014441 | 0.017939 | 0.023068 | 0.028754 | 0.036271 | 0.045614 | 0.058069 | 0.073874 | 0.094576 | 0.12111 | 0.15588 | 0.20027 | 0.2566 | 0.32289 | 0.39693 | 0.47149 | 0.53464 | 0.57592 | 0.58749 | 0.56901 | 0.52696 | 0.47094 | 0.40966 | 0.35018 | 0.29633 | 0.24993 | 0.21217 | 0.18272 | 0.16189 | 0.14492 | 0.13366 | 0.12856 | 0.12752 | 0.12988 | 0.13813 | 0.14441 | 0.16577 | 0.18416 | 0.20379 | 0.22754 | 0.25211 | 0.27626 | 0.29768
Time 8.5 h, -40 celcius | 0.0037257 | 0.0037257 | 0.0022099 | 0.0055149 | 0.0027352 | 0.0046061 | 0.0028107 | 0.0031095 | 0.0040131 | 0.0050574 | 0.0061923 | 0.007861 | 0.0096387 | 0.01196 | 0.014763 | 0.018394 | 0.022849 | 0.028488 | 0.035374 | 0.043956 | 0.054453 | 0.067485 | 0.083211 | 0.10248 | 0.12627 | 0.15518 | 0.19059 | 0.23316 | 0.28257 | 0.33712 | 0.39303 | 0.44437 | 0.48479 | 0.5087 | 0.51347 | 0.50011 | 0.47159 | 0.43276 | 0.38638 | 0.3379 | 0.29199 | 0.24717 | 0.21022 | 0.17522 | 0.14921 | 0.12645 | 0.10764 | 0.094062 | 0.085294 | 0.077832 | 0.076166 | 0.074374
Std. : Water (1:0)_ 5 g_17C | 0.40745 | 0.37758 | 0.33167 | 0.27797 | 0.22443 | 0.19417 | 0.16169 | 0.13849 | 0.12143 | 0.10697 | 0.096405 | 0.087034 | 0.081871 | 0.078656 | 0.076272 | 0.075226 | 0.07448 | 0.073601 | 0.073296 | 0.073701 | 0.074654 | 0.075483 | 0.075353 | 0.07631 | 0.0769 | 0.077719 | 0.079173 | 0.081297 | 0.084848 | 0.090883 | 0.098449 | 0.10715 | 0.11892 | 0.13057 | 0.14131 | 0.15224 | 0.161 | 0.1718 | 0.18315 | 0.19373 | 0.20789 | 0.22266 | 0.23803 | 0.25543 | 0.2743 | 0.29629 | 0.32039 | 0.34659
Std. : Water (1:3)_5 g_22.3C | 0.35022 | 0.37378 | 0.41173 | 0.55328 | 0.59494 | 0.61621 | 0.60526 | 0.57334 | 0.52449 | 0.45124 | 0.3927 | 0.33463 | 0.28521 | 0.24069 | 0.2067 | 0.17971 | 0.15895 | 0.14354 | 0.13098 | 0.1209 | 0.11373 | 0.10865 | 0.10508 | 0.10263 | 0.10025 | 0.099388 | 0.099009 | 0.09943 | 0.10029 | 0.10249 | 0.10653 | 0.11105 | 0.11816 | 0.12566 | 0.13766 | 0.14944 | 0.16023 | 0.17158 | 0.18231 | 0.19465 | 0.20803 | 0.22162 | 0.23412 | 0.25639 | 0.27568 | 0.29801 | 0.32242 | 0.35086 | 0.38221 | 0.41669
Fitting-Time 0.23 h (20.3 C) Liquid | 0.00080148 | 0.001009 | 0.0012703 | 0.0015992 | 0.0020132 | 0.0025345 | 0.0031908 | 0.0040169 | 0.005057 | 0.0063663 | 0.0080146 | 0.010089 | 0.012702 | 0.01599 | 0.020128 | 0.025337 | 0.031892 | 0.040136 | 0.050507 | 0.063538 | 0.07988 | 0.10037 | 0.126 | 0.15786 | 0.1972 | 0.24538 | 0.30315 | 0.37047 | 0.44655 | 0.52567 | 0.59992 | 0.65886 | 0.68897 | 0.68388 | 0.64493 | 0.58011 | 0.50283 | 0.42213 | 0.34988 | 0.28493 | 0.23165 | 0.18163 | 0.14738 | 0.11839 | 0.089082 | 0.071374 | 0.059526 | 0.041708 | 0.035758 | 0.029805 | 0.023848
Fitting-Time 5.03 h (-20 C) Frozen | 0.0048296 | 0.0060799 | 0.0076542 | 0.0096359 | 0.01213 | 0.01527 | 0.019222 | 0.024195 | 0.030452 | 0.038322 | 0.048213 | 0.060635 | 0.076215 | 0.095709 | 0.12001 | 0.15014 | 0.18717 | 0.232 | 0.28517 | 0.34602 | 0.41191 | 0.4775 | 0.53412 | 0.57135 | 0.58091 | 0.56037 | 0.51476 | 0.45386 | 0.38691 | 0.32252 | 0.26465 | 0.2141 | 0.17243 | 0.138 | 0.11029 | 0.087693 | 0.069745 | 0.055173 | 0.044035 | 0.034967 | 0.027983 | 0.021692 | 0.017495 | 0.013998 | 0.010499 | 0.0083994 | 0.0069995 | 0.0048997 | 0.0041998 | 0.0034999 | 0.0027999
(I)
Notes: D:\PhD\Assignment\Conference ppt\Plot used in smi 2019 Conference_v1

## Slide 14: Dielectric Loss Mechanisms
Maxwell-Wagner (MW) polarization of the glass wall of the TVIS vial at +20 ˚C, with a dielectric loss peak frequency of 17.8 kHz
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 0.23 h, 20.3 celcius | 1.0162 | 1.0162 | 1.0162 | 1.0162 | 1.0176 | 1.024 | 1.0293 | 1.0371 | 1.0465 | 1.0575 | 1.0742 | 1.0971 | 1.131 | 1.18 | 1.2538 | 1.3532 | 1.4803 | 1.6281 | 1.7827 | 1.9288 | 2.0567 | 2.1587 | 2.2374 | 2.2971 | 2.3414 | 2.3756 | 2.4038 | 2.4272 | 2.4476 | 2.4662 | 2.4852 | 2.5003 | 2.5152 | 2.5286 | 2.5419 | 2.5531 | 2.5615 | 2.5717 | 2.5802 | 2.589 | 2.5992 | 2.6061 | 2.6201 | 2.625 | 2.6478 | 2.6645 | 2.686 | 2.708 | 2.7359 | 2.7645 | 2.7969
Time 5.03 h, -20 celcius | 0.084287 | 0.084287 | 0.084287 | 0.084287 | 0.083003 | 0.084127 | 0.084469 | 0.085543 | 0.086743 | 0.087572 | 0.088047 | 0.088435 | 0.088561 | 0.088505 | 0.088289 | 0.08811 | 0.08822 | 0.08935 | 0.092855 | 0.10126 | 0.11875 | 0.15045 | 0.2025 | 0.28088 | 0.38519 | 0.51085 | 0.64628 | 0.77735 | 0.89303 | 0.98847 | 1.0619 | 1.1159 | 1.155 | 1.183 | 1.203 | 1.2167 | 1.2267 | 1.2373 | 1.2444 | 1.2521 | 1.2603 | 1.2702 | 1.2796 | 1.2944 | 1.3089 | 1.3285 | 1.3531 | 1.3812 | 1.4143 | 1.4539 | 1.499
Time 8.5 h, -40 celcius | 0.083695 | 0.083695 | 0.083695 | 0.083695 | 0.083695 | 0.08324 | 0.083622 | 0.083847 | 0.084834 | 0.086128 | 0.086714 | 0.087381 | 0.087886 | 0.088558 | 0.089009 | 0.089429 | 0.089952 | 0.090576 | 0.09152 | 0.092833 | 0.094685 | 0.097293 | 0.10083 | 0.10562 | 0.11226 | 0.12153 | 0.13502 | 0.15512 | 0.18502 | 0.22832 | 0.28819 | 0.36571 | 0.45912 | 0.56375 | 0.67242 | 0.77798 | 0.87404 | 0.95702 | 1.0284 | 1.0845 | 1.1268 | 1.1614 | 1.1873 | 1.2052 | 1.2167 | 1.2287 | 1.2359 | 1.2394 | 1.2468 | 1.2512 | 1.2552 | 1.2597
Std. : Water (1:0)_ 5 g_17C | 2.3304 | 2.4387 | 2.4947 | 2.5334 | 2.5734 | 2.595 | 2.6101 | 2.6237 | 2.6363 | 2.6458 | 2.6525 | 2.6605 | 2.6675 | 2.6754 | 2.6841 | 2.6936 | 2.7028 | 2.7121 | 2.7221 | 2.732 | 2.7416 | 2.7519 | 2.7627 | 2.773 | 2.7842 | 2.7949 | 2.8043 | 2.8148 | 2.8246 | 2.8354 | 2.8466 | 2.8598 | 2.8739 | 2.8899 | 2.908 | 2.9285 | 2.9493 | 2.972 | 2.9971 | 3.0212 | 3.0504 | 3.076 | 3.1069 | 3.1381 | 3.1714 | 3.2072 | 3.2459 | 3.2864
Std. : Water (1:3)_5 g_22.3C | 1.4671 | 1.6855 | 1.744 | 1.8824 | 1.9929 | 2.1292 | 2.3085 | 2.4363 | 2.5381 | 2.631 | 2.6959 | 2.7426 | 2.7767 | 2.8037 | 2.8219 | 2.8373 | 2.8512 | 2.8646 | 2.8761 | 2.8875 | 2.8997 | 2.9118 | 2.9235 | 2.9355 | 2.9478 | 2.9606 | 2.9734 | 2.9866 | 2.9991 | 3.0123 | 3.0249 | 3.039 | 3.0531 | 3.0685 | 3.0846 | 3.1036 | 3.1223 | 3.1442 | 3.1674 | 3.1919 | 3.2199 | 3.2461 | 3.2797 | 3.3073 | 3.3426 | 3.3775 | 3.4158 | 3.4569 | 3.5017 | 3.5489
Fitting-Time 0.23 h (20.3 C) Liquid | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.4119 | 2.4119 | 2.4119 | 2.4119 | 2.4118 | 2.4117 | 2.4115 | 2.4112 | 2.4108 | 2.4101 | 2.409 | 2.4073 | 2.4046 | 2.4004 | 2.3937 | 2.3832 | 2.367 | 2.342 | 2.3043 | 2.2484 | 2.1697 | 2.0642 | 1.9301 | 1.7781 | 1.6192 | 1.4715 | 1.3445 | 1.2461 | 1.1731 | 1.1243 | 1.0907 | 1.0692 | 1.0535 | 1.0451 | 1.0395 | 1.035 | 1.0329 | 1.0318 | 1.0305 | 1.0302 | 1.0299 | 1.0297
Fitting-Time 5.03 h (-20 C) Frozen | 1.2389 | 1.2389 | 1.2389 | 1.2389 | 1.2388 | 1.2387 | 1.2386 | 1.2384 | 1.2381 | 1.2377 | 1.2369 | 1.2358 | 1.2339 | 1.231 | 1.2264 | 1.2192 | 1.208 | 1.1906 | 1.1642 | 1.1248 | 1.0679 | 0.98926 | 0.88724 | 0.76529 | 0.63274 | 0.50242 | 0.38716 | 0.29408 | 0.22349 | 0.17372 | 0.13979 | 0.11692 | 0.10222 | 0.092682 | 0.086624 | 0.08272 | 0.080267 | 0.078693 | 0.077739 | 0.077121 | 0.076743 | 0.076474 | 0.076332 | 0.076237 | 0.076164 | 0.07613 | 0.076111 | 0.076089 | 0.076084 | 0.076079 | 0.076076
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 0.23 h, 20.3 celcius | 0.021874 | 0.027746 | 0.035469 | 0.045629 | 0.058992 | 0.069416 | 0.087012 | 0.11672 | 0.14702 | 0.18167 | 0.23092 | 0.28348 | 0.3471 | 0.41787 | 0.49639 | 0.57052 | 0.6315 | 0.66609 | 0.66868 | 0.64015 | 0.58783 | 0.52385 | 0.45806 | 0.3962 | 0.3406 | 0.2936 | 0.25428 | 0.22269 | 0.19622 | 0.17597 | 0.15804 | 0.14279 | 0.13021 | 0.12053 | 0.11142 | 0.10593 | 0.099251 | 0.10112 | 0.096409 | 0.099903 | 0.099379 | 0.11049 | 0.12227 | 0.13167 | 0.14829 | 0.16415 | 0.18118 | 0.20571 | 0.22768 | 0.25157 | 0.27443
Time 5.03 h, -20 celcius | 0.0046903 | 0.0060276 | 0.0048957 | 0.0052689 | 0.0080706 | 0.0070995 | 0.0085621 | 0.011385 | 0.014441 | 0.017939 | 0.023068 | 0.028754 | 0.036271 | 0.045614 | 0.058069 | 0.073874 | 0.094576 | 0.12111 | 0.15588 | 0.20027 | 0.2566 | 0.32289 | 0.39693 | 0.47149 | 0.53464 | 0.57592 | 0.58749 | 0.56901 | 0.52696 | 0.47094 | 0.40966 | 0.35018 | 0.29633 | 0.24993 | 0.21217 | 0.18272 | 0.16189 | 0.14492 | 0.13366 | 0.12856 | 0.12752 | 0.12988 | 0.13813 | 0.14441 | 0.16577 | 0.18416 | 0.20379 | 0.22754 | 0.25211 | 0.27626 | 0.29768
Time 8.5 h, -40 celcius | 0.0037257 | 0.0037257 | 0.0022099 | 0.0055149 | 0.0027352 | 0.0046061 | 0.0028107 | 0.0031095 | 0.0040131 | 0.0050574 | 0.0061923 | 0.007861 | 0.0096387 | 0.01196 | 0.014763 | 0.018394 | 0.022849 | 0.028488 | 0.035374 | 0.043956 | 0.054453 | 0.067485 | 0.083211 | 0.10248 | 0.12627 | 0.15518 | 0.19059 | 0.23316 | 0.28257 | 0.33712 | 0.39303 | 0.44437 | 0.48479 | 0.5087 | 0.51347 | 0.50011 | 0.47159 | 0.43276 | 0.38638 | 0.3379 | 0.29199 | 0.24717 | 0.21022 | 0.17522 | 0.14921 | 0.12645 | 0.10764 | 0.094062 | 0.085294 | 0.077832 | 0.076166 | 0.074374
Std. : Water (1:0)_ 5 g_17C | 0.40745 | 0.37758 | 0.33167 | 0.27797 | 0.22443 | 0.19417 | 0.16169 | 0.13849 | 0.12143 | 0.10697 | 0.096405 | 0.087034 | 0.081871 | 0.078656 | 0.076272 | 0.075226 | 0.07448 | 0.073601 | 0.073296 | 0.073701 | 0.074654 | 0.075483 | 0.075353 | 0.07631 | 0.0769 | 0.077719 | 0.079173 | 0.081297 | 0.084848 | 0.090883 | 0.098449 | 0.10715 | 0.11892 | 0.13057 | 0.14131 | 0.15224 | 0.161 | 0.1718 | 0.18315 | 0.19373 | 0.20789 | 0.22266 | 0.23803 | 0.25543 | 0.2743 | 0.29629 | 0.32039 | 0.34659
Std. : Water (1:3)_5 g_22.3C | 0.35022 | 0.37378 | 0.41173 | 0.55328 | 0.59494 | 0.61621 | 0.60526 | 0.57334 | 0.52449 | 0.45124 | 0.3927 | 0.33463 | 0.28521 | 0.24069 | 0.2067 | 0.17971 | 0.15895 | 0.14354 | 0.13098 | 0.1209 | 0.11373 | 0.10865 | 0.10508 | 0.10263 | 0.10025 | 0.099388 | 0.099009 | 0.09943 | 0.10029 | 0.10249 | 0.10653 | 0.11105 | 0.11816 | 0.12566 | 0.13766 | 0.14944 | 0.16023 | 0.17158 | 0.18231 | 0.19465 | 0.20803 | 0.22162 | 0.23412 | 0.25639 | 0.27568 | 0.29801 | 0.32242 | 0.35086 | 0.38221 | 0.41669
Fitting-Time 0.23 h (20.3 C) Liquid | 0.00080148 | 0.001009 | 0.0012703 | 0.0015992 | 0.0020132 | 0.0025345 | 0.0031908 | 0.0040169 | 0.005057 | 0.0063663 | 0.0080146 | 0.010089 | 0.012702 | 0.01599 | 0.020128 | 0.025337 | 0.031892 | 0.040136 | 0.050507 | 0.063538 | 0.07988 | 0.10037 | 0.126 | 0.15786 | 0.1972 | 0.24538 | 0.30315 | 0.37047 | 0.44655 | 0.52567 | 0.59992 | 0.65886 | 0.68897 | 0.68388 | 0.64493 | 0.58011 | 0.50283 | 0.42213 | 0.34988 | 0.28493 | 0.23165 | 0.18163 | 0.14738 | 0.11839 | 0.089082 | 0.071374 | 0.059526 | 0.041708 | 0.035758 | 0.029805 | 0.023848
Fitting-Time 5.03 h (-20 C) Frozen | 0.0048296 | 0.0060799 | 0.0076542 | 0.0096359 | 0.01213 | 0.01527 | 0.019222 | 0.024195 | 0.030452 | 0.038322 | 0.048213 | 0.060635 | 0.076215 | 0.095709 | 0.12001 | 0.15014 | 0.18717 | 0.232 | 0.28517 | 0.34602 | 0.41191 | 0.4775 | 0.53412 | 0.57135 | 0.58091 | 0.56037 | 0.51476 | 0.45386 | 0.38691 | 0.32252 | 0.26465 | 0.2141 | 0.17243 | 0.138 | 0.11029 | 0.087693 | 0.069745 | 0.055173 | 0.044035 | 0.034967 | 0.027983 | 0.021692 | 0.017495 | 0.013998 | 0.010499 | 0.0083994 | 0.0069995 | 0.0048997 | 0.0041998 | 0.0034999 | 0.0027999
18 GHz
20.3 °C
18 kHz
Real part Capacitance
Imaginary part Capacitance
Measurement vial
+
-
-
+
+
+
+
+
+
+
+
+
+
+
-
-
-
-
-
-
-
-
-
-
+
+
+
+
+
+
+
+
+
+
-
-
-
-
-
-
-
-
-
-
Water
Glass wall
Electronic polarization
Atomic polarization
o
H
H
δ-
δ+
δ+
Dipolar polarization
Grotthus mechanism
o
H
H
δ-
δ+
δ+
o
H
H
δ-
δ+
δ+
H
o
H
H
o
H
H
o
H
H+
Ionic diffusion
H+
HCO3−
Na+
Cl−
H+
H+
OH-
OH-
H+
(I)
(II)
20.3 °C
Notes: D:\PhD\Assignment\Conference ppt\Plot used in smi 2019 Conference_v1

## Slide 15: Dielectric Loss Mechanisms
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 0.23 h, 20.3 celcius | 1.0162 | 1.0162 | 1.0162 | 1.0162 | 1.0176 | 1.024 | 1.0293 | 1.0371 | 1.0465 | 1.0575 | 1.0742 | 1.0971 | 1.131 | 1.18 | 1.2538 | 1.3532 | 1.4803 | 1.6281 | 1.7827 | 1.9288 | 2.0567 | 2.1587 | 2.2374 | 2.2971 | 2.3414 | 2.3756 | 2.4038 | 2.4272 | 2.4476 | 2.4662 | 2.4852 | 2.5003 | 2.5152 | 2.5286 | 2.5419 | 2.5531 | 2.5615 | 2.5717 | 2.5802 | 2.589 | 2.5992 | 2.6061 | 2.6201 | 2.625 | 2.6478 | 2.6645 | 2.686 | 2.708 | 2.7359 | 2.7645 | 2.7969
Time 5.03 h, -20 celcius | 0.084287 | 0.084287 | 0.084287 | 0.084287 | 0.083003 | 0.084127 | 0.084469 | 0.085543 | 0.086743 | 0.087572 | 0.088047 | 0.088435 | 0.088561 | 0.088505 | 0.088289 | 0.08811 | 0.08822 | 0.08935 | 0.092855 | 0.10126 | 0.11875 | 0.15045 | 0.2025 | 0.28088 | 0.38519 | 0.51085 | 0.64628 | 0.77735 | 0.89303 | 0.98847 | 1.0619 | 1.1159 | 1.155 | 1.183 | 1.203 | 1.2167 | 1.2267 | 1.2373 | 1.2444 | 1.2521 | 1.2603 | 1.2702 | 1.2796 | 1.2944 | 1.3089 | 1.3285 | 1.3531 | 1.3812 | 1.4143 | 1.4539 | 1.499
Time 8.5 h, -40 celcius | 0.083695 | 0.083695 | 0.083695 | 0.083695 | 0.083695 | 0.08324 | 0.083622 | 0.083847 | 0.084834 | 0.086128 | 0.086714 | 0.087381 | 0.087886 | 0.088558 | 0.089009 | 0.089429 | 0.089952 | 0.090576 | 0.09152 | 0.092833 | 0.094685 | 0.097293 | 0.10083 | 0.10562 | 0.11226 | 0.12153 | 0.13502 | 0.15512 | 0.18502 | 0.22832 | 0.28819 | 0.36571 | 0.45912 | 0.56375 | 0.67242 | 0.77798 | 0.87404 | 0.95702 | 1.0284 | 1.0845 | 1.1268 | 1.1614 | 1.1873 | 1.2052 | 1.2167 | 1.2287 | 1.2359 | 1.2394 | 1.2468 | 1.2512 | 1.2552 | 1.2597
Std. : Water (1:0)_ 5 g_17C | 2.3304 | 2.4387 | 2.4947 | 2.5334 | 2.5734 | 2.595 | 2.6101 | 2.6237 | 2.6363 | 2.6458 | 2.6525 | 2.6605 | 2.6675 | 2.6754 | 2.6841 | 2.6936 | 2.7028 | 2.7121 | 2.7221 | 2.732 | 2.7416 | 2.7519 | 2.7627 | 2.773 | 2.7842 | 2.7949 | 2.8043 | 2.8148 | 2.8246 | 2.8354 | 2.8466 | 2.8598 | 2.8739 | 2.8899 | 2.908 | 2.9285 | 2.9493 | 2.972 | 2.9971 | 3.0212 | 3.0504 | 3.076 | 3.1069 | 3.1381 | 3.1714 | 3.2072 | 3.2459 | 3.2864
Std. : Water (1:3)_5 g_22.3C | 1.4671 | 1.6855 | 1.744 | 1.8824 | 1.9929 | 2.1292 | 2.3085 | 2.4363 | 2.5381 | 2.631 | 2.6959 | 2.7426 | 2.7767 | 2.8037 | 2.8219 | 2.8373 | 2.8512 | 2.8646 | 2.8761 | 2.8875 | 2.8997 | 2.9118 | 2.9235 | 2.9355 | 2.9478 | 2.9606 | 2.9734 | 2.9866 | 2.9991 | 3.0123 | 3.0249 | 3.039 | 3.0531 | 3.0685 | 3.0846 | 3.1036 | 3.1223 | 3.1442 | 3.1674 | 3.1919 | 3.2199 | 3.2461 | 3.2797 | 3.3073 | 3.3426 | 3.3775 | 3.4158 | 3.4569 | 3.5017 | 3.5489
Fitting-Time 0.23 h (20.3 C) Liquid | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.4119 | 2.4119 | 2.4119 | 2.4119 | 2.4118 | 2.4117 | 2.4115 | 2.4112 | 2.4108 | 2.4101 | 2.409 | 2.4073 | 2.4046 | 2.4004 | 2.3937 | 2.3832 | 2.367 | 2.342 | 2.3043 | 2.2484 | 2.1697 | 2.0642 | 1.9301 | 1.7781 | 1.6192 | 1.4715 | 1.3445 | 1.2461 | 1.1731 | 1.1243 | 1.0907 | 1.0692 | 1.0535 | 1.0451 | 1.0395 | 1.035 | 1.0329 | 1.0318 | 1.0305 | 1.0302 | 1.0299 | 1.0297
Fitting-Time 5.03 h (-20 C) Frozen | 1.2389 | 1.2389 | 1.2389 | 1.2389 | 1.2388 | 1.2387 | 1.2386 | 1.2384 | 1.2381 | 1.2377 | 1.2369 | 1.2358 | 1.2339 | 1.231 | 1.2264 | 1.2192 | 1.208 | 1.1906 | 1.1642 | 1.1248 | 1.0679 | 0.98926 | 0.88724 | 0.76529 | 0.63274 | 0.50242 | 0.38716 | 0.29408 | 0.22349 | 0.17372 | 0.13979 | 0.11692 | 0.10222 | 0.092682 | 0.086624 | 0.08272 | 0.080267 | 0.078693 | 0.077739 | 0.077121 | 0.076743 | 0.076474 | 0.076332 | 0.076237 | 0.076164 | 0.07613 | 0.076111 | 0.076089 | 0.076084 | 0.076079 | 0.076076
Maxwell-Wagner (MW) polarization of the glass wall of the TVIS vial at +20 ˚C, with a dielectric loss peak frequency of 17.8 kHz
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 0.23 h, 20.3 celcius | 0.021874 | 0.027746 | 0.035469 | 0.045629 | 0.058992 | 0.069416 | 0.087012 | 0.11672 | 0.14702 | 0.18167 | 0.23092 | 0.28348 | 0.3471 | 0.41787 | 0.49639 | 0.57052 | 0.6315 | 0.66609 | 0.66868 | 0.64015 | 0.58783 | 0.52385 | 0.45806 | 0.3962 | 0.3406 | 0.2936 | 0.25428 | 0.22269 | 0.19622 | 0.17597 | 0.15804 | 0.14279 | 0.13021 | 0.12053 | 0.11142 | 0.10593 | 0.099251 | 0.10112 | 0.096409 | 0.099903 | 0.099379 | 0.11049 | 0.12227 | 0.13167 | 0.14829 | 0.16415 | 0.18118 | 0.20571 | 0.22768 | 0.25157 | 0.27443
Time 5.03 h, -20 celcius | 0.0046903 | 0.0060276 | 0.0048957 | 0.0052689 | 0.0080706 | 0.0070995 | 0.0085621 | 0.011385 | 0.014441 | 0.017939 | 0.023068 | 0.028754 | 0.036271 | 0.045614 | 0.058069 | 0.073874 | 0.094576 | 0.12111 | 0.15588 | 0.20027 | 0.2566 | 0.32289 | 0.39693 | 0.47149 | 0.53464 | 0.57592 | 0.58749 | 0.56901 | 0.52696 | 0.47094 | 0.40966 | 0.35018 | 0.29633 | 0.24993 | 0.21217 | 0.18272 | 0.16189 | 0.14492 | 0.13366 | 0.12856 | 0.12752 | 0.12988 | 0.13813 | 0.14441 | 0.16577 | 0.18416 | 0.20379 | 0.22754 | 0.25211 | 0.27626 | 0.29768
Time 8.5 h, -40 celcius | 0.0037257 | 0.0037257 | 0.0022099 | 0.0055149 | 0.0027352 | 0.0046061 | 0.0028107 | 0.0031095 | 0.0040131 | 0.0050574 | 0.0061923 | 0.007861 | 0.0096387 | 0.01196 | 0.014763 | 0.018394 | 0.022849 | 0.028488 | 0.035374 | 0.043956 | 0.054453 | 0.067485 | 0.083211 | 0.10248 | 0.12627 | 0.15518 | 0.19059 | 0.23316 | 0.28257 | 0.33712 | 0.39303 | 0.44437 | 0.48479 | 0.5087 | 0.51347 | 0.50011 | 0.47159 | 0.43276 | 0.38638 | 0.3379 | 0.29199 | 0.24717 | 0.21022 | 0.17522 | 0.14921 | 0.12645 | 0.10764 | 0.094062 | 0.085294 | 0.077832 | 0.076166 | 0.074374
Std. : Water (1:0)_ 5 g_17C | 0.40745 | 0.37758 | 0.33167 | 0.27797 | 0.22443 | 0.19417 | 0.16169 | 0.13849 | 0.12143 | 0.10697 | 0.096405 | 0.087034 | 0.081871 | 0.078656 | 0.076272 | 0.075226 | 0.07448 | 0.073601 | 0.073296 | 0.073701 | 0.074654 | 0.075483 | 0.075353 | 0.07631 | 0.0769 | 0.077719 | 0.079173 | 0.081297 | 0.084848 | 0.090883 | 0.098449 | 0.10715 | 0.11892 | 0.13057 | 0.14131 | 0.15224 | 0.161 | 0.1718 | 0.18315 | 0.19373 | 0.20789 | 0.22266 | 0.23803 | 0.25543 | 0.2743 | 0.29629 | 0.32039 | 0.34659
Std. : Water (1:3)_5 g_22.3C | 0.35022 | 0.37378 | 0.41173 | 0.55328 | 0.59494 | 0.61621 | 0.60526 | 0.57334 | 0.52449 | 0.45124 | 0.3927 | 0.33463 | 0.28521 | 0.24069 | 0.2067 | 0.17971 | 0.15895 | 0.14354 | 0.13098 | 0.1209 | 0.11373 | 0.10865 | 0.10508 | 0.10263 | 0.10025 | 0.099388 | 0.099009 | 0.09943 | 0.10029 | 0.10249 | 0.10653 | 0.11105 | 0.11816 | 0.12566 | 0.13766 | 0.14944 | 0.16023 | 0.17158 | 0.18231 | 0.19465 | 0.20803 | 0.22162 | 0.23412 | 0.25639 | 0.27568 | 0.29801 | 0.32242 | 0.35086 | 0.38221 | 0.41669
Fitting-Time 0.23 h (20.3 C) Liquid | 0.00080148 | 0.001009 | 0.0012703 | 0.0015992 | 0.0020132 | 0.0025345 | 0.0031908 | 0.0040169 | 0.005057 | 0.0063663 | 0.0080146 | 0.010089 | 0.012702 | 0.01599 | 0.020128 | 0.025337 | 0.031892 | 0.040136 | 0.050507 | 0.063538 | 0.07988 | 0.10037 | 0.126 | 0.15786 | 0.1972 | 0.24538 | 0.30315 | 0.37047 | 0.44655 | 0.52567 | 0.59992 | 0.65886 | 0.68897 | 0.68388 | 0.64493 | 0.58011 | 0.50283 | 0.42213 | 0.34988 | 0.28493 | 0.23165 | 0.18163 | 0.14738 | 0.11839 | 0.089082 | 0.071374 | 0.059526 | 0.041708 | 0.035758 | 0.029805 | 0.023848
Fitting-Time 5.03 h (-20 C) Frozen | 0.0048296 | 0.0060799 | 0.0076542 | 0.0096359 | 0.01213 | 0.01527 | 0.019222 | 0.024195 | 0.030452 | 0.038322 | 0.048213 | 0.060635 | 0.076215 | 0.095709 | 0.12001 | 0.15014 | 0.18717 | 0.232 | 0.28517 | 0.34602 | 0.41191 | 0.4775 | 0.53412 | 0.57135 | 0.58091 | 0.56037 | 0.51476 | 0.45386 | 0.38691 | 0.32252 | 0.26465 | 0.2141 | 0.17243 | 0.138 | 0.11029 | 0.087693 | 0.069745 | 0.055173 | 0.044035 | 0.034967 | 0.027983 | 0.021692 | 0.017495 | 0.013998 | 0.010499 | 0.0083994 | 0.0069995 | 0.0048997 | 0.0041998 | 0.0034999 | 0.0027999
(I)
18 GHz
20.3 °C
20.3 °C
18 kHz
0.0056 M KCl
(II)
0.0056 M KCl
TVIS range
Real part Capacitance
Imaginary part Capacitance
Measurement vial
+
-
-
+
+
+
+
+
+
+
+
+
+
+
-
-
-
-
-
-
-
-
-
-
+
+
+
+
+
+
+
+
+
+
-
-
-
-
-
-
-
-
-
-
Water
Glass wall
Electronic polarization
Atomic polarization
o
H
H
δ-
δ+
δ+
Dipolar polarization
Grotthus mechanism
o
H
H
δ-
δ+
δ+
o
H
H
δ-
δ+
δ+
H
o
H
H
o
H
H
o
H
H+
Ionic diffusion
H+
HCO3−
Na+
Cl−
H+
H+
OH-
H+
K+
K+
Cl−
Notes: D:\PhD\Assignment\Conference ppt\Plot used in smi 2019 Conference_v1

## Slide 16: Dielectric Loss Mechanisms
The dielectric polarization of ice at −20 ˚C, with a dielectric loss peak frequencies of 2.57 kHz
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 0.23 h, 20.3 celcius | 1.0162 | 1.0162 | 1.0162 | 1.0162 | 1.0176 | 1.024 | 1.0293 | 1.0371 | 1.0465 | 1.0575 | 1.0742 | 1.0971 | 1.131 | 1.18 | 1.2538 | 1.3532 | 1.4803 | 1.6281 | 1.7827 | 1.9288 | 2.0567 | 2.1587 | 2.2374 | 2.2971 | 2.3414 | 2.3756 | 2.4038 | 2.4272 | 2.4476 | 2.4662 | 2.4852 | 2.5003 | 2.5152 | 2.5286 | 2.5419 | 2.5531 | 2.5615 | 2.5717 | 2.5802 | 2.589 | 2.5992 | 2.6061 | 2.6201 | 2.625 | 2.6478 | 2.6645 | 2.686 | 2.708 | 2.7359 | 2.7645 | 2.7969
Time 5.03 h, -20 celcius | 0.084287 | 0.084287 | 0.084287 | 0.084287 | 0.083003 | 0.084127 | 0.084469 | 0.085543 | 0.086743 | 0.087572 | 0.088047 | 0.088435 | 0.088561 | 0.088505 | 0.088289 | 0.08811 | 0.08822 | 0.08935 | 0.092855 | 0.10126 | 0.11875 | 0.15045 | 0.2025 | 0.28088 | 0.38519 | 0.51085 | 0.64628 | 0.77735 | 0.89303 | 0.98847 | 1.0619 | 1.1159 | 1.155 | 1.183 | 1.203 | 1.2167 | 1.2267 | 1.2373 | 1.2444 | 1.2521 | 1.2603 | 1.2702 | 1.2796 | 1.2944 | 1.3089 | 1.3285 | 1.3531 | 1.3812 | 1.4143 | 1.4539 | 1.499
Time 8.5 h, -40 celcius | 0.083695 | 0.083695 | 0.083695 | 0.083695 | 0.083695 | 0.08324 | 0.083622 | 0.083847 | 0.084834 | 0.086128 | 0.086714 | 0.087381 | 0.087886 | 0.088558 | 0.089009 | 0.089429 | 0.089952 | 0.090576 | 0.09152 | 0.092833 | 0.094685 | 0.097293 | 0.10083 | 0.10562 | 0.11226 | 0.12153 | 0.13502 | 0.15512 | 0.18502 | 0.22832 | 0.28819 | 0.36571 | 0.45912 | 0.56375 | 0.67242 | 0.77798 | 0.87404 | 0.95702 | 1.0284 | 1.0845 | 1.1268 | 1.1614 | 1.1873 | 1.2052 | 1.2167 | 1.2287 | 1.2359 | 1.2394 | 1.2468 | 1.2512 | 1.2552 | 1.2597
Std. : Water (1:0)_ 5 g_17C | 2.3304 | 2.4387 | 2.4947 | 2.5334 | 2.5734 | 2.595 | 2.6101 | 2.6237 | 2.6363 | 2.6458 | 2.6525 | 2.6605 | 2.6675 | 2.6754 | 2.6841 | 2.6936 | 2.7028 | 2.7121 | 2.7221 | 2.732 | 2.7416 | 2.7519 | 2.7627 | 2.773 | 2.7842 | 2.7949 | 2.8043 | 2.8148 | 2.8246 | 2.8354 | 2.8466 | 2.8598 | 2.8739 | 2.8899 | 2.908 | 2.9285 | 2.9493 | 2.972 | 2.9971 | 3.0212 | 3.0504 | 3.076 | 3.1069 | 3.1381 | 3.1714 | 3.2072 | 3.2459 | 3.2864
Std. : Water (1:3)_5 g_22.3C | 1.4671 | 1.6855 | 1.744 | 1.8824 | 1.9929 | 2.1292 | 2.3085 | 2.4363 | 2.5381 | 2.631 | 2.6959 | 2.7426 | 2.7767 | 2.8037 | 2.8219 | 2.8373 | 2.8512 | 2.8646 | 2.8761 | 2.8875 | 2.8997 | 2.9118 | 2.9235 | 2.9355 | 2.9478 | 2.9606 | 2.9734 | 2.9866 | 2.9991 | 3.0123 | 3.0249 | 3.039 | 3.0531 | 3.0685 | 3.0846 | 3.1036 | 3.1223 | 3.1442 | 3.1674 | 3.1919 | 3.2199 | 3.2461 | 3.2797 | 3.3073 | 3.3426 | 3.3775 | 3.4158 | 3.4569 | 3.5017 | 3.5489
Fitting-Time 0.23 h (20.3 C) Liquid | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.4119 | 2.4119 | 2.4119 | 2.4119 | 2.4118 | 2.4117 | 2.4115 | 2.4112 | 2.4108 | 2.4101 | 2.409 | 2.4073 | 2.4046 | 2.4004 | 2.3937 | 2.3832 | 2.367 | 2.342 | 2.3043 | 2.2484 | 2.1697 | 2.0642 | 1.9301 | 1.7781 | 1.6192 | 1.4715 | 1.3445 | 1.2461 | 1.1731 | 1.1243 | 1.0907 | 1.0692 | 1.0535 | 1.0451 | 1.0395 | 1.035 | 1.0329 | 1.0318 | 1.0305 | 1.0302 | 1.0299 | 1.0297
Fitting-Time 5.03 h (-20 C) Frozen | 1.2389 | 1.2389 | 1.2389 | 1.2389 | 1.2388 | 1.2387 | 1.2386 | 1.2384 | 1.2381 | 1.2377 | 1.2369 | 1.2358 | 1.2339 | 1.231 | 1.2264 | 1.2192 | 1.208 | 1.1906 | 1.1642 | 1.1248 | 1.0679 | 0.98926 | 0.88724 | 0.76529 | 0.63274 | 0.50242 | 0.38716 | 0.29408 | 0.22349 | 0.17372 | 0.13979 | 0.11692 | 0.10222 | 0.092682 | 0.086624 | 0.08272 | 0.080267 | 0.078693 | 0.077739 | 0.077121 | 0.076743 | 0.076474 | 0.076332 | 0.076237 | 0.076164 | 0.07613 | 0.076111 | 0.076089 | 0.076084 | 0.076079 | 0.076076
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 0.23 h, 20.3 celcius | 0.021874 | 0.027746 | 0.035469 | 0.045629 | 0.058992 | 0.069416 | 0.087012 | 0.11672 | 0.14702 | 0.18167 | 0.23092 | 0.28348 | 0.3471 | 0.41787 | 0.49639 | 0.57052 | 0.6315 | 0.66609 | 0.66868 | 0.64015 | 0.58783 | 0.52385 | 0.45806 | 0.3962 | 0.3406 | 0.2936 | 0.25428 | 0.22269 | 0.19622 | 0.17597 | 0.15804 | 0.14279 | 0.13021 | 0.12053 | 0.11142 | 0.10593 | 0.099251 | 0.10112 | 0.096409 | 0.099903 | 0.099379 | 0.11049 | 0.12227 | 0.13167 | 0.14829 | 0.16415 | 0.18118 | 0.20571 | 0.22768 | 0.25157 | 0.27443
Time 5.03 h, -20 celcius | 0.0046903 | 0.0060276 | 0.0048957 | 0.0052689 | 0.0080706 | 0.0070995 | 0.0085621 | 0.011385 | 0.014441 | 0.017939 | 0.023068 | 0.028754 | 0.036271 | 0.045614 | 0.058069 | 0.073874 | 0.094576 | 0.12111 | 0.15588 | 0.20027 | 0.2566 | 0.32289 | 0.39693 | 0.47149 | 0.53464 | 0.57592 | 0.58749 | 0.56901 | 0.52696 | 0.47094 | 0.40966 | 0.35018 | 0.29633 | 0.24993 | 0.21217 | 0.18272 | 0.16189 | 0.14492 | 0.13366 | 0.12856 | 0.12752 | 0.12988 | 0.13813 | 0.14441 | 0.16577 | 0.18416 | 0.20379 | 0.22754 | 0.25211 | 0.27626 | 0.29768
Time 8.5 h, -40 celcius | 0.0037257 | 0.0037257 | 0.0022099 | 0.0055149 | 0.0027352 | 0.0046061 | 0.0028107 | 0.0031095 | 0.0040131 | 0.0050574 | 0.0061923 | 0.007861 | 0.0096387 | 0.01196 | 0.014763 | 0.018394 | 0.022849 | 0.028488 | 0.035374 | 0.043956 | 0.054453 | 0.067485 | 0.083211 | 0.10248 | 0.12627 | 0.15518 | 0.19059 | 0.23316 | 0.28257 | 0.33712 | 0.39303 | 0.44437 | 0.48479 | 0.5087 | 0.51347 | 0.50011 | 0.47159 | 0.43276 | 0.38638 | 0.3379 | 0.29199 | 0.24717 | 0.21022 | 0.17522 | 0.14921 | 0.12645 | 0.10764 | 0.094062 | 0.085294 | 0.077832 | 0.076166 | 0.074374
Std. : Water (1:0)_ 5 g_17C | 0.40745 | 0.37758 | 0.33167 | 0.27797 | 0.22443 | 0.19417 | 0.16169 | 0.13849 | 0.12143 | 0.10697 | 0.096405 | 0.087034 | 0.081871 | 0.078656 | 0.076272 | 0.075226 | 0.07448 | 0.073601 | 0.073296 | 0.073701 | 0.074654 | 0.075483 | 0.075353 | 0.07631 | 0.0769 | 0.077719 | 0.079173 | 0.081297 | 0.084848 | 0.090883 | 0.098449 | 0.10715 | 0.11892 | 0.13057 | 0.14131 | 0.15224 | 0.161 | 0.1718 | 0.18315 | 0.19373 | 0.20789 | 0.22266 | 0.23803 | 0.25543 | 0.2743 | 0.29629 | 0.32039 | 0.34659
Std. : Water (1:3)_5 g_22.3C | 0.35022 | 0.37378 | 0.41173 | 0.55328 | 0.59494 | 0.61621 | 0.60526 | 0.57334 | 0.52449 | 0.45124 | 0.3927 | 0.33463 | 0.28521 | 0.24069 | 0.2067 | 0.17971 | 0.15895 | 0.14354 | 0.13098 | 0.1209 | 0.11373 | 0.10865 | 0.10508 | 0.10263 | 0.10025 | 0.099388 | 0.099009 | 0.09943 | 0.10029 | 0.10249 | 0.10653 | 0.11105 | 0.11816 | 0.12566 | 0.13766 | 0.14944 | 0.16023 | 0.17158 | 0.18231 | 0.19465 | 0.20803 | 0.22162 | 0.23412 | 0.25639 | 0.27568 | 0.29801 | 0.32242 | 0.35086 | 0.38221 | 0.41669
Fitting-Time 0.23 h (20.3 C) Liquid | 0.00080148 | 0.001009 | 0.0012703 | 0.0015992 | 0.0020132 | 0.0025345 | 0.0031908 | 0.0040169 | 0.005057 | 0.0063663 | 0.0080146 | 0.010089 | 0.012702 | 0.01599 | 0.020128 | 0.025337 | 0.031892 | 0.040136 | 0.050507 | 0.063538 | 0.07988 | 0.10037 | 0.126 | 0.15786 | 0.1972 | 0.24538 | 0.30315 | 0.37047 | 0.44655 | 0.52567 | 0.59992 | 0.65886 | 0.68897 | 0.68388 | 0.64493 | 0.58011 | 0.50283 | 0.42213 | 0.34988 | 0.28493 | 0.23165 | 0.18163 | 0.14738 | 0.11839 | 0.089082 | 0.071374 | 0.059526 | 0.041708 | 0.035758 | 0.029805 | 0.023848
Fitting-Time 5.03 h (-20 C) Frozen | 0.0048296 | 0.0060799 | 0.0076542 | 0.0096359 | 0.01213 | 0.01527 | 0.019222 | 0.024195 | 0.030452 | 0.038322 | 0.048213 | 0.060635 | 0.076215 | 0.095709 | 0.12001 | 0.15014 | 0.18717 | 0.232 | 0.28517 | 0.34602 | 0.41191 | 0.4775 | 0.53412 | 0.57135 | 0.58091 | 0.56037 | 0.51476 | 0.45386 | 0.38691 | 0.32252 | 0.26465 | 0.2141 | 0.17243 | 0.138 | 0.11029 | 0.087693 | 0.069745 | 0.055173 | 0.044035 | 0.034967 | 0.027983 | 0.021692 | 0.017495 | 0.013998 | 0.010499 | 0.0083994 | 0.0069995 | 0.0048997 | 0.0041998 | 0.0034999 | 0.0027999
(I)
18 GHz
20.3 °C
20.3 °C
18 kHz
0.0056 M KCl
0.0056 M KCl
(II)
-20 °C
(III)
2.6 kHz
-20 °C
Real part Capacitance
Imaginary part Capacitance
Measurement vial
Co
+
-
-
+
+
+
+
+
+
+
+
+
+
+
-
-
-
-
-
-
-
-
-
-
+
+
+
+
+
+
+
+
+
+
-
-
-
-
-
-
-
-
-
-
Ice
Glass wall
Electronic polarization
Atomic polarization
L- and D- orientation defects
Notes: D:\PhD\Assignment\Conference ppt\Plot used in smi 2019 Conference_v1

## Slide 17: Dielectric Loss Mechanisms
The dielectric polarization of ice at −40 ˚C with a dielectric loss peak frequencies of 537 Hz.
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 0.23 h, 20.3 celcius | 1.0162 | 1.0162 | 1.0162 | 1.0162 | 1.0176 | 1.024 | 1.0293 | 1.0371 | 1.0465 | 1.0575 | 1.0742 | 1.0971 | 1.131 | 1.18 | 1.2538 | 1.3532 | 1.4803 | 1.6281 | 1.7827 | 1.9288 | 2.0567 | 2.1587 | 2.2374 | 2.2971 | 2.3414 | 2.3756 | 2.4038 | 2.4272 | 2.4476 | 2.4662 | 2.4852 | 2.5003 | 2.5152 | 2.5286 | 2.5419 | 2.5531 | 2.5615 | 2.5717 | 2.5802 | 2.589 | 2.5992 | 2.6061 | 2.6201 | 2.625 | 2.6478 | 2.6645 | 2.686 | 2.708 | 2.7359 | 2.7645 | 2.7969
Time 5.03 h, -20 celcius | 0.084287 | 0.084287 | 0.084287 | 0.084287 | 0.083003 | 0.084127 | 0.084469 | 0.085543 | 0.086743 | 0.087572 | 0.088047 | 0.088435 | 0.088561 | 0.088505 | 0.088289 | 0.08811 | 0.08822 | 0.08935 | 0.092855 | 0.10126 | 0.11875 | 0.15045 | 0.2025 | 0.28088 | 0.38519 | 0.51085 | 0.64628 | 0.77735 | 0.89303 | 0.98847 | 1.0619 | 1.1159 | 1.155 | 1.183 | 1.203 | 1.2167 | 1.2267 | 1.2373 | 1.2444 | 1.2521 | 1.2603 | 1.2702 | 1.2796 | 1.2944 | 1.3089 | 1.3285 | 1.3531 | 1.3812 | 1.4143 | 1.4539 | 1.499
Time 8.5 h, -40 celcius | 0.083695 | 0.083695 | 0.083695 | 0.083695 | 0.083695 | 0.08324 | 0.083622 | 0.083847 | 0.084834 | 0.086128 | 0.086714 | 0.087381 | 0.087886 | 0.088558 | 0.089009 | 0.089429 | 0.089952 | 0.090576 | 0.09152 | 0.092833 | 0.094685 | 0.097293 | 0.10083 | 0.10562 | 0.11226 | 0.12153 | 0.13502 | 0.15512 | 0.18502 | 0.22832 | 0.28819 | 0.36571 | 0.45912 | 0.56375 | 0.67242 | 0.77798 | 0.87404 | 0.95702 | 1.0284 | 1.0845 | 1.1268 | 1.1614 | 1.1873 | 1.2052 | 1.2167 | 1.2287 | 1.2359 | 1.2394 | 1.2468 | 1.2512 | 1.2552 | 1.2597
Std. : Water (1:0)_ 5 g_17C | 2.3304 | 2.4387 | 2.4947 | 2.5334 | 2.5734 | 2.595 | 2.6101 | 2.6237 | 2.6363 | 2.6458 | 2.6525 | 2.6605 | 2.6675 | 2.6754 | 2.6841 | 2.6936 | 2.7028 | 2.7121 | 2.7221 | 2.732 | 2.7416 | 2.7519 | 2.7627 | 2.773 | 2.7842 | 2.7949 | 2.8043 | 2.8148 | 2.8246 | 2.8354 | 2.8466 | 2.8598 | 2.8739 | 2.8899 | 2.908 | 2.9285 | 2.9493 | 2.972 | 2.9971 | 3.0212 | 3.0504 | 3.076 | 3.1069 | 3.1381 | 3.1714 | 3.2072 | 3.2459 | 3.2864
Std. : Water (1:3)_5 g_22.3C | 1.4671 | 1.6855 | 1.744 | 1.8824 | 1.9929 | 2.1292 | 2.3085 | 2.4363 | 2.5381 | 2.631 | 2.6959 | 2.7426 | 2.7767 | 2.8037 | 2.8219 | 2.8373 | 2.8512 | 2.8646 | 2.8761 | 2.8875 | 2.8997 | 2.9118 | 2.9235 | 2.9355 | 2.9478 | 2.9606 | 2.9734 | 2.9866 | 2.9991 | 3.0123 | 3.0249 | 3.039 | 3.0531 | 3.0685 | 3.0846 | 3.1036 | 3.1223 | 3.1442 | 3.1674 | 3.1919 | 3.2199 | 3.2461 | 3.2797 | 3.3073 | 3.3426 | 3.3775 | 3.4158 | 3.4569 | 3.5017 | 3.5489
Fitting-Time 0.23 h (20.3 C) Liquid | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.4119 | 2.4119 | 2.4119 | 2.4119 | 2.4118 | 2.4117 | 2.4115 | 2.4112 | 2.4108 | 2.4101 | 2.409 | 2.4073 | 2.4046 | 2.4004 | 2.3937 | 2.3832 | 2.367 | 2.342 | 2.3043 | 2.2484 | 2.1697 | 2.0642 | 1.9301 | 1.7781 | 1.6192 | 1.4715 | 1.3445 | 1.2461 | 1.1731 | 1.1243 | 1.0907 | 1.0692 | 1.0535 | 1.0451 | 1.0395 | 1.035 | 1.0329 | 1.0318 | 1.0305 | 1.0302 | 1.0299 | 1.0297
Fitting-Time 5.03 h (-20 C) Frozen | 1.2389 | 1.2389 | 1.2389 | 1.2389 | 1.2388 | 1.2387 | 1.2386 | 1.2384 | 1.2381 | 1.2377 | 1.2369 | 1.2358 | 1.2339 | 1.231 | 1.2264 | 1.2192 | 1.208 | 1.1906 | 1.1642 | 1.1248 | 1.0679 | 0.98926 | 0.88724 | 0.76529 | 0.63274 | 0.50242 | 0.38716 | 0.29408 | 0.22349 | 0.17372 | 0.13979 | 0.11692 | 0.10222 | 0.092682 | 0.086624 | 0.08272 | 0.080267 | 0.078693 | 0.077739 | 0.077121 | 0.076743 | 0.076474 | 0.076332 | 0.076237 | 0.076164 | 0.07613 | 0.076111 | 0.076089 | 0.076084 | 0.076079 | 0.076076
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 0.23 h, 20.3 celcius | 0.021874 | 0.027746 | 0.035469 | 0.045629 | 0.058992 | 0.069416 | 0.087012 | 0.11672 | 0.14702 | 0.18167 | 0.23092 | 0.28348 | 0.3471 | 0.41787 | 0.49639 | 0.57052 | 0.6315 | 0.66609 | 0.66868 | 0.64015 | 0.58783 | 0.52385 | 0.45806 | 0.3962 | 0.3406 | 0.2936 | 0.25428 | 0.22269 | 0.19622 | 0.17597 | 0.15804 | 0.14279 | 0.13021 | 0.12053 | 0.11142 | 0.10593 | 0.099251 | 0.10112 | 0.096409 | 0.099903 | 0.099379 | 0.11049 | 0.12227 | 0.13167 | 0.14829 | 0.16415 | 0.18118 | 0.20571 | 0.22768 | 0.25157 | 0.27443
Time 5.03 h, -20 celcius | 0.0046903 | 0.0060276 | 0.0048957 | 0.0052689 | 0.0080706 | 0.0070995 | 0.0085621 | 0.011385 | 0.014441 | 0.017939 | 0.023068 | 0.028754 | 0.036271 | 0.045614 | 0.058069 | 0.073874 | 0.094576 | 0.12111 | 0.15588 | 0.20027 | 0.2566 | 0.32289 | 0.39693 | 0.47149 | 0.53464 | 0.57592 | 0.58749 | 0.56901 | 0.52696 | 0.47094 | 0.40966 | 0.35018 | 0.29633 | 0.24993 | 0.21217 | 0.18272 | 0.16189 | 0.14492 | 0.13366 | 0.12856 | 0.12752 | 0.12988 | 0.13813 | 0.14441 | 0.16577 | 0.18416 | 0.20379 | 0.22754 | 0.25211 | 0.27626 | 0.29768
Time 8.5 h, -40 celcius | 0.0037257 | 0.0037257 | 0.0022099 | 0.0055149 | 0.0027352 | 0.0046061 | 0.0028107 | 0.0031095 | 0.0040131 | 0.0050574 | 0.0061923 | 0.007861 | 0.0096387 | 0.01196 | 0.014763 | 0.018394 | 0.022849 | 0.028488 | 0.035374 | 0.043956 | 0.054453 | 0.067485 | 0.083211 | 0.10248 | 0.12627 | 0.15518 | 0.19059 | 0.23316 | 0.28257 | 0.33712 | 0.39303 | 0.44437 | 0.48479 | 0.5087 | 0.51347 | 0.50011 | 0.47159 | 0.43276 | 0.38638 | 0.3379 | 0.29199 | 0.24717 | 0.21022 | 0.17522 | 0.14921 | 0.12645 | 0.10764 | 0.094062 | 0.085294 | 0.077832 | 0.076166 | 0.074374
Std. : Water (1:0)_ 5 g_17C | 0.40745 | 0.37758 | 0.33167 | 0.27797 | 0.22443 | 0.19417 | 0.16169 | 0.13849 | 0.12143 | 0.10697 | 0.096405 | 0.087034 | 0.081871 | 0.078656 | 0.076272 | 0.075226 | 0.07448 | 0.073601 | 0.073296 | 0.073701 | 0.074654 | 0.075483 | 0.075353 | 0.07631 | 0.0769 | 0.077719 | 0.079173 | 0.081297 | 0.084848 | 0.090883 | 0.098449 | 0.10715 | 0.11892 | 0.13057 | 0.14131 | 0.15224 | 0.161 | 0.1718 | 0.18315 | 0.19373 | 0.20789 | 0.22266 | 0.23803 | 0.25543 | 0.2743 | 0.29629 | 0.32039 | 0.34659
Std. : Water (1:3)_5 g_22.3C | 0.35022 | 0.37378 | 0.41173 | 0.55328 | 0.59494 | 0.61621 | 0.60526 | 0.57334 | 0.52449 | 0.45124 | 0.3927 | 0.33463 | 0.28521 | 0.24069 | 0.2067 | 0.17971 | 0.15895 | 0.14354 | 0.13098 | 0.1209 | 0.11373 | 0.10865 | 0.10508 | 0.10263 | 0.10025 | 0.099388 | 0.099009 | 0.09943 | 0.10029 | 0.10249 | 0.10653 | 0.11105 | 0.11816 | 0.12566 | 0.13766 | 0.14944 | 0.16023 | 0.17158 | 0.18231 | 0.19465 | 0.20803 | 0.22162 | 0.23412 | 0.25639 | 0.27568 | 0.29801 | 0.32242 | 0.35086 | 0.38221 | 0.41669
Fitting-Time 0.23 h (20.3 C) Liquid | 0.00080148 | 0.001009 | 0.0012703 | 0.0015992 | 0.0020132 | 0.0025345 | 0.0031908 | 0.0040169 | 0.005057 | 0.0063663 | 0.0080146 | 0.010089 | 0.012702 | 0.01599 | 0.020128 | 0.025337 | 0.031892 | 0.040136 | 0.050507 | 0.063538 | 0.07988 | 0.10037 | 0.126 | 0.15786 | 0.1972 | 0.24538 | 0.30315 | 0.37047 | 0.44655 | 0.52567 | 0.59992 | 0.65886 | 0.68897 | 0.68388 | 0.64493 | 0.58011 | 0.50283 | 0.42213 | 0.34988 | 0.28493 | 0.23165 | 0.18163 | 0.14738 | 0.11839 | 0.089082 | 0.071374 | 0.059526 | 0.041708 | 0.035758 | 0.029805 | 0.023848
Fitting-Time 5.03 h (-20 C) Frozen | 0.0048296 | 0.0060799 | 0.0076542 | 0.0096359 | 0.01213 | 0.01527 | 0.019222 | 0.024195 | 0.030452 | 0.038322 | 0.048213 | 0.060635 | 0.076215 | 0.095709 | 0.12001 | 0.15014 | 0.18717 | 0.232 | 0.28517 | 0.34602 | 0.41191 | 0.4775 | 0.53412 | 0.57135 | 0.58091 | 0.56037 | 0.51476 | 0.45386 | 0.38691 | 0.32252 | 0.26465 | 0.2141 | 0.17243 | 0.138 | 0.11029 | 0.087693 | 0.069745 | 0.055173 | 0.044035 | 0.034967 | 0.027983 | 0.021692 | 0.017495 | 0.013998 | 0.010499 | 0.0083994 | 0.0069995 | 0.0048997 | 0.0041998 | 0.0034999 | 0.0027999
(I)
18 GHz
20.3 °C
20.3 °C
18 kHz
(II)
-20 °C
2.6 kHz
-20 °C
-40 °C
-40 °C
537 Hz
(IV)
0.0056 M KCl
0.0056 M KCl
(III)
Real part Capacitance
Imaginary part Capacitance
Measurement vial
Co
+
-
-
+
+
+
+
+
+
+
+
+
+
+
-
-
-
-
-
-
-
-
-
-
+
+
+
+
+
+
+
+
+
+
-
-
-
-
-
-
-
-
-
-
Ice
Glass wall
Electronic polarization
Atomic polarization
Ionic defects
(similar to Grotthus mechanism)
Notes: D:\PhD\Assignment\Conference ppt\Plot used in smi 2019 Conference_v1

## Slide 18: Dielectric Loss Mechanisms
The polarization of the water dipole in liquid water at 20 ˚C, with a dielectric loss peak frequency of ~ 18 GHz
Maxwell-Wagner (MW) polarization of the glass wall of the TVIS vial at +20 ˚C, with a dielectric loss peak frequency of 17.8 kHz
The dielectric polarization of ice at −20 ˚C, with a dielectric loss peak frequencies of 2.57 kHz
The dielectric polarization of ice at −40 ˚C with a dielectric loss peak frequencies of 537 Hz.
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 0.23 h, 20.3 celcius | 1.0162 | 1.0162 | 1.0162 | 1.0162 | 1.0176 | 1.024 | 1.0293 | 1.0371 | 1.0465 | 1.0575 | 1.0742 | 1.0971 | 1.131 | 1.18 | 1.2538 | 1.3532 | 1.4803 | 1.6281 | 1.7827 | 1.9288 | 2.0567 | 2.1587 | 2.2374 | 2.2971 | 2.3414 | 2.3756 | 2.4038 | 2.4272 | 2.4476 | 2.4662 | 2.4852 | 2.5003 | 2.5152 | 2.5286 | 2.5419 | 2.5531 | 2.5615 | 2.5717 | 2.5802 | 2.589 | 2.5992 | 2.6061 | 2.6201 | 2.625 | 2.6478 | 2.6645 | 2.686 | 2.708 | 2.7359 | 2.7645 | 2.7969
Time 5.03 h, -20 celcius | 0.084287 | 0.084287 | 0.084287 | 0.084287 | 0.083003 | 0.084127 | 0.084469 | 0.085543 | 0.086743 | 0.087572 | 0.088047 | 0.088435 | 0.088561 | 0.088505 | 0.088289 | 0.08811 | 0.08822 | 0.08935 | 0.092855 | 0.10126 | 0.11875 | 0.15045 | 0.2025 | 0.28088 | 0.38519 | 0.51085 | 0.64628 | 0.77735 | 0.89303 | 0.98847 | 1.0619 | 1.1159 | 1.155 | 1.183 | 1.203 | 1.2167 | 1.2267 | 1.2373 | 1.2444 | 1.2521 | 1.2603 | 1.2702 | 1.2796 | 1.2944 | 1.3089 | 1.3285 | 1.3531 | 1.3812 | 1.4143 | 1.4539 | 1.499
Time 8.5 h, -40 celcius | 0.083695 | 0.083695 | 0.083695 | 0.083695 | 0.083695 | 0.08324 | 0.083622 | 0.083847 | 0.084834 | 0.086128 | 0.086714 | 0.087381 | 0.087886 | 0.088558 | 0.089009 | 0.089429 | 0.089952 | 0.090576 | 0.09152 | 0.092833 | 0.094685 | 0.097293 | 0.10083 | 0.10562 | 0.11226 | 0.12153 | 0.13502 | 0.15512 | 0.18502 | 0.22832 | 0.28819 | 0.36571 | 0.45912 | 0.56375 | 0.67242 | 0.77798 | 0.87404 | 0.95702 | 1.0284 | 1.0845 | 1.1268 | 1.1614 | 1.1873 | 1.2052 | 1.2167 | 1.2287 | 1.2359 | 1.2394 | 1.2468 | 1.2512 | 1.2552 | 1.2597
Std. : Water (1:0)_ 5 g_17C | 2.3304 | 2.4387 | 2.4947 | 2.5334 | 2.5734 | 2.595 | 2.6101 | 2.6237 | 2.6363 | 2.6458 | 2.6525 | 2.6605 | 2.6675 | 2.6754 | 2.6841 | 2.6936 | 2.7028 | 2.7121 | 2.7221 | 2.732 | 2.7416 | 2.7519 | 2.7627 | 2.773 | 2.7842 | 2.7949 | 2.8043 | 2.8148 | 2.8246 | 2.8354 | 2.8466 | 2.8598 | 2.8739 | 2.8899 | 2.908 | 2.9285 | 2.9493 | 2.972 | 2.9971 | 3.0212 | 3.0504 | 3.076 | 3.1069 | 3.1381 | 3.1714 | 3.2072 | 3.2459 | 3.2864
Std. : Water (1:3)_5 g_22.3C | 1.4671 | 1.6855 | 1.744 | 1.8824 | 1.9929 | 2.1292 | 2.3085 | 2.4363 | 2.5381 | 2.631 | 2.6959 | 2.7426 | 2.7767 | 2.8037 | 2.8219 | 2.8373 | 2.8512 | 2.8646 | 2.8761 | 2.8875 | 2.8997 | 2.9118 | 2.9235 | 2.9355 | 2.9478 | 2.9606 | 2.9734 | 2.9866 | 2.9991 | 3.0123 | 3.0249 | 3.039 | 3.0531 | 3.0685 | 3.0846 | 3.1036 | 3.1223 | 3.1442 | 3.1674 | 3.1919 | 3.2199 | 3.2461 | 3.2797 | 3.3073 | 3.3426 | 3.3775 | 3.4158 | 3.4569 | 3.5017 | 3.5489
Fitting-Time 0.23 h (20.3 C) Liquid | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.4119 | 2.4119 | 2.4119 | 2.4119 | 2.4118 | 2.4117 | 2.4115 | 2.4112 | 2.4108 | 2.4101 | 2.409 | 2.4073 | 2.4046 | 2.4004 | 2.3937 | 2.3832 | 2.367 | 2.342 | 2.3043 | 2.2484 | 2.1697 | 2.0642 | 1.9301 | 1.7781 | 1.6192 | 1.4715 | 1.3445 | 1.2461 | 1.1731 | 1.1243 | 1.0907 | 1.0692 | 1.0535 | 1.0451 | 1.0395 | 1.035 | 1.0329 | 1.0318 | 1.0305 | 1.0302 | 1.0299 | 1.0297
Fitting-Time 5.03 h (-20 C) Frozen | 1.2389 | 1.2389 | 1.2389 | 1.2389 | 1.2388 | 1.2387 | 1.2386 | 1.2384 | 1.2381 | 1.2377 | 1.2369 | 1.2358 | 1.2339 | 1.231 | 1.2264 | 1.2192 | 1.208 | 1.1906 | 1.1642 | 1.1248 | 1.0679 | 0.98926 | 0.88724 | 0.76529 | 0.63274 | 0.50242 | 0.38716 | 0.29408 | 0.22349 | 0.17372 | 0.13979 | 0.11692 | 0.10222 | 0.092682 | 0.086624 | 0.08272 | 0.080267 | 0.078693 | 0.077739 | 0.077121 | 0.076743 | 0.076474 | 0.076332 | 0.076237 | 0.076164 | 0.07613 | 0.076111 | 0.076089 | 0.076084 | 0.076079 | 0.076076
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 0.23 h, 20.3 celcius | 0.021874 | 0.027746 | 0.035469 | 0.045629 | 0.058992 | 0.069416 | 0.087012 | 0.11672 | 0.14702 | 0.18167 | 0.23092 | 0.28348 | 0.3471 | 0.41787 | 0.49639 | 0.57052 | 0.6315 | 0.66609 | 0.66868 | 0.64015 | 0.58783 | 0.52385 | 0.45806 | 0.3962 | 0.3406 | 0.2936 | 0.25428 | 0.22269 | 0.19622 | 0.17597 | 0.15804 | 0.14279 | 0.13021 | 0.12053 | 0.11142 | 0.10593 | 0.099251 | 0.10112 | 0.096409 | 0.099903 | 0.099379 | 0.11049 | 0.12227 | 0.13167 | 0.14829 | 0.16415 | 0.18118 | 0.20571 | 0.22768 | 0.25157 | 0.27443
Time 5.03 h, -20 celcius | 0.0046903 | 0.0060276 | 0.0048957 | 0.0052689 | 0.0080706 | 0.0070995 | 0.0085621 | 0.011385 | 0.014441 | 0.017939 | 0.023068 | 0.028754 | 0.036271 | 0.045614 | 0.058069 | 0.073874 | 0.094576 | 0.12111 | 0.15588 | 0.20027 | 0.2566 | 0.32289 | 0.39693 | 0.47149 | 0.53464 | 0.57592 | 0.58749 | 0.56901 | 0.52696 | 0.47094 | 0.40966 | 0.35018 | 0.29633 | 0.24993 | 0.21217 | 0.18272 | 0.16189 | 0.14492 | 0.13366 | 0.12856 | 0.12752 | 0.12988 | 0.13813 | 0.14441 | 0.16577 | 0.18416 | 0.20379 | 0.22754 | 0.25211 | 0.27626 | 0.29768
Time 8.5 h, -40 celcius | 0.0037257 | 0.0037257 | 0.0022099 | 0.0055149 | 0.0027352 | 0.0046061 | 0.0028107 | 0.0031095 | 0.0040131 | 0.0050574 | 0.0061923 | 0.007861 | 0.0096387 | 0.01196 | 0.014763 | 0.018394 | 0.022849 | 0.028488 | 0.035374 | 0.043956 | 0.054453 | 0.067485 | 0.083211 | 0.10248 | 0.12627 | 0.15518 | 0.19059 | 0.23316 | 0.28257 | 0.33712 | 0.39303 | 0.44437 | 0.48479 | 0.5087 | 0.51347 | 0.50011 | 0.47159 | 0.43276 | 0.38638 | 0.3379 | 0.29199 | 0.24717 | 0.21022 | 0.17522 | 0.14921 | 0.12645 | 0.10764 | 0.094062 | 0.085294 | 0.077832 | 0.076166 | 0.074374
Std. : Water (1:0)_ 5 g_17C | 0.40745 | 0.37758 | 0.33167 | 0.27797 | 0.22443 | 0.19417 | 0.16169 | 0.13849 | 0.12143 | 0.10697 | 0.096405 | 0.087034 | 0.081871 | 0.078656 | 0.076272 | 0.075226 | 0.07448 | 0.073601 | 0.073296 | 0.073701 | 0.074654 | 0.075483 | 0.075353 | 0.07631 | 0.0769 | 0.077719 | 0.079173 | 0.081297 | 0.084848 | 0.090883 | 0.098449 | 0.10715 | 0.11892 | 0.13057 | 0.14131 | 0.15224 | 0.161 | 0.1718 | 0.18315 | 0.19373 | 0.20789 | 0.22266 | 0.23803 | 0.25543 | 0.2743 | 0.29629 | 0.32039 | 0.34659
Std. : Water (1:3)_5 g_22.3C | 0.35022 | 0.37378 | 0.41173 | 0.55328 | 0.59494 | 0.61621 | 0.60526 | 0.57334 | 0.52449 | 0.45124 | 0.3927 | 0.33463 | 0.28521 | 0.24069 | 0.2067 | 0.17971 | 0.15895 | 0.14354 | 0.13098 | 0.1209 | 0.11373 | 0.10865 | 0.10508 | 0.10263 | 0.10025 | 0.099388 | 0.099009 | 0.09943 | 0.10029 | 0.10249 | 0.10653 | 0.11105 | 0.11816 | 0.12566 | 0.13766 | 0.14944 | 0.16023 | 0.17158 | 0.18231 | 0.19465 | 0.20803 | 0.22162 | 0.23412 | 0.25639 | 0.27568 | 0.29801 | 0.32242 | 0.35086 | 0.38221 | 0.41669
Fitting-Time 0.23 h (20.3 C) Liquid | 0.00080148 | 0.001009 | 0.0012703 | 0.0015992 | 0.0020132 | 0.0025345 | 0.0031908 | 0.0040169 | 0.005057 | 0.0063663 | 0.0080146 | 0.010089 | 0.012702 | 0.01599 | 0.020128 | 0.025337 | 0.031892 | 0.040136 | 0.050507 | 0.063538 | 0.07988 | 0.10037 | 0.126 | 0.15786 | 0.1972 | 0.24538 | 0.30315 | 0.37047 | 0.44655 | 0.52567 | 0.59992 | 0.65886 | 0.68897 | 0.68388 | 0.64493 | 0.58011 | 0.50283 | 0.42213 | 0.34988 | 0.28493 | 0.23165 | 0.18163 | 0.14738 | 0.11839 | 0.089082 | 0.071374 | 0.059526 | 0.041708 | 0.035758 | 0.029805 | 0.023848
Fitting-Time 5.03 h (-20 C) Frozen | 0.0048296 | 0.0060799 | 0.0076542 | 0.0096359 | 0.01213 | 0.01527 | 0.019222 | 0.024195 | 0.030452 | 0.038322 | 0.048213 | 0.060635 | 0.076215 | 0.095709 | 0.12001 | 0.15014 | 0.18717 | 0.232 | 0.28517 | 0.34602 | 0.41191 | 0.4775 | 0.53412 | 0.57135 | 0.58091 | 0.56037 | 0.51476 | 0.45386 | 0.38691 | 0.32252 | 0.26465 | 0.2141 | 0.17243 | 0.138 | 0.11029 | 0.087693 | 0.069745 | 0.055173 | 0.044035 | 0.034967 | 0.027983 | 0.021692 | 0.017495 | 0.013998 | 0.010499 | 0.0083994 | 0.0069995 | 0.0048997 | 0.0041998 | 0.0034999 | 0.0027999
Real part Capacitance
Imaginary part Capacitance
(I)
18 GHz
20.3 °C
20.3 °C
18 kHz
-20 °C
(III)
2.6 kHz
-20 °C
-40 °C
-40 °C
537 Hz
(IV)
(II)
0.0056 M KCl
0.0056 M KCl
Notes: D:\PhD\Assignment\Conference ppt\Plot used in smi 2019 Conference_v1

## Slide 19
TVIS Applications
TVIS parameters for different facets of the freezing process

## Slide 20: Methodology
Freezing step of freeze drying process
Thermocouple (TC)
TVIS vial
TC vial
TC in the nearest neighbour vial at a height equivalent to the mid point of the volume bounded by the electrodes
Shelf temperature
Product temperature
Chart: Time / h
Notes: D:\PhD\Assignment\Conference ppt\Plot used in smi 2019 Conference_v1
Process log for Sugar
Data spectrum_Real Capa-S5 (135z, 4.5 h, -43.9 C)
Data spectrum_Img Capa-S5 (135z, 4.5 h, -43.9 C)

## Slide 21: Methodology
Freezing step of freeze drying process
In-line TVIS measurement
Thermocouple (TC)
TVIS vial
TC vial
TC in the nearest neighbour vial at a height equivalent to the mid point of the volume bounded by the electrodes
Shelf temperature
Product temperature
Chart: Time / h
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
135.z | 2.2958 | 2.2665 | 2.2338 | 2.1996 | 2.1627 | 2.1229 | 2.0815 | 2.0395 | 1.9939 | 1.9478 | 1.9015 | 1.8545 | 1.8081 | 1.7613 | 1.7152 | 1.6733 | 1.6328 | 1.593 | 1.5561 | 1.5212 | 1.4883 | 1.4567 | 1.4247 | 1.3904 | 1.35 | 1.3007 | 1.2373 | 1.1557 | 1.051 | 0.92437 | 0.77971 | 0.62511 | 0.48011 | 0.35624 | 0.26346 | 0.20048 | 0.16263 | 0.14169 | 0.13139 | 0.12641 | 0.12379 | 0.12214 | 0.12034 | 0.11797 | 0.114 | 0.11172 | 0.093235
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
135.z | 2.2958 | 2.2665 | 2.2338 | 2.1996 | 2.1627 | 2.1229 | 2.0815 | 2.0395 | 1.9939 | 1.9478 | 1.9015 | 1.8545 | 1.8081 | 1.7613 | 1.7152 | 1.6733 | 1.6328 | 1.593 | 1.5561 | 1.5212 | 1.4883 | 1.4567 | 1.4247 | 1.3904 | 1.35 | 1.3007 | 1.2373 | 1.1557 | 1.051 | 0.92437 | 0.77971 | 0.62511 | 0.48011 | 0.35624 | 0.26346 | 0.20048 | 0.16263 | 0.14169 | 0.13139 | 0.12641 | 0.12379 | 0.12214 | 0.12034 | 0.11797 | 0.114 | 0.11172 | 0.093235
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
135.z | 0.2096 | 0.22054 | 0.23152 | 0.24243 | 0.2538 | 0.2645 | 0.27437 | 0.28534 | 0.28854 | 0.29617 | 0.29907 | 0.30222 | 0.30213 | 0.30143 | 0.29916 | 0.29589 | 0.29181 | 0.28961 | 0.28866 | 0.29101 | 0.29708 | 0.30917 | 0.32872 | 0.35705 | 0.3947 | 0.44202 | 0.49692 | 0.5557 | 0.61245 | 0.65663 | 0.67796 | 0.66767 | 0.62387 | 0.55258 | 0.46738 | 0.38017 | 0.30196 | 0.23564 | 0.18477 | 0.14422 | 0.11363 | 0.087151 | 0.069446 | 0.054131 | 0.039552 | 0.03036 | 0.022405
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
135.z | 0.2096 | 0.22054 | 0.23152 | 0.24243 | 0.2538 | 0.2645 | 0.27437 | 0.28534 | 0.28854 | 0.29617 | 0.29907 | 0.30222 | 0.30213 | 0.30143 | 0.29916 | 0.29589 | 0.29181 | 0.28961 | 0.28866 | 0.29101 | 0.29708 | 0.30917 | 0.32872 | 0.35705 | 0.3947 | 0.44202 | 0.49692 | 0.5557 | 0.61245 | 0.65663 | 0.67796 | 0.66767 | 0.62387 | 0.55258 | 0.46738 | 0.38017 | 0.30196 | 0.23564 | 0.18477 | 0.14422 | 0.11363 | 0.087151 | 0.069446 | 0.054131 | 0.039552 | 0.03036 | 0.022405
Notes: D:\PhD\Assignment\Conference ppt\Plot used in smi 2019 Conference_v1
Process log for Sugar
Data spectrum_Real Capa-S5 (135z, 4.5 h, -43.9 C)
Data spectrum_Img Capa-S5 (135z, 4.5 h, -43.9 C)

## Slide 22: Methodology
Freezing step of freeze drying process
In-line TVIS measurement
Identifying TVIS parameters using LyoView software
Determination of Ice nucleation event
Thermocouple (TC)
TVIS vial
TC vial
TC in the nearest neighbour vial at a height equivalent to the mid point of the volume bounded by the electrodes
Shelf temperature
Product temperature
(10 Hz) (0.2MHz)
Chart: Time / h
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
135.z | 2.2958 | 2.2665 | 2.2338 | 2.1996 | 2.1627 | 2.1229 | 2.0815 | 2.0395 | 1.9939 | 1.9478 | 1.9015 | 1.8545 | 1.8081 | 1.7613 | 1.7152 | 1.6733 | 1.6328 | 1.593 | 1.5561 | 1.5212 | 1.4883 | 1.4567 | 1.4247 | 1.3904 | 1.35 | 1.3007 | 1.2373 | 1.1557 | 1.051 | 0.92437 | 0.77971 | 0.62511 | 0.48011 | 0.35624 | 0.26346 | 0.20048 | 0.16263 | 0.14169 | 0.13139 | 0.12641 | 0.12379 | 0.12214 | 0.12034 | 0.11797 | 0.114 | 0.11172 | 0.093235
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
135.z | 2.2958 | 2.2665 | 2.2338 | 2.1996 | 2.1627 | 2.1229 | 2.0815 | 2.0395 | 1.9939 | 1.9478 | 1.9015 | 1.8545 | 1.8081 | 1.7613 | 1.7152 | 1.6733 | 1.6328 | 1.593 | 1.5561 | 1.5212 | 1.4883 | 1.4567 | 1.4247 | 1.3904 | 1.35 | 1.3007 | 1.2373 | 1.1557 | 1.051 | 0.92437 | 0.77971 | 0.62511 | 0.48011 | 0.35624 | 0.26346 | 0.20048 | 0.16263 | 0.14169 | 0.13139 | 0.12641 | 0.12379 | 0.12214 | 0.12034 | 0.11797 | 0.114 | 0.11172 | 0.093235
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
135.z | 0.2096 | 0.22054 | 0.23152 | 0.24243 | 0.2538 | 0.2645 | 0.27437 | 0.28534 | 0.28854 | 0.29617 | 0.29907 | 0.30222 | 0.30213 | 0.30143 | 0.29916 | 0.29589 | 0.29181 | 0.28961 | 0.28866 | 0.29101 | 0.29708 | 0.30917 | 0.32872 | 0.35705 | 0.3947 | 0.44202 | 0.49692 | 0.5557 | 0.61245 | 0.65663 | 0.67796 | 0.66767 | 0.62387 | 0.55258 | 0.46738 | 0.38017 | 0.30196 | 0.23564 | 0.18477 | 0.14422 | 0.11363 | 0.087151 | 0.069446 | 0.054131 | 0.039552 | 0.03036 | 0.022405
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
135.z | 0.2096 | 0.22054 | 0.23152 | 0.24243 | 0.2538 | 0.2645 | 0.27437 | 0.28534 | 0.28854 | 0.29617 | 0.29907 | 0.30222 | 0.30213 | 0.30143 | 0.29916 | 0.29589 | 0.29181 | 0.28961 | 0.28866 | 0.29101 | 0.29708 | 0.30917 | 0.32872 | 0.35705 | 0.3947 | 0.44202 | 0.49692 | 0.5557 | 0.61245 | 0.65663 | 0.67796 | 0.66767 | 0.62387 | 0.55258 | 0.46738 | 0.38017 | 0.30196 | 0.23564 | 0.18477 | 0.14422 | 0.11363 | 0.087151 | 0.069446 | 0.054131 | 0.039552 | 0.03036 | 0.022405
(10 Hz)
(0.2 MHz)
0.2 MHz
10 Hz
Notes: D:\PhD\Assignment\Conference ppt\Plot used in smi 2019 Conference_v1
Process log for Sugar
Data spectrum_Real Capa-S5 (135z, 4.5 h, -43.9 C)
Data spectrum_Img Capa-S5 (135z, 4.5 h, -43.9 C)

## Slide 23
An example : 5 %w/v sucrose in 0.55 %w/v NaCl

## Slide 24
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
0.2 MHz
10 Hz
TC4
Notes: 45
img5935
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 25
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
TC4
Notes: 46
Im
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)
g5936

## Slide 26
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
TC4
Notes: 47
Img5937
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 27
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
TC4
Notes: 48
img5938
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 28
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
TC4
Notes: 49
5939
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 29
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
TC4
Notes: 50
img5940
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 30
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
TC4
Notes: 51
Img5941
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 31
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
TC4
Notes: 52
img5942
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 32
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
TC4
Notes: 53
Img5943
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 33
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
Ice nucleation
TC4
Notes: 54
img5944
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 34
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
TC4
Ice nucleation
Notes: 55
img5945
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 35
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
TC4
Ice nucleation
Notes: 56
img5946
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 36
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
TC4
Ice nucleation
Notes: 57
img5947
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 37
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
TC4
Ice nucleation
Notes: 58
img5948
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 38
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
TC4
Ice nucleation
Notes: 59
img5949
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 39
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
TC4
Ice nucleation
Notes: 60
img5950
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 40
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
TC4
Ice nucleation
Notes: 62
Use Img5951 because IMG5952 record 1 minute late
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 41
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
TC4
Ice nucleation
Notes: 45
img5935
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 42
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
TC4
Ice nucleation
Notes: 63
Img5952 (picture record faster than spectrum 1 minute)
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 43
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
TC4
Ice nucleation
Notes: 64
Img5953 (picture record faster than spectrum 1 minute)
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 44
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
TC4
Ice nucleation
Notes: 65
Img5954 (picture record faster than spectrum 1 minute)
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 45
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
TC4
Ice nucleation
Notes: 66
Img5955 (picture record faster than spectrum 1 minute)
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 46
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
TC4
Ice nucleation
Notes: 67
Img5956 (picture record faster than spectrum 1 minute)
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 47
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
TC4
Ice nucleation
Notes: 68
Img5957 (picture record faster than spectrum 1 minute)
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 48
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Ice nucleation
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
TC4
Ice nucleation
Notes: 69
Img5958 (picture record faster than spectrum 1 minute)
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 49
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
TC4
Ice nucleation
Notes: 70
Img5959 (picture record faster than spectrum 1 minute)
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 50
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
TC4
Ice nucleation
Notes: 71
Img5960 (picture record faster than spectrum 1 minute)
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 51
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
TC4
Ice nucleation
Notes: 72
Img5961 (picture record faster than spectrum 1 minute)
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 52
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
TC4
Ice nucleation
Notes: 73
Img5962(picture record faster than spectrum 1 minute)
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 53
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
TC4
Ice nucleation
Notes: 74
Img5963 (picture record faster than spectrum 1 minute)
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 54
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
TC4
Ice nucleation
Notes: 75
Img5964 (picture record faster than spectrum 1 minute)
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 55
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
Solidification end point
TC4
Ice nucleation
Notes: 76
Img5965 (picture record faster than spectrum 1 minute)
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 56
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Solidification end point
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
TC4
Ice nucleation
Notes: 77
Img5966 (picture record faster than spectrum 1 minute)
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 57
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Solidification end point
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
TC4
Ice nucleation
Notes: 78
Img5967 (picture record faster than spectrum 1 minute)
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 58
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Solidification end point
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
Temperature stabilization
(loss of excessive heat)
Sensitive
Sensitive
TC4
Ice nucleation
Notes: 84
Img5973 (picture record faster than spectrum 1 minute)
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 59
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Solidification end point
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
TC4
Ice nucleation
Temperature stabilization
(loss of excessive heat)
Sensitive
Sensitive
Notes: 90
Img5979 (picture record faster than spectrum 1 minute)
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 60
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Solidification end point
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
TC4
Ice nucleation
Temperature stabilization
(loss of excessive heat)
Sensitive
Sensitive
Notes: 105
Img5994 (picture record faster than spectrum 1 minute)
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 61
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Solidification end point
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
TC4
Ice nucleation
Temperature stabilization
(loss of excessive heat)
Sensitive
Sensitive
Notes: 120
Img6009 (picture record faster than spectrum 1 minute)
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 62
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 0.24919 | 0.24719 | 0.24829 | 0.2523 | 0.25961 | 0.26809 | 0.28004 | 0.28377 | 0.29969 | 0.30728 | 0.31219 | 0.3131 | 0.30662 | 0.29798 | 0.28516 | 0.26464 | 0.24389 | 0.22297 | 0.20368 | 0.18329 | 0.1665 | 0.15192 | 0.13804 | 0.12475 | 0.11408 | 0.10396 | 0.095976 | 0.087985 | 0.079948 | 0.073559 | 0.067931 | 0.063688 | 0.05842 | 0.053689 | 0.050521 | 0.046781 | 0.044054 | 0.04288 | 0.042284 | 0.040497 | 0.04309 | 0.040869 | 0.038748 | 0.032321 | 0.026731 | 0.023775 | 0.015772
31.z | 0.22514 | 0.22265 | 0.22334 | 0.2266 | 0.23444 | 0.24479 | 0.26048 | 0.2756 | 0.30149 | 0.3249 | 0.33464 | 0.33806 | 0.33508 | 0.32185 | 0.30528 | 0.28098 | 0.25534 | 0.22985 | 0.20762 | 0.18666 | 0.16772 | 0.15108 | 0.13701 | 0.12445 | 0.11298 | 0.10297 | 0.094565 | 0.08623 | 0.079124 | 0.072071 | 0.066195 | 0.061555 | 0.056722 | 0.051837 | 0.048657 | 0.045855 | 0.042536 | 0.041531 | 0.040471 | 0.039511 | 0.037294 | 0.038526 | 0.035781 | 0.030191 | 0.023156 | 0.017736 | -0.060078
32.z | 0.24101 | 0.24143 | 0.2451 | 0.25024 | 0.25876 | 0.26853 | 0.28122 | 0.28711 | 0.29853 | 0.30617 | 0.31161 | 0.30847 | 0.30281 | 0.29162 | 0.27497 | 0.25649 | 0.23302 | 0.21265 | 0.19468 | 0.17396 | 0.15736 | 0.14281 | 0.12976 | 0.11827 | 0.10815 | 0.098973 | 0.091153 | 0.083395 | 0.076447 | 0.071919 | 0.066851 | 0.062359 | 0.057358 | 0.052931 | 0.049287 | 0.046437 | 0.043556 | 0.042163 | 0.041825 | 0.041231 | 0.043449 | 0.041486 | 0.039265 | 0.033334 | 0.027682 | 0.02486 | 0.017592
33.z | 0.23949 | 0.24261 | 0.24653 | 0.25252 | 0.2617 | 0.27222 | 0.28242 | 0.29331 | 0.30102 | 0.30755 | 0.3092 | 0.30641 | 0.30123 | 0.28818 | 0.27324 | 0.25101 | 0.22843 | 0.20849 | 0.18842 | 0.17 | 0.15316 | 0.14012 | 0.12652 | 0.11523 | 0.10609 | 0.096325 | 0.089031 | 0.08174 | 0.07514 | 0.069465 | 0.064461 | 0.060026 | 0.055505 | 0.051925 | 0.048412 | 0.045745 | 0.043768 | 0.042969 | 0.041472 | 0.041061 | 0.043904 | 0.04173 | 0.039322 | 0.033312 | 0.027496 | 0.024588 | 0.01769
34.z | 0.237 | 0.2391 | 0.24429 | 0.25166 | 0.26176 | 0.27355 | 0.28588 | 0.29347 | 0.30114 | 0.3059 | 0.30961 | 0.30496 | 0.29519 | 0.28325 | 0.26699 | 0.24552 | 0.22278 | 0.20241 | 0.18216 | 0.16379 | 0.14909 | 0.13478 | 0.12319 | 0.11206 | 0.10228 | 0.093792 | 0.086544 | 0.079676 | 0.073191 | 0.068636 | 0.062664 | 0.058449 | 0.054104 | 0.050335 | 0.047712 | 0.044032 | 0.041983 | 0.040664 | 0.039579 | 0.038649 | 0.041285 | 0.039218 | 0.037054 | 0.031136 | 0.025692 | 0.023261 | 0.015007
35.z | 0.23551 | 0.2388 | 0.24443 | 0.25176 | 0.26209 | 0.27378 | 0.28462 | 0.29105 | 0.29973 | 0.30673 | 0.30487 | 0.29973 | 0.29162 | 0.27802 | 0.26247 | 0.2405 | 0.21794 | 0.19739 | 0.17846 | 0.16044 | 0.14463 | 0.13082 | 0.11925 | 0.10922 | 0.099687 | 0.091967 | 0.085296 | 0.078667 | 0.071893 | 0.066741 | 0.062246 | 0.058176 | 0.053776 | 0.049574 | 0.046948 | 0.043707 | 0.041123 | 0.040346 | 0.039157 | 0.038299 | 0.040416 | 0.038983 | 0.036493 | 0.030626 | 0.025032 | 0.022427 | 0.014581
36.z | 0.23398 | 0.23852 | 0.24439 | 0.25272 | 0.2634 | 0.27434 | 0.28573 | 0.2884 | 0.30106 | 0.30394 | 0.30429 | 0.29915 | 0.2898 | 0.27407 | 0.25521 | 0.23554 | 0.21246 | 0.19386 | 0.17336 | 0.15535 | 0.14225 | 0.12688 | 0.11583 | 0.10595 | 0.097228 | 0.088882 | 0.082819 | 0.076155 | 0.070305 | 0.067208 | 0.060697 | 0.057113 | 0.053072 | 0.049311 | 0.044737 | 0.042675 | 0.040302 | 0.039335 | 0.038424 | 0.037848 | 0.040278 | 0.038218 | 0.036378 | 0.030473 | 0.024499 | 0.020882 | -0.055076
37.z | 0.2334 | 0.23859 | 0.2456 | 0.25413 | 0.26496 | 0.27593 | 0.28717 | 0.29404 | 0.30019 | 0.30392 | 0.30157 | 0.29465 | 0.28509 | 0.26832 | 0.25178 | 0.22909 | 0.20776 | 0.1868 | 0.16817 | 0.15146 | 0.13428 | 0.1241 | 0.11226 | 0.10297 | 0.093851 | 0.086592 | 0.080411 | 0.074453 | 0.068333 | 0.063631 | 0.059477 | 0.055152 | 0.051609 | 0.047527 | 0.045422 | 0.043348 | 0.040635 | 0.039874 | 0.038773 | 0.03801 | 0.040597 | 0.039122 | 0.037147 | 0.031583 | 0.026014 | 0.024283 | 0.016493
38.z | 0.23336 | 0.23968 | 0.24751 | 0.25646 | 0.26802 | 0.27871 | 0.28848 | 0.28954 | 0.30257 | 0.30344 | 0.30012 | 0.29124 | 0.28017 | 0.26408 | 0.24726 | 0.22231 | 0.20094 | 0.18075 | 0.16242 | 0.14745 | 0.13125 | 0.1202 | 0.10858 | 0.099321 | 0.091296 | 0.084358 | 0.078604 | 0.072713 | 0.06705 | 0.062578 | 0.058084 | 0.054152 | 0.050844 | 0.047104 | 0.043831 | 0.041753 | 0.039927 | 0.039158 | 0.038399 | 0.037894 | 0.039851 | 0.038348 | 0.036503 | 0.031014 | 0.025091 | 0.022825 | 0.015263
39.z | 0.23295 | 0.24014 | 0.24838 | 0.25803 | 0.26991 | 0.28159 | 0.29035 | 0.29706 | 0.30144 | 0.30301 | 0.29753 | 0.28847 | 0.27502 | 0.25981 | 0.23946 | 0.21844 | 0.1963 | 0.17629 | 0.15884 | 0.14181 | 0.12771 | 0.11635 | 0.10515 | 0.096226 | 0.088241 | 0.081268 | 0.075666 | 0.069711 | 0.064576 | 0.060939 | 0.055841 | 0.052773 | 0.049076 | 0.046282 | 0.041791 | 0.040066 | 0.038735 | 0.037638 | 0.036624 | 0.035696 | 0.03862 | 0.036788 | 0.034775 | 0.029352 | 0.023644 | 0.020761 | 0.013853
40.z | 0.23293 | 0.2408 | 0.24982 | 0.25974 | 0.27155 | 0.28164 | 0.29197 | 0.29997 | 0.29923 | 0.29992 | 0.29287 | 0.28378 | 0.27247 | 0.25388 | 0.23605 | 0.2152 | 0.19042 | 0.17097 | 0.15392 | 0.13812 | 0.12405 | 0.11236 | 0.10304 | 0.093938 | 0.087477 | 0.080141 | 0.074644 | 0.06917 | 0.064389 | 0.059586 | 0.055493 | 0.052518 | 0.048781 | 0.045342 | 0.042149 | 0.040196 | 0.038713 | 0.037506 | 0.036961 | 0.03556 | 0.039085 | 0.037688 | 0.035844 | 0.030717 | 0.025446 | 0.023859 | 0.01724
41.z | 0.23337 | 0.24156 | 0.25131 | 0.26156 | 0.27342 | 0.28376 | 0.29335 | 0.29052 | 0.30024 | 0.29719 | 0.29255 | 0.28178 | 0.26586 | 0.24989 | 0.22824 | 0.20799 | 0.18544 | 0.16722 | 0.14916 | 0.13577 | 0.12096 | 0.11021 | 0.10006 | 0.092363 | 0.084529 | 0.078263 | 0.073283 | 0.067667 | 0.062528 | 0.058535 | 0.055006 | 0.051097 | 0.048062 | 0.044696 | 0.043785 | 0.040851 | 0.038477 | 0.038129 | 0.036529 | 0.035617 | 0.03905 | 0.037377 | 0.035834 | 0.031028 | 0.02608 | 0.024162 | 0.018598
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
0.2 MHz
10 Hz
Solidification end point
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625
30.z | 3.3814 | 3.3542 | 3.3242 | 3.2946 | 3.2623 | 3.2279 | 3.187 | 3.1433 | 3.099 | 3.051 | 2.9956 | 2.9439 | 2.89 | 2.8384 | 2.79 | 2.7428 | 2.7041 | 2.6679 | 2.6363 | 2.6103 | 2.5874 | 2.5664 | 2.5476 | 2.5309 | 2.5171 | 2.5026 | 2.4899 | 2.479 | 2.4691 | 2.4599 | 2.4506 | 2.4418 | 2.4346 | 2.428 | 2.4216 | 2.4154 | 2.4105 | 2.4035 | 2.3969 | 2.3919 | 2.3837 | 2.3759 | 2.369 | 2.3584 | 2.3473 | 2.3439 | 2.3412
31.z | 3.3842 | 3.3627 | 3.337 | 3.3129 | 3.2883 | 3.2604 | 3.2295 | 3.1893 | 3.1571 | 3.104 | 3.0471 | 2.9856 | 2.9235 | 2.8659 | 2.8117 | 2.7649 | 2.7233 | 2.6857 | 2.6561 | 2.6303 | 2.6074 | 2.5868 | 2.5697 | 2.5536 | 2.5389 | 2.5256 | 2.5135 | 2.5035 | 2.4934 | 2.4837 | 2.4748 | 2.467 | 2.4602 | 2.4535 | 2.4474 | 2.4405 | 2.437 | 2.4307 | 2.4241 | 2.4183 | 2.4099 | 2.402 | 2.3931 | 2.3817 | 2.3693 | 2.3635 | 2.3129
32.z | 3.3563 | 3.3296 | 3.2988 | 3.2703 | 3.2362 | 3.202 | 3.1604 | 3.117 | 3.0698 | 3.021 | 2.9679 | 2.9141 | 2.8617 | 2.8092 | 2.7629 | 2.7194 | 2.6812 | 2.6481 | 2.6203 | 2.5938 | 2.5724 | 2.5529 | 2.536 | 2.5204 | 2.5057 | 2.4932 | 2.4814 | 2.471 | 2.4613 | 2.4524 | 2.4429 | 2.4348 | 2.4276 | 2.4204 | 2.414 | 2.4078 | 2.4042 | 2.3978 | 2.3921 | 2.3856 | 2.3777 | 2.3696 | 2.3618 | 2.3512 | 2.3402 | 2.3378 | 2.3341
33.z | 3.3488 | 3.3217 | 3.2917 | 3.2602 | 3.2269 | 3.1892 | 3.1475 | 3.1109 | 3.0552 | 3.0052 | 2.9523 | 2.898 | 2.8456 | 2.7963 | 2.7486 | 2.7084 | 2.6708 | 2.638 | 2.6107 | 2.5863 | 2.5656 | 2.5468 | 2.5301 | 2.5147 | 2.5005 | 2.4889 | 2.4771 | 2.4671 | 2.4575 | 2.4484 | 2.4401 | 2.4319 | 2.4249 | 2.4181 | 2.4113 | 2.4059 | 2.402 | 2.3939 | 2.3898 | 2.3838 | 2.3755 | 2.3666 | 2.3596 | 2.3499 | 2.3382 | 2.3362 | 2.3328
34.z | 3.3342 | 3.3083 | 3.2782 | 3.2474 | 3.2137 | 3.1764 | 3.135 | 3.0956 | 3.0387 | 2.9902 | 2.9352 | 2.8831 | 2.8299 | 2.7808 | 2.7364 | 2.6955 | 2.6598 | 2.6281 | 2.6019 | 2.5762 | 2.5582 | 2.5402 | 2.5236 | 2.5093 | 2.4953 | 2.4838 | 2.4725 | 2.4628 | 2.4533 | 2.4445 | 2.436 | 2.4284 | 2.4216 | 2.4143 | 2.4085 | 2.4028 | 2.3978 | 2.3918 | 2.3862 | 2.3806 | 2.3725 | 2.3646 | 2.3581 | 2.3476 | 2.3375 | 2.3352 | 2.3319
35.z | 3.3233 | 3.2956 | 3.2648 | 3.2336 | 3.1992 | 3.1602 | 3.1182 | 3.0757 | 3.0242 | 2.9736 | 2.9187 | 2.8672 | 2.8159 | 2.7676 | 2.7238 | 2.6843 | 2.6492 | 2.6196 | 2.5928 | 2.5705 | 2.5514 | 2.5333 | 2.5183 | 2.504 | 2.49 | 2.4791 | 2.4681 | 2.4583 | 2.4494 | 2.4407 | 2.4326 | 2.4249 | 2.418 | 2.4115 | 2.4056 | 2.399 | 2.3943 | 2.388 | 2.3824 | 2.3769 | 2.3676 | 2.3609 | 2.3547 | 2.3449 | 2.3343 | 2.3321 | 2.3291
36.z | 3.3122 | 3.2847 | 3.2536 | 3.2212 | 3.186 | 3.1476 | 3.1028 | 3.0631 | 3.0067 | 2.9566 | 2.9024 | 2.8512 | 2.8023 | 2.7519 | 2.7089 | 2.6718 | 2.6385 | 2.6087 | 2.5846 | 2.5631 | 2.5445 | 2.5271 | 2.5125 | 2.4988 | 2.4852 | 2.4747 | 2.464 | 2.4546 | 2.4456 | 2.4377 | 2.4285 | 2.4212 | 2.4143 | 2.4075 | 2.4026 | 2.3965 | 2.3922 | 2.3844 | 2.3808 | 2.3751 | 2.3662 | 2.3582 | 2.3517 | 2.3432 | 2.3313 | 2.326 | 2.2779
37.z | 3.3005 | 3.2726 | 3.241 | 3.2079 | 3.1721 | 3.1321 | 3.088 | 3.0349 | 2.9924 | 2.94 | 2.8873 | 2.8359 | 2.7857 | 2.7393 | 2.6982 | 2.6602 | 2.6285 | 2.6002 | 2.5761 | 2.5554 | 2.5366 | 2.521 | 2.5067 | 2.4938 | 2.481 | 2.4702 | 2.4598 | 2.4505 | 2.4417 | 2.4332 | 2.4249 | 2.4182 | 2.4117 | 2.4046 | 2.3955 | 2.3908 | 2.3863 | 2.3812 | 2.3761 | 2.3704 | 2.3619 | 2.3553 | 2.3486 | 2.34 | 2.3284 | 2.3276 | 2.3245
38.z | 3.2884 | 3.2597 | 3.2281 | 3.1937 | 3.1569 | 3.1167 | 3.0715 | 3.0168 | 2.9731 | 2.9198 | 2.8703 | 2.8164 | 2.7677 | 2.7227 | 2.6827 | 2.6469 | 2.6164 | 2.5896 | 2.5665 | 2.5476 | 2.5296 | 2.5142 | 2.5012 | 2.4884 | 2.4764 | 2.4653 | 2.4564 | 2.4467 | 2.4382 | 2.4303 | 2.4216 | 2.4149 | 2.4078 | 2.401 | 2.396 | 2.3897 | 2.3859 | 2.3794 | 2.3738 | 2.368 | 2.3598 | 2.3525 | 2.3451 | 2.3375 | 2.3263 | 2.325 | 2.3209
39.z | 3.2769 | 3.2483 | 3.2161 | 3.1809 | 3.1437 | 3.1012 | 3.0552 | 3.0057 | 2.9581 | 2.9045 | 2.853 | 2.801 | 2.7523 | 2.7093 | 2.6696 | 2.6354 | 2.6059 | 2.5796 | 2.5579 | 2.5382 | 2.5217 | 2.5067 | 2.4938 | 2.4814 | 2.4697 | 2.4591 | 2.4493 | 2.441 | 2.4325 | 2.4221 | 2.417 | 2.4104 | 2.4038 | 2.3968 | 2.3925 | 2.3871 | 2.3834 | 2.3758 | 2.3718 | 2.3662 | 2.3577 | 2.3507 | 2.3438 | 2.3356 | 2.3234 | 2.3228 | 2.3197
40.z | 3.2653 | 3.2357 | 3.2025 | 3.1671 | 3.1288 | 3.0847 | 3.0398 | 2.996 | 2.9396 | 2.8881 | 2.8349 | 2.7862 | 2.7389 | 2.6954 | 2.656 | 2.6241 | 2.5961 | 2.5704 | 2.549 | 2.5313 | 2.5144 | 2.5012 | 2.4881 | 2.4757 | 2.4644 | 2.4541 | 2.4441 | 2.436 | 2.428 | 2.4194 | 2.4117 | 2.4052 | 2.3986 | 2.3927 | 2.3873 | 2.3813 | 2.3775 | 2.3714 | 2.3649 | 2.3606 | 2.3526 | 2.3456 | 2.3387 | 2.3306 | 2.3196 | 2.3179 | 2.3145
41.z | 3.2542 | 3.2244 | 3.1909 | 3.1543 | 3.1153 | 3.0719 | 3.0235 | 2.9794 | 2.9239 | 2.8729 | 2.8213 | 2.7721 | 2.7242 | 2.6843 | 2.6455 | 2.6146 | 2.5871 | 2.5629 | 2.5425 | 2.5238 | 2.5097 | 2.4949 | 2.483 | 2.4722 | 2.4604 | 2.4501 | 2.4406 | 2.4325 | 2.4241 | 2.4162 | 2.4086 | 2.4023 | 2.3956 | 2.3893 | 2.3812 | 2.3759 | 2.3714 | 2.3667 | 2.3613 | 2.3554 | 2.3483 | 2.3428 | 2.3361 | 2.3278 | 2.3166 | 2.3164 | 2.3148
TC4
Ice nucleation
Temperature stabilization
(loss of excessive heat)
Sensitive
Sensitive
Notes: 135
Img6024 (picture record faster than spectrum 1 minute)
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 63
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
Nucleation point from TC
Notes: E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 64
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
Nucleation point from TC
Sensitive to nucleation
Notes: E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 65
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
Nucleation point from TC
Solidification
Sensitive to nucleation
Notes: E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 66
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
Nucleation point from TC
Solidification
Good indication for temperature
Notes: E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 67
Freezing Step - 5%Sucrose+0.55%NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Temp. / °C
Chart: C″PEAK /pF
Chart: Log FPEAK
Nucleation point from TC
Solidification
Sensitive to nucleation
Good indication for temperature
Notes: E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 68
Chart: C′(10 Hz)/pF
Chart: Time/h
Chart: Log FPEAK
Chart: Temp.
Chart: C″PEAK /pF
High conductive sample (HC)
Chart: Temp.
Chart: C″PEAK /pF
Chart: Log FPEAK
Chart: C′(10 Hz)/pF
Chart: Time/h
Low conductive sample (LC)
Stage | TVIS Parameters | 
 | LC | HC
Nucleation |  | 
Temperature indication |  | 
Solidification | (0.2MHz) | (0.2MHz)
Stage | TVIS Parameters | 
 | LC | HC
Nucleation |  | 
Temperature indication |  | 
Solidification |  | 
Temperature indication
Temperature indication
Temperature indication
Temperature indication
Notes: E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 69
TVIS Applications
Determination of Ice Nucleation Temperature ()

## Slide 70: Ice Nucleation Temperature
Chart: Time / h
TVIS record every 120 seconds
(2nd spectrum start after 1st spectrum 10 seconds)
Thermocouple
TC at the middle of solution bounded within electrode region
(defined as Temperature node)
TVIS vial
TC vial
Thermocouple position
Freezing step
Shelf Temperature
Freezing from 20 °C to -45 °C with 0.5 °C/min
Notes: E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 71: Ice Nucleation Temperature
Chart: Temperature/ °C
In case the TVIS vial nucleates before TC vial, the nucleation temperature in the TVIS vial can be inferred directly from TC temperatures in the nearest neighbor vials
However, if TVIS vial nucleates later than TC vial, the nucleation temperature can be predicted by fitting a curve to the plot of the average temperature from thermocouple vials against TVIS parameter (i.e. )
In this presentation, the ice nucleation temperature of sample (5 %w/v sucrose in 0.55 %w/v NaCl) extrapolate from fitting because TVIS vials form ice after TC vials
Thermocouple
Shelf Temperature
Chart: Time / h
Notes: E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 72: Ice Nucleation Temperature
In case the TVIS vial nucleates before TC vial, the nucleation temperature in the TVIS vial can be inferred directly from TC temperatures in the nearest neighbor vials
However, if TVIS vial nucleates later than TC vial, the nucleation temperature can be predicted by fitting a curve to the plot of the average temperature from thermocouple vials against TVIS parameter (i.e. )
In this presentation, the ice nucleation temperature of sample (5 %w/v sucrose in 0.55 %w/v NaCl) extrapolate from fitting because TVIS vials form ice after TC vials
Thermocouple
Shelf Temperature
Chart: Time / h
Calibration period
Chart: Temperature/ °C
Notes: E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 73: Ice Nucleation Temperature
Chart: C′(10 Hz) / pF
| 3.4845039999999998 | 3.4830739999999998 | 3.4814319999999999 | 3.4797609999999999 | 3.4789569999999999 | 3.4743539999999999 | 3.4654280000000002 | 3.4595660000000001 | 3.4469810000000001 | 3.4385810000000001 | 3.4296760000000002 | 3.4200349999999999 | 3.4120300000000001 | 3.399162 | 3.3921489999999999 | 3.3813659999999999 | 3.3841489999999999 | 3.3562630000000002 | 3.3488020000000001 | 3.3342010000000002 | 3.3232460000000001 | 3.3121749999999999 | 3.3004530000000001 | 3.28844 | 3.2769710000000001 | 3.265301 | 3.2542140000000002 | 3.241895 | 3.2303839999999999 | 3.2178960000000001 | 3.200088 | 3.1839759999999999 | 3.1683400000000002 | 3.1513969999999998 | 3.135875 | 3.1190220000000002 | 3.1030690000000001 | 3.0870739999999999
Tn from C(10 Hz)-S5 | 19.3 | 19.3 | 19 | 18.6 | 18 | 17.4 | 16.7 | 16.1 | 15.3 | 14.6 | 13.7 | 13 | 12.1 | 11.2 | 10.4 | 9.5 | 8.7 | 7.8 | 6.9 | 6.1 | 5.1 | 4.3 | 3.4 | 2.6 | 1.8 | 1 | 0.2 | -0.5 | -1.3 | -2.4 | -3.6 | -4.7 | -5.8 | -6.7 | -7.7 | -8.7 | -9.7 | -10.6
Tn from C(10 Hz)-S5 Trendline | 0.2 | -0.5 | -1.3 | -2.4 | -3.6 | -4.7 | -5.8 | -6.7 | -7.7 | -8.7 | -9.7 | -10.6
In case the TVIS vial nucleates before TC vial, the nucleation temperature in the TVIS vial can be inferred directly from TC temperatures in the nearest neighbor vials
However, if TVIS vial nucleates later than TC vial, the nucleation temperature can be predicted by fitting a curve to the plot of the average temperature from thermocouple vials against TVIS parameter (i.e. )
In this presentation, the ice nucleation temperature of sample (5 %w/v sucrose in 0.55 %w/v NaCl) extrapolate from fitting because TVIS vials form ice after TC vials
Chart: Temperature/ °C
Thermocouple
Shelf Temperature
Chart: Time / h
Temperature prediction
Notes: E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 74
TVIS Applications
Ice crystallization time

## Slide 75: Temperature Profile of Frozen Water
Re-heating
−18 °C
−38 °C
0.2 MHz
10 Hz
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 4.4 h | 0.084931 | 0.086522 | 0.084416 | 0.083584 | 0.083174 | 0.083511 | 0.083751 | 0.08481 | 0.086097 | 0.087016 | 0.08765 | 0.088181 | 0.088496 | 0.088762 | 0.089099 | 0.089472 | 0.089925 | 0.090513 | 0.091358 | 0.092525 | 0.094183 | 0.096518 | 0.099835 | 0.10479 | 0.11237 | 0.12448 | 0.14398 | 0.17467 | 0.22057 | 0.28526 | 0.36955 | 0.47045 | 0.58177 | 0.69523 | 0.80256 | 0.89741 | 0.97772 | 1.0421 | 1.0945 | 1.1331 | 1.1623 | 1.1841 | 1.1989 | 1.2068 | 1.2187 | 1.2253 | 1.2327 | 1.237 | 1.2399 | 1.2486 | 1.255
Time 4.5 h | 0.085366 | 0.086803 | 0.084615 | 0.083791 | 0.08325 | 0.083655 | 0.08388 | 0.084857 | 0.086134 | 0.087029 | 0.087739 | 0.088261 | 0.088595 | 0.088921 | 0.089249 | 0.089597 | 0.090054 | 0.090652 | 0.091444 | 0.092572 | 0.094147 | 0.096469 | 0.1002 | 0.10625 | 0.11628 | 0.13294 | 0.16019 | 0.20224 | 0.26295 | 0.34453 | 0.4451 | 0.55868 | 0.67675 | 0.78985 | 0.89073 | 0.97528 | 1.0447 | 1.0963 | 1.1358 | 1.1645 | 1.186 | 1.201 | 1.2113 | 1.2236 | 1.2271 | 1.2334 | 1.2393 | 1.2455 | 1.2503 | 1.2572 | 1.2649
Time 4.6 h | 0.085421 | 0.086928 | 0.084724 | 0.083845 | 0.083347 | 0.08379 | 0.084043 | 0.085026 | 0.086304 | 0.087178 | 0.087815 | 0.088397 | 0.088756 | 0.089066 | 0.089366 | 0.089722 | 0.09011 | 0.090631 | 0.091373 | 0.092445 | 0.094129 | 0.096729 | 0.10147 | 0.10987 | 0.1245 | 0.14913 | 0.18836 | 0.24664 | 0.32654 | 0.42723 | 0.54333 | 0.66518 | 0.78294 | 0.88848 | 0.97692 | 1.0466 | 1.0994 | 1.1385 | 1.1666 | 1.1884 | 1.2027 | 1.2152 | 1.2223 | 1.2362 | 1.2355 | 1.2434 | 1.246 | 1.2536 | 1.263 | 1.2718 | 1.2824
Time 4.7 h | 0.085219 | 0.053673 | 0.07561 | 0.08397 | 0.083322 | 0.083877 | 0.08413 | 0.085153 | 0.086388 | 0.087288 | 0.087947 | 0.088541 | 0.088887 | 0.089203 | 0.089467 | 0.089732 | 0.090058 | 0.090474 | 0.091108 | 0.092199 | 0.094218 | 0.097969 | 0.10503 | 0.11817 | 0.14112 | 0.17863 | 0.23583 | 0.31561 | 0.41746 | 0.5363 | 0.66212 | 0.78332 | 0.89159 | 0.98177 | 1.0528 | 1.1059 | 1.144 | 1.1727 | 1.1928 | 1.2063 | 1.2173 | 1.2263 | 1.2337 | 1.2352 | 1.2458 | 1.2532 | 1.2583 | 1.2687 | 1.2782 | 1.2933 | 1.3076
Time 4.8 h | 0.085767 | 0.086823 | 0.084565 | 0.083702 | 0.083251 | 0.083937 | 0.084213 | 0.085213 | 0.08646 | 0.08732 | 0.087939 | 0.088605 | 0.088971 | 0.089203 | 0.089353 | 0.089534 | 0.089704 | 0.089999 | 0.090613 | 0.092049 | 0.095196 | 0.10173 | 0.11443 | 0.13773 | 0.1764 | 0.23592 | 0.3197 | 0.42654 | 0.54993 | 0.67941 | 0.80269 | 0.91088 | 0.99931 | 1.0678 | 1.1184 | 1.1539 | 1.1788 | 1.1983 | 1.2124 | 1.2227 | 1.2311 | 1.2384 | 1.2447 | 1.2451 | 1.2597 | 1.2688 | 1.2791 | 1.2911 | 1.3078 | 1.3258 | 1.3483
Time 4.9 h | 0.085491 | 0.087251 | 0.08505 | 0.084128 | 0.083174 | 0.083948 | 0.084267 | 0.085306 | 0.08667 | 0.087534 | 0.088121 | 0.088619 | 0.088926 | 0.089085 | 0.089128 | 0.089112 | 0.089161 | 0.08943 | 0.090419 | 0.093078 | 0.099266 | 0.11187 | 0.13539 | 0.17568 | 0.23771 | 0.32517 | 0.43674 | 0.56448 | 0.69641 | 0.82031 | 0.92766 | 1.014 | 1.08 | 1.1283 | 1.1629 | 1.1868 | 1.205 | 1.2169 | 1.2264 | 1.2348 | 1.2422 | 1.2502 | 1.2567 | 1.2693 | 1.277 | 1.2885 | 1.3045 | 1.3218 | 1.3429 | 1.3705 | 1.4013
Time 5 h | 0.085899 | 0.087566 | 0.085217 | 0.084216 | 0.083061 | 0.084092 | 0.084384 | 0.085508 | 0.086793 | 0.087615 | 0.088068 | 0.088506 | 0.088675 | 0.088682 | 0.088573 | 0.088418 | 0.088473 | 0.089248 | 0.091786 | 0.098085 | 0.11166 | 0.13702 | 0.18042 | 0.24808 | 0.34172 | 0.45928 | 0.59178 | 0.72515 | 0.84771 | 0.95135 | 1.0334 | 1.0952 | 1.14 | 1.172 | 1.1948 | 1.2102 | 1.2225 | 1.2305 | 1.2409 | 1.247 | 1.2556 | 1.2643 | 1.2739 | 1.2804 | 1.3003 | 1.3163 | 1.3401 | 1.3649 | 1.3972 | 1.4316 | 1.4722
Time 5.1 h | 0.086471 | 0.087693 | 0.085343 | 0.084296 | 0.082843 | 0.084117 | 0.084511 | 0.08554 | 0.086831 | 0.08761 | 0.088012 | 0.088277 | 0.088283 | 0.088083 | 0.087726 | 0.087608 | 0.088197 | 0.090659 | 0.097298 | 0.11173 | 0.13971 | 0.18675 | 0.25888 | 0.35922 | 0.48223 | 0.61799 | 0.7525 | 0.87288 | 0.97254 | 1.0503 | 1.1079 | 1.1501 | 1.1795 | 1.2007 | 1.2162 | 1.2273 | 1.2375 | 1.2456 | 1.2527 | 1.2602 | 1.27 | 1.2804 | 1.294 | 1.3078 | 1.3301 | 1.3551 | 1.3841 | 1.417 | 1.4564 | 1.5034 | 1.5541
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 4.4 h | 0.003361 | 0.0040465 | 0.0025946 | 0.0026214 | 0.0044553 | 0.0026323 | 0.0029087 | 0.0037366 | 0.0047149 | 0.0057564 | 0.0073092 | 0.0089323 | 0.011077 | 0.013703 | 0.017072 | 0.021236 | 0.026498 | 0.03301 | 0.041177 | 0.051293 | 0.064055 | 0.079683 | 0.099225 | 0.12384 | 0.15437 | 0.19231 | 0.23844 | 0.29224 | 0.35137 | 0.41125 | 0.46468 | 0.50437 | 0.52462 | 0.52356 | 0.50251 | 0.46708 | 0.42347 | 0.37253 | 0.32299 | 0.27537 | 0.23266 | 0.1978 | 0.16633 | 0.14424 | 0.12252 | 0.10889 | 0.097697 | 0.093226 | 0.09166 | 0.092155 | 0.093069
Time 4.5 h | 0.0034249 | 0.0041377 | 0.0027283 | 0.0027814 | 0.0046887 | 0.0029259 | 0.0032733 | 0.0042363 | 0.0053358 | 0.0065296 | 0.0082746 | 0.010141 | 0.012601 | 0.015587 | 0.01947 | 0.024233 | 0.030264 | 0.037737 | 0.047171 | 0.058901 | 0.073903 | 0.092482 | 0.11581 | 0.14538 | 0.18208 | 0.2274 | 0.28124 | 0.3416 | 0.40418 | 0.46243 | 0.50811 | 0.53414 | 0.53731 | 0.51809 | 0.48197 | 0.43577 | 0.38293 | 0.3311 | 0.28177 | 0.23748 | 0.20077 | 0.16802 | 0.14416 | 0.11848 | 0.11171 | 0.10187 | 0.09506 | 0.092791 | 0.094429 | 0.097904 | 0.10371
Time 4.6 h | 0.0036091 | 0.0043103 | 0.0029276 | 0.0030069 | 0.0049863 | 0.0032919 | 0.0037411 | 0.0048548 | 0.006126 | 0.0075084 | 0.0095269 | 0.011725 | 0.014579 | 0.018062 | 0.02259 | 0.028182 | 0.035264 | 0.044078 | 0.055262 | 0.069282 | 0.08729 | 0.10998 | 0.13862 | 0.17503 | 0.21981 | 0.27381 | 0.33549 | 0.40061 | 0.46249 | 0.51282 | 0.5432 | 0.54939 | 0.53148 | 0.49439 | 0.44517 | 0.39079 | 0.33577 | 0.28667 | 0.23998 | 0.20372 | 0.17181 | 0.14592 | 0.12721 | 0.11818 | 0.10489 | 0.098824 | 0.097908 | 0.099283 | 0.10394 | 0.11202 | 0.12205
Time 4.7 h | 0.0040254 | 0.0023197 | 0.0060162 | 0.0032852 | 0.0053707 | 0.003773 | 0.0043423 | 0.005672 | 0.0071637 | 0.0088006 | 0.011181 | 0.013771 | 0.017191 | 0.021329 | 0.026752 | 0.033478 | 0.04203 | 0.052722 | 0.066401 | 0.083738 | 0.10613 | 0.13436 | 0.17037 | 0.21573 | 0.27035 | 0.33347 | 0.40107 | 0.46603 | 0.51937 | 0.55253 | 0.55996 | 0.54171 | 0.50311 | 0.45097 | 0.39455 | 0.33876 | 0.28887 | 0.24104 | 0.20387 | 0.17339 | 0.14886 | 0.13225 | 0.11753 | 0.11 | 0.10482 | 0.10438 | 0.1074 | 0.1138 | 0.12408 | 0.13639 | 0.14993
Time 4.8 h | 0.0039527 | 0.0047899 | 0.0034912 | 0.0036734 | 0.0059298 | 0.004463 | 0.0052209 | 0.0068624 | 0.0086692 | 0.010687 | 0.013626 | 0.016821 | 0.021041 | 0.026213 | 0.03302 | 0.041432 | 0.052261 | 0.065911 | 0.083539 | 0.10606 | 0.13545 | 0.17252 | 0.21928 | 0.27665 | 0.34237 | 0.41271 | 0.47977 | 0.53353 | 0.56486 | 0.56874 | 0.5461 | 0.50317 | 0.44813 | 0.38872 | 0.33224 | 0.28045 | 0.23905 | 0.20172 | 0.17254 | 0.14998 | 0.13098 | 0.12232 | 0.11386 | 0.11155 | 0.11412 | 0.11812 | 0.12658 | 0.13848 | 0.15361 | 0.17099 | 0.18866
Time 4.9 h | 0.0043188 | 0.0052277 | 0.0039868 | 0.0042288 | 0.0066529 | 0.0053479 | 0.0063485 | 0.0083796 | 0.010604 | 0.013107 | 0.016765 | 0.020805 | 0.02609 | 0.032614 | 0.041224 | 0.052023 | 0.065957 | 0.083695 | 0.10683 | 0.13651 | 0.17523 | 0.22353 | 0.28261 | 0.35128 | 0.42376 | 0.49222 | 0.54583 | 0.5749 | 0.57491 | 0.54814 | 0.50162 | 0.4443 | 0.38384 | 0.32634 | 0.27564 | 0.23272 | 0.19908 | 0.1717 | 0.1495 | 0.13558 | 0.12502 | 0.11938 | 0.11838 | 0.12437 | 0.13062 | 0.13992 | 0.15492 | 0.1723 | 0.19073 | 0.21306 | 0.23377
Time 5 h | 0.0049133 | 0.0058042 | 0.0046338 | 0.0049712 | 0.0076664 | 0.0066031 | 0.0079479 | 0.010533 | 0.013358 | 0.01657 | 0.021298 | 0.026516 | 0.033397 | 0.041936 | 0.053268 | 0.067599 | 0.086277 | 0.11024 | 0.14161 | 0.18183 | 0.2334 | 0.29549 | 0.36662 | 0.44167 | 0.50969 | 0.5605 | 0.58424 | 0.57776 | 0.54458 | 0.49349 | 0.43374 | 0.3731 | 0.31662 | 0.26729 | 0.22557 | 0.19336 | 0.16657 | 0.14984 | 0.13717 | 0.1291 | 0.12498 | 0.12759 | 0.13258 | 0.1444 | 0.15601 | 0.17194 | 0.19151 | 0.21398 | 0.23437 | 0.26066 | 0.28203
Time 5.1 h | 0.0051615 | 0.0065394 | 0.0054847 | 0.0059393 | 0.0090001 | 0.0082333 | 0.010001 | 0.013324 | 0.016914 | 0.021051 | 0.027147 | 0.03396 | 0.042946 | 0.054219 | 0.069319 | 0.088593 | 0.11394 | 0.14657 | 0.18909 | 0.24251 | 0.30812 | 0.38167 | 0.45767 | 0.52576 | 0.57299 | 0.59132 | 0.57824 | 0.5397 | 0.48481 | 0.42344 | 0.36292 | 0.3071 | 0.25963 | 0.22024 | 0.18846 | 0.16573 | 0.14947 | 0.13846 | 0.13032 | 0.13113 | 0.1334 | 0.1422 | 0.15348 | 0.16641 | 0.18887 | 0.2096 | 0.23238 | 0.25699 | 0.28353 | 0.3063 | 0.32649
−38 °C
−18 °C
Chart: Temperature/°C
| -38.4 | -37.299999999999997 | -36.5 | -35.700000000000003 | -34.799999999999997 | -33.799999999999997 | -32.799999999999997 | -32.1 | -31 | -30.2 | -28.9 | -27.9 | -27 | -26 | -25.1 | -24 | -23 | -22 | -21.1 | -20 | -19.100000000000001 | -18.2
Log(Frequency) | 2.7551 | 2.7816 | 2.8077 | 2.8379 | 2.8682 | 2.9001 | 2.9302 | 2.9619 | 2.9942 | 3.0287 | 3.0652 | 3.1011 | 3.139 | 3.177 | 3.2151 | 3.2535 | 3.2935 | 3.3335 | 3.3735 | 3.4135 | 3.4535 | 3.4959
C''PEAK | 0.51695 | 0.51973 | 0.52247 | 0.52612 | 0.52968 | 0.53324 | 0.53671 | 0.54012 | 0.54345 | 0.54669 | 0.54988 | 0.55289 | 0.5558 | 0.55853 | 0.56117 | 0.5636 | 0.56593 | 0.56819 | 0.57022 | 0.57217 | 0.57397 | 0.5757
Chart: Temperature/°C
| -38.4 | -37.299999999999997 | -36.5 | -35.700000000000003 | -34.799999999999997 | -33.799999999999997 | -32.799999999999997 | -32.1 | -31 | -30.2 | -28.9 | -27.9 | -27 | -26 | -25.1 | -24 | -23 | -22 | -21.1 | -20 | -19.100000000000001 | -18.2
Tslice (10 Hz) | 1.255 | 1.2575 | 1.2617 | 1.2649 | 1.2706 | 1.2764 | 1.2824 | 1.2902 | 1.2979 | 1.3076 | 1.3199 | 1.3323 | 1.3483 | 1.3646 | 1.3828 | 1.4013 | 1.425 | 1.4477 | 1.4722 | 1.499 | 1.5263 | 1.5541
Tslice (100 Hz) | 1.1623 | 1.1717 | 1.1777 | 1.186 | 1.1918 | 1.1977 | 1.2027 | 1.2078 | 1.2128 | 1.2173 | 1.2209 | 1.2251 | 1.2311 | 1.2338 | 1.2379 | 1.2422 | 1.2475 | 1.2511 | 1.2556 | 1.2603 | 1.2651 | 1.27
Tslice (0.2MHz) | 0.08481 | 0.084804 | 0.084613 | 0.084857 | 0.08494 | 0.084921 | 0.085026 | 0.085057 | 0.085114 | 0.085153 | 0.08514 | 0.08519 | 0.085213 | 0.085331 | 0.085285 | 0.085306 | 0.085392 | 0.085401 | 0.085507 | 0.085543 | 0.08555 | 0.08554
10 Hz
100 Hz
0.2 MHz
Notes: D:\PhD\Experiment\2016_06_29_Full Shelf Water Study\281016 (Different height from base)\3 mm from base

## Slide 76: Temperature Profile of Frozen Water
Chart: Temperature/°C
| -38.4 | -37.299999999999997 | -36.5 | -35.700000000000003 | -34.799999999999997 | -33.799999999999997 | -32.799999999999997 | -32.1 | -31 | -30.2 | -28.9 | -27.9 | -27 | -26 | -25.1 | -24 | -23 | -22 | -21.1 | -20 | -19.100000000000001 | -18.2
Tslice (10 Hz) | 1 | 1.002 | 1.0053 | 1.0078 | 1.0124 | 1.017 | 1.0218 | 1.028 | 1.0342 | 1.0419 | 1.0517 | 1.0616 | 1.0743 | 1.0873 | 1.1018 | 1.1166 | 1.1355 | 1.1535 | 1.173 | 1.1944 | 1.2161 | 1.2383
Tslice (100 Hz) | 1 | 1.0081 | 1.0133 | 1.0204 | 1.0254 | 1.0305 | 1.0348 | 1.0392 | 1.0435 | 1.0473 | 1.0504 | 1.0541 | 1.0592 | 1.0616 | 1.065 | 1.0688 | 1.0733 | 1.0764 | 1.0803 | 1.0843 | 1.0884 | 1.0927
Tslice (0.2MHz) | 1 | 0.99993 | 0.99767 | 1.0006 | 1.0015 | 1.0013 | 1.0025 | 1.0029 | 1.0036 | 1.004 | 1.0039 | 1.0045 | 1.0048 | 1.0061 | 1.0056 | 1.0058 | 1.0069 | 1.007 | 1.0082 | 1.0086 | 1.0087 | 1.0086
Re-heating
−18 °C
−38 °C
0.2 MHz
10 Hz
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 4.4 h | 0.084931 | 0.086522 | 0.084416 | 0.083584 | 0.083174 | 0.083511 | 0.083751 | 0.08481 | 0.086097 | 0.087016 | 0.08765 | 0.088181 | 0.088496 | 0.088762 | 0.089099 | 0.089472 | 0.089925 | 0.090513 | 0.091358 | 0.092525 | 0.094183 | 0.096518 | 0.099835 | 0.10479 | 0.11237 | 0.12448 | 0.14398 | 0.17467 | 0.22057 | 0.28526 | 0.36955 | 0.47045 | 0.58177 | 0.69523 | 0.80256 | 0.89741 | 0.97772 | 1.0421 | 1.0945 | 1.1331 | 1.1623 | 1.1841 | 1.1989 | 1.2068 | 1.2187 | 1.2253 | 1.2327 | 1.237 | 1.2399 | 1.2486 | 1.255
Time 4.5 h | 0.085366 | 0.086803 | 0.084615 | 0.083791 | 0.08325 | 0.083655 | 0.08388 | 0.084857 | 0.086134 | 0.087029 | 0.087739 | 0.088261 | 0.088595 | 0.088921 | 0.089249 | 0.089597 | 0.090054 | 0.090652 | 0.091444 | 0.092572 | 0.094147 | 0.096469 | 0.1002 | 0.10625 | 0.11628 | 0.13294 | 0.16019 | 0.20224 | 0.26295 | 0.34453 | 0.4451 | 0.55868 | 0.67675 | 0.78985 | 0.89073 | 0.97528 | 1.0447 | 1.0963 | 1.1358 | 1.1645 | 1.186 | 1.201 | 1.2113 | 1.2236 | 1.2271 | 1.2334 | 1.2393 | 1.2455 | 1.2503 | 1.2572 | 1.2649
Time 4.6 h | 0.085421 | 0.086928 | 0.084724 | 0.083845 | 0.083347 | 0.08379 | 0.084043 | 0.085026 | 0.086304 | 0.087178 | 0.087815 | 0.088397 | 0.088756 | 0.089066 | 0.089366 | 0.089722 | 0.09011 | 0.090631 | 0.091373 | 0.092445 | 0.094129 | 0.096729 | 0.10147 | 0.10987 | 0.1245 | 0.14913 | 0.18836 | 0.24664 | 0.32654 | 0.42723 | 0.54333 | 0.66518 | 0.78294 | 0.88848 | 0.97692 | 1.0466 | 1.0994 | 1.1385 | 1.1666 | 1.1884 | 1.2027 | 1.2152 | 1.2223 | 1.2362 | 1.2355 | 1.2434 | 1.246 | 1.2536 | 1.263 | 1.2718 | 1.2824
Time 4.7 h | 0.085219 | 0.053673 | 0.07561 | 0.08397 | 0.083322 | 0.083877 | 0.08413 | 0.085153 | 0.086388 | 0.087288 | 0.087947 | 0.088541 | 0.088887 | 0.089203 | 0.089467 | 0.089732 | 0.090058 | 0.090474 | 0.091108 | 0.092199 | 0.094218 | 0.097969 | 0.10503 | 0.11817 | 0.14112 | 0.17863 | 0.23583 | 0.31561 | 0.41746 | 0.5363 | 0.66212 | 0.78332 | 0.89159 | 0.98177 | 1.0528 | 1.1059 | 1.144 | 1.1727 | 1.1928 | 1.2063 | 1.2173 | 1.2263 | 1.2337 | 1.2352 | 1.2458 | 1.2532 | 1.2583 | 1.2687 | 1.2782 | 1.2933 | 1.3076
Time 4.8 h | 0.085767 | 0.086823 | 0.084565 | 0.083702 | 0.083251 | 0.083937 | 0.084213 | 0.085213 | 0.08646 | 0.08732 | 0.087939 | 0.088605 | 0.088971 | 0.089203 | 0.089353 | 0.089534 | 0.089704 | 0.089999 | 0.090613 | 0.092049 | 0.095196 | 0.10173 | 0.11443 | 0.13773 | 0.1764 | 0.23592 | 0.3197 | 0.42654 | 0.54993 | 0.67941 | 0.80269 | 0.91088 | 0.99931 | 1.0678 | 1.1184 | 1.1539 | 1.1788 | 1.1983 | 1.2124 | 1.2227 | 1.2311 | 1.2384 | 1.2447 | 1.2451 | 1.2597 | 1.2688 | 1.2791 | 1.2911 | 1.3078 | 1.3258 | 1.3483
Time 4.9 h | 0.085491 | 0.087251 | 0.08505 | 0.084128 | 0.083174 | 0.083948 | 0.084267 | 0.085306 | 0.08667 | 0.087534 | 0.088121 | 0.088619 | 0.088926 | 0.089085 | 0.089128 | 0.089112 | 0.089161 | 0.08943 | 0.090419 | 0.093078 | 0.099266 | 0.11187 | 0.13539 | 0.17568 | 0.23771 | 0.32517 | 0.43674 | 0.56448 | 0.69641 | 0.82031 | 0.92766 | 1.014 | 1.08 | 1.1283 | 1.1629 | 1.1868 | 1.205 | 1.2169 | 1.2264 | 1.2348 | 1.2422 | 1.2502 | 1.2567 | 1.2693 | 1.277 | 1.2885 | 1.3045 | 1.3218 | 1.3429 | 1.3705 | 1.4013
Time 5 h | 0.085899 | 0.087566 | 0.085217 | 0.084216 | 0.083061 | 0.084092 | 0.084384 | 0.085508 | 0.086793 | 0.087615 | 0.088068 | 0.088506 | 0.088675 | 0.088682 | 0.088573 | 0.088418 | 0.088473 | 0.089248 | 0.091786 | 0.098085 | 0.11166 | 0.13702 | 0.18042 | 0.24808 | 0.34172 | 0.45928 | 0.59178 | 0.72515 | 0.84771 | 0.95135 | 1.0334 | 1.0952 | 1.14 | 1.172 | 1.1948 | 1.2102 | 1.2225 | 1.2305 | 1.2409 | 1.247 | 1.2556 | 1.2643 | 1.2739 | 1.2804 | 1.3003 | 1.3163 | 1.3401 | 1.3649 | 1.3972 | 1.4316 | 1.4722
Time 5.1 h | 0.086471 | 0.087693 | 0.085343 | 0.084296 | 0.082843 | 0.084117 | 0.084511 | 0.08554 | 0.086831 | 0.08761 | 0.088012 | 0.088277 | 0.088283 | 0.088083 | 0.087726 | 0.087608 | 0.088197 | 0.090659 | 0.097298 | 0.11173 | 0.13971 | 0.18675 | 0.25888 | 0.35922 | 0.48223 | 0.61799 | 0.7525 | 0.87288 | 0.97254 | 1.0503 | 1.1079 | 1.1501 | 1.1795 | 1.2007 | 1.2162 | 1.2273 | 1.2375 | 1.2456 | 1.2527 | 1.2602 | 1.27 | 1.2804 | 1.294 | 1.3078 | 1.3301 | 1.3551 | 1.3841 | 1.417 | 1.4564 | 1.5034 | 1.5541
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 4.4 h | 0.003361 | 0.0040465 | 0.0025946 | 0.0026214 | 0.0044553 | 0.0026323 | 0.0029087 | 0.0037366 | 0.0047149 | 0.0057564 | 0.0073092 | 0.0089323 | 0.011077 | 0.013703 | 0.017072 | 0.021236 | 0.026498 | 0.03301 | 0.041177 | 0.051293 | 0.064055 | 0.079683 | 0.099225 | 0.12384 | 0.15437 | 0.19231 | 0.23844 | 0.29224 | 0.35137 | 0.41125 | 0.46468 | 0.50437 | 0.52462 | 0.52356 | 0.50251 | 0.46708 | 0.42347 | 0.37253 | 0.32299 | 0.27537 | 0.23266 | 0.1978 | 0.16633 | 0.14424 | 0.12252 | 0.10889 | 0.097697 | 0.093226 | 0.09166 | 0.092155 | 0.093069
Time 4.5 h | 0.0034249 | 0.0041377 | 0.0027283 | 0.0027814 | 0.0046887 | 0.0029259 | 0.0032733 | 0.0042363 | 0.0053358 | 0.0065296 | 0.0082746 | 0.010141 | 0.012601 | 0.015587 | 0.01947 | 0.024233 | 0.030264 | 0.037737 | 0.047171 | 0.058901 | 0.073903 | 0.092482 | 0.11581 | 0.14538 | 0.18208 | 0.2274 | 0.28124 | 0.3416 | 0.40418 | 0.46243 | 0.50811 | 0.53414 | 0.53731 | 0.51809 | 0.48197 | 0.43577 | 0.38293 | 0.3311 | 0.28177 | 0.23748 | 0.20077 | 0.16802 | 0.14416 | 0.11848 | 0.11171 | 0.10187 | 0.09506 | 0.092791 | 0.094429 | 0.097904 | 0.10371
Time 4.6 h | 0.0036091 | 0.0043103 | 0.0029276 | 0.0030069 | 0.0049863 | 0.0032919 | 0.0037411 | 0.0048548 | 0.006126 | 0.0075084 | 0.0095269 | 0.011725 | 0.014579 | 0.018062 | 0.02259 | 0.028182 | 0.035264 | 0.044078 | 0.055262 | 0.069282 | 0.08729 | 0.10998 | 0.13862 | 0.17503 | 0.21981 | 0.27381 | 0.33549 | 0.40061 | 0.46249 | 0.51282 | 0.5432 | 0.54939 | 0.53148 | 0.49439 | 0.44517 | 0.39079 | 0.33577 | 0.28667 | 0.23998 | 0.20372 | 0.17181 | 0.14592 | 0.12721 | 0.11818 | 0.10489 | 0.098824 | 0.097908 | 0.099283 | 0.10394 | 0.11202 | 0.12205
Time 4.7 h | 0.0040254 | 0.0023197 | 0.0060162 | 0.0032852 | 0.0053707 | 0.003773 | 0.0043423 | 0.005672 | 0.0071637 | 0.0088006 | 0.011181 | 0.013771 | 0.017191 | 0.021329 | 0.026752 | 0.033478 | 0.04203 | 0.052722 | 0.066401 | 0.083738 | 0.10613 | 0.13436 | 0.17037 | 0.21573 | 0.27035 | 0.33347 | 0.40107 | 0.46603 | 0.51937 | 0.55253 | 0.55996 | 0.54171 | 0.50311 | 0.45097 | 0.39455 | 0.33876 | 0.28887 | 0.24104 | 0.20387 | 0.17339 | 0.14886 | 0.13225 | 0.11753 | 0.11 | 0.10482 | 0.10438 | 0.1074 | 0.1138 | 0.12408 | 0.13639 | 0.14993
Time 4.8 h | 0.0039527 | 0.0047899 | 0.0034912 | 0.0036734 | 0.0059298 | 0.004463 | 0.0052209 | 0.0068624 | 0.0086692 | 0.010687 | 0.013626 | 0.016821 | 0.021041 | 0.026213 | 0.03302 | 0.041432 | 0.052261 | 0.065911 | 0.083539 | 0.10606 | 0.13545 | 0.17252 | 0.21928 | 0.27665 | 0.34237 | 0.41271 | 0.47977 | 0.53353 | 0.56486 | 0.56874 | 0.5461 | 0.50317 | 0.44813 | 0.38872 | 0.33224 | 0.28045 | 0.23905 | 0.20172 | 0.17254 | 0.14998 | 0.13098 | 0.12232 | 0.11386 | 0.11155 | 0.11412 | 0.11812 | 0.12658 | 0.13848 | 0.15361 | 0.17099 | 0.18866
Time 4.9 h | 0.0043188 | 0.0052277 | 0.0039868 | 0.0042288 | 0.0066529 | 0.0053479 | 0.0063485 | 0.0083796 | 0.010604 | 0.013107 | 0.016765 | 0.020805 | 0.02609 | 0.032614 | 0.041224 | 0.052023 | 0.065957 | 0.083695 | 0.10683 | 0.13651 | 0.17523 | 0.22353 | 0.28261 | 0.35128 | 0.42376 | 0.49222 | 0.54583 | 0.5749 | 0.57491 | 0.54814 | 0.50162 | 0.4443 | 0.38384 | 0.32634 | 0.27564 | 0.23272 | 0.19908 | 0.1717 | 0.1495 | 0.13558 | 0.12502 | 0.11938 | 0.11838 | 0.12437 | 0.13062 | 0.13992 | 0.15492 | 0.1723 | 0.19073 | 0.21306 | 0.23377
Time 5 h | 0.0049133 | 0.0058042 | 0.0046338 | 0.0049712 | 0.0076664 | 0.0066031 | 0.0079479 | 0.010533 | 0.013358 | 0.01657 | 0.021298 | 0.026516 | 0.033397 | 0.041936 | 0.053268 | 0.067599 | 0.086277 | 0.11024 | 0.14161 | 0.18183 | 0.2334 | 0.29549 | 0.36662 | 0.44167 | 0.50969 | 0.5605 | 0.58424 | 0.57776 | 0.54458 | 0.49349 | 0.43374 | 0.3731 | 0.31662 | 0.26729 | 0.22557 | 0.19336 | 0.16657 | 0.14984 | 0.13717 | 0.1291 | 0.12498 | 0.12759 | 0.13258 | 0.1444 | 0.15601 | 0.17194 | 0.19151 | 0.21398 | 0.23437 | 0.26066 | 0.28203
Time 5.1 h | 0.0051615 | 0.0065394 | 0.0054847 | 0.0059393 | 0.0090001 | 0.0082333 | 0.010001 | 0.013324 | 0.016914 | 0.021051 | 0.027147 | 0.03396 | 0.042946 | 0.054219 | 0.069319 | 0.088593 | 0.11394 | 0.14657 | 0.18909 | 0.24251 | 0.30812 | 0.38167 | 0.45767 | 0.52576 | 0.57299 | 0.59132 | 0.57824 | 0.5397 | 0.48481 | 0.42344 | 0.36292 | 0.3071 | 0.25963 | 0.22024 | 0.18846 | 0.16573 | 0.14947 | 0.13846 | 0.13032 | 0.13113 | 0.1334 | 0.1422 | 0.15348 | 0.16641 | 0.18887 | 0.2096 | 0.23238 | 0.25699 | 0.28353 | 0.3063 | 0.32649
−38 °C
−18 °C
Chart: Temperature/°C
| -38.4 | -37.299999999999997 | -36.5 | -35.700000000000003 | -34.799999999999997 | -33.799999999999997 | -32.799999999999997 | -32.1 | -31 | -30.2 | -28.9 | -27.9 | -27 | -26 | -25.1 | -24 | -23 | -22 | -21.1 | -20 | -19.100000000000001 | -18.2
Log(Frequency) | 2.7551 | 2.7816 | 2.8077 | 2.8379 | 2.8682 | 2.9001 | 2.9302 | 2.9619 | 2.9942 | 3.0287 | 3.0652 | 3.1011 | 3.139 | 3.177 | 3.2151 | 3.2535 | 3.2935 | 3.3335 | 3.3735 | 3.4135 | 3.4535 | 3.4959
C''PEAK | 0.51695 | 0.51973 | 0.52247 | 0.52612 | 0.52968 | 0.53324 | 0.53671 | 0.54012 | 0.54345 | 0.54669 | 0.54988 | 0.55289 | 0.5558 | 0.55853 | 0.56117 | 0.5636 | 0.56593 | 0.56819 | 0.57022 | 0.57217 | 0.57397 | 0.5757
10 Hz
100 Hz
0.2 MHz
Notes: D:\PhD\Experiment\2016_06_29_Full Shelf Water Study\281016 (Different height from base)\3 mm from base

## Slide 77: Solidification End Point
Chart: Time/h
Solidification end point
Select 0.2 MHz as the frequency for determining the end of ice solidification owing to the fact that it is not impacted by temperature
The capacitance of ice has almost no temperature dependence at frequency above the relaxation frequency of ice (~1 kHz) such as (0.2 MHz).
Any changes in (0.2 MHz) either with time or temperature, can be associated with the completion of ice formation on freezing
Notes: E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 78: Solidification End Point
Chart: C′(10 Hz)/pF
Chart: Time/h
Ice nucleation
Solidification end point
Select 0.2 MHz as the frequency for determining the end of ice solidification owing to the fact that it is not impacted by temperature
The capacitance of ice has almost no temperature dependence at frequency above the relaxation frequency of ice (~1 kHz) such as (0.2 MHz).
Any changes in (0.2 MHz) either with time or temperature, can be associated with the completion of ice formation on freezing
Notes: E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 79: Solidification End Point
Chart: C′(10 Hz)/pF
Chart: Time/h
Ice solidification time
Ice nucleation
Solidification end point
Select 0.2 MHz as the frequency for determining the end of ice solidification owing to the fact that it is not impacted by temperature
The capacitance of ice has almost no temperature dependence at frequency above the relaxation frequency of ice (~1 kHz) such as (0.2 MHz).
Any changes in (0.2 MHz) either with time or temperature, can be associated with the completion of ice formation on freezing or the glass transition (Next section)
Notes: E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 80: Examples
5%Sucrose
Water
5%Sucrose in 0.26% NaCl
5%Sucrose in 0.55% NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
1.7 h (-14.1 °C)
1.93 h
~14 mins
Chart: C′(10 Hz)/pF
Chart: Time/h
1.97 h (-10.5 °C)
2.30 h
~20 mins
Chart: C′(10 Hz)/pF
Chart: Time/h
1.73 h (-11.7 °C)
2.30 h
Chart: C′(10 Hz)/pF
Chart: Time/h
1.8 h (-11.9 °C)
2.53 h
~44 mins
~34 mins
Solidification time for ice formation increases with the salt concentration
Notes: D:\PhD\Experiment\2016_06_29_Full Shelf Water Study\281016 (Different height from base)\3 mm from base
E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 81
TVIS Applications
Determination of in-vial Glass Transition temperature ()

## Slide 82: Glass Transition Temperature
Chart: Time /h
TVIS record every 120 seconds
(2nd spectrum start after 1st spectrum 10 seconds)
Thermocouple
TC at the middle of solution bounded within electrode region
(defined as Temperature node)
TVIS vial
TC vial
Thermocouple position
Annealing
Ramp from -50 °C to -10 °C with 0.2 /min
Shelf Temperature
Notes: E:\PhD\Experiment\2017_11_01_Glass Transition Study of Sugar\171117-Sucrose

## Slide 83: during Re-heating
5% Sucrose in
0% NaCl
5% Sucrose in
0.26% NaCl
5% Sucrose in
0.55% NaCl
Chart: Heat flow/mW
Chart: Temperature/°C
-37.2 oC
-35.7 oC
1.5 °C
Chart: Heat flow/mW
Chart: Temperature/°C
-41.7 oC
-40.7 oC
1 °C
Chart: Temperature/°C
Chart: Heat flow/mW
-45.2 oC
-45.7 oC
0.5 °C
Onset of glass transition observed by the inflection in the temperature profile of the TVIS parameter C′(0.2MHz)
Notes: E:\PhD\Experiment\2017_11_01_Glass Transition Study of Sugar\171117-Sucrose
File name: Summary Tslice CH1 to CH5_v3, DSC Data_Sucrose-Salt_v3, 2nd Re-heating

## Slide 84
TVIS Applications
Determination of Collapse Phenomena

## Slide 85: Collapse Phenomena 5%w/v Lactose in water
Chart: Temp. /oC
Chart: Pressure / torr
Chart: Time/h
Chart: Dry Layer Thickness (cm)
1.75cm2·Torr·h·g-1
Notes: Version 14
If we are now calculating Rp then we only need to show the product temperature (Tfpeak) and the condenser temperature
Then below that the values for partial pressure – to replace the log Fpeak plot
Then the drying rate (ice vs time) and not c”peak vs time on the bottom left
Replace the top right with my new Rp values
Full shelf 5%w/v Lactose in ultrapure water 23Jan17-CH3 (Fill factor 0.7 = 3.04 g)-3mm above base electrode
From the literature (Pikal (1985) PDA Journal of Pharmaceutical Science and Technology, 39(3), pp.115-139.): Rs ~0.33 Torr·h·g-1.
The temperatures we are using (calculated from log Fpeak) have not been smoothed and the values for C"peak (although taken from the moving average) both have scatter to them - hence why the plot of Rp vs dry layer is quite noisy
Rp starts at 2 (which I presume is the stopper resistance) then it increases as the dry layer develops in thickness up to 3 mm - which is the point at which micro-collapse has happened. At that point, the overall resistance decreases which means that microcollapse has also occurred in the layer above the ice - which already has no ice. If there was no microcollapse in that layer then Rp would stay constant after 3 mm of the dry layer develops. Instead, Rp drops to a new value at ~ 4mm dry layer (when microcollapse occurs) and doesn't change after that. I presume that is because a more dramatic microcollapse occurs in the layers that develop below 4 mm - which are seen on our SEM as large pores and so the layers below the top 3 mm do not have a measureable dry layer resistance.
Milton et al. study product resistance by comparing the direct vial method and MTM (data shown in literature are from direct vial technique according to Pikal's method). They concluded that the product resistance value from MTM agree with direct vial technique and has the same pattern. For lactose, they reported that product resistance increases continuously with thickness up to the onset of collapse (below about -30 Celcius).

## Slide 86: Collapse Phenomena 5%w/v Lactose in water
Middle layer
Micro-collapse
Top layer
Fine pores
Bottom layer
Full collapse
500X
3 mm
500X
500X
Chart: Dry Layer Thickness (cm)
3 mm
Notes: Full shelf 5%w/v Lactose in ultrapure water 23Jan17-CH3 (Fill factor 0.7 = 3.04 g)-3mm above base electrode
Img: 5%lactose_16Nov16-Ng03-10a (ch3)

## Slide 87: Future Work
Development mapping a drying characteristics
  heat transfer coefficients ()
  dry layer resistance ()
Instrument Development
  Commercial C-TVIS (2018)
  Non-contact TVIS (2018-19)
    Micro-well screening
    Vial clusters in batch FD
  TVIS - Shuttle (2019-20)
Non-invasive real time information for characterising the freeze drying

## Slide 88: Further Reading

## Slide 89: New instrumentation development Z-FDM impedance spectroscopy enabled freeze-drying microscope
Grant Funding:
FASTLyo Innovate UK feasibility study . Completed Feb 28 2019
Application :
Development of micro-collapsed formulations with fast drying potential

## Slide 90: Acknowledgements, Recent Projects & Collaborators
De Montfort University, School of Pharmacy
  Evgeny Polygalov: co-inventor of TVIS instrument
  Yowwares Jeeraruangrattana. PhD student
  Bhaskar Pandya. PhD student
  Irina Ermolina. Senior Lecturer
Biopharmaceutical Stability at
Room Temperature
Analytical Technologies for the
Stabilization of Biopharmaceuticals
Government Support for industry
LyoDEA
Lyophilization process analytics
By dielectric analysis
@LyoGroupDMU
Our data
Our Twitter Page
Our WebPage