## Slide 1
Through Vial Impedance Spectroscopy (TVIS)
Dual-electrode System for Process Parameter Determination
Geoff Smith
Leicester School of Pharmacy, De Montfort University, United Kingdom
ISL-FD Midwest Chapter Annual Meeting, April 12th 2018
Midwest Conference Centre, Northlake, IL 60164, USA

## Slide 2: Outline
Description of TVIS measurement system
TVIS dielectric loss mechanisms
First time report on the use of dual-electrode system and its applications
    Ice region specific temperature prediction (, )
    Drying rate determination
    Heat transfer coefficient () determination
Acknowledgements
Notes: Change picture for you at freeze-dryer

## Slide 3
Through Vial Impedance Spectroscopy (TVIS)
Description of Measurement System

## Slide 4: Introduction to the TVIS System
Impedance spectroscopy characterizes the ability of materials to conduct electricity under an applied an oscillating voltage (of varying frequency)
Impedance measurements across a vial rather than within the vial
Hence “Through Vial Impedance Spectroscopy”
Features
  Single vial “non-product invasive”
  Both freezing and drying characterised in a single technique
  Non-perturbing to the packing of vials
  Stopper mechanism unaffected
TVIS
SV product temperature | 
SV sublimation rate | 
SV end point | 

## Slide 5
Freeze drying chamber
Stimulating voltage
Resultant current
LyoDEATM measurement software
Junction
box
TVIS system
(I to V convertor)
Pass-through
TVIS
measurement vial
LyoViewTM analysis software
Notes: Version 14 change : don’t animate first picture – make it appear with the slide
Change photo for the pass-through to the new one
V4 edited slide

## Slide 6
Through Vial Impedance Spectroscopy (TVIS)
Dielectric Loss Mechanisms

## Slide 7
Measurement vial
Glass
Glass
1.1
mm
22
mm
1.1
mm
10 mm
Electronic polarization distortion of electrons relative to the nuclei
+
-
Atomic polarization distortion of nuclei across a heteroatom bond by stretching and bending
-
+
Instantaneous polarization dominant mechanism in the glass wall
Electronic polarization distortion of electrons relative to the nuclei
+
-
Atomic polarization distortion of nuclei across a heteroatom bond by stretching and bending
-
+
Instantaneous polarization dominant mechanism in the glass wall
TVIS response for empty vial
+
+
+
+
+
+
+
+
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
-
-
-
-
-
-
-
-
TVIS response for empty vial

## Slide 8
Measurement vial
Glass
Glass
1.1
mm
22
mm
1.1
mm
10 mm
Electronic polarization distortion of electrons relative to the nuclei
+
-
Atomic polarization distortion of nuclei across a heteroatom bond by stretching and bending
-
+
Instantaneous polarization dominant mechanism in the glass wall
Electronic polarization distortion of electrons relative to the nuclei
+
-
Atomic polarization distortion of nuclei across a heteroatom bond by stretching and bending
-
+
Instantaneous polarization dominant mechanism in the glass wall
MW (space-charge) polarization at glass wall – sample interface
Space charge polarization
(weak frequency dependence)
Space charge polarization
(weak frequency dependence)
+
+
+
+
+
+
+
+
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
-
-
-
-
-
-
-
-
TVIS response for empty vial
OH−
OH−
OH−
OH−
H+
H+
H+
H+
TVIS response for empty vial

## Slide 9
Measurement vial
Glass
Glass
1.1
mm
22
mm
1.1
mm
10 mm
Electronic polarization distortion of electrons relative to the nuclei
+
-
Atomic polarization distortion of nuclei across a heteroatom bond by stretching and bending
-
+
Instantaneous polarization dominant mechanism in the glass wall
Electronic polarization distortion of electrons relative to the nuclei
+
-
Atomic polarization distortion of nuclei across a heteroatom bond by stretching and bending
-
+
Instantaneous polarization dominant mechanism in the glass wall
MW (space-charge) polarization at glass wall – sample interface
MW (space-charge) polarization at glass wall – sample interface
TVIS response for empty vial
Space charge polarization
(weak frequency dependence)
Space charge polarization
(weak frequency dependence)
+
+
+
+
+
+
+
+
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
-
-
-
-
-
-
-
-
TVIS response for empty vial
o
H
H
δ-
δ+
δ+
Dipolar polarization
re-orientation/alignment of permanent dipoles in liquid water (Debye-like relaxation)
Polarization mechanisms in liquid water
(relaxation time,  ~ 9 ps at 20oC)
Proton-hopping
Conduction of protons in liquid water occurs through the Grotthuss "hop-turn" mechanism
Conductivity in pure water
TVIS response for liquid water
OH−
OH−
OH−
OH−
H+
H+
H+
H+
OH−
H+

## Slide 10
Dominant at T > 235 K (approx. −40 oC)
Generation/migration of L- and D- orientation defects
in ice Ih
Dominant at T < 235 K (approx. −40 oC)
Generation/migration of H3O+/OH− ion pairs (ionic defects) in ice Ih (similar to the Grotthus mechanism)
Polarization mechanism in ice
TVIS response for frozen water (ice)
Measurement vial
Glass
Glass
1.1
mm
22
mm
1.1
mm
10 mm
Electronic polarization distortion of electrons relative to the nuclei
+
-
Atomic polarization distortion of nuclei across a heteroatom bond by stretching and bending
-
+
Instantaneous polarization dominant mechanism in the glass wall
Electronic polarization distortion of electrons relative to the nuclei
+
-
Atomic polarization distortion of nuclei across a heteroatom bond by stretching and bending
-
+
Instantaneous polarization dominant mechanism in the glass wall
MW (space-charge) polarization at glass wall – sample interface
MW (space-charge) polarization at glass wall – sample interface
TVIS response for empty vial
Space charge polarization
(weak frequency dependence)
Space charge polarization
(weak frequency dependence)
+
+
+
+
+
+
+
+
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
-
-
-
-
-
-
-
-
TVIS response for empty vial

## Slide 11: Frozen Water and Dielectric Relaxation of Ice
: The polarization of the water dipole in liquid water at 20 ˚C, with a dielectric loss peak frequency of ~ 18 GHz
Chart: Log Frequency
| 6.1 | 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 0.23 h, 20.3 celcius | 1.0162 | 1.0162 | 1.0162 | 1.0162 | 1.0162 | 1.0176 | 1.024 | 1.0293 | 1.0371 | 1.0465 | 1.0575 | 1.0742 | 1.0971 | 1.131 | 1.18 | 1.2538 | 1.3532 | 1.4803 | 1.6281 | 1.7827 | 1.9288 | 2.0567 | 2.1587 | 2.2374 | 2.2971 | 2.3414 | 2.3756 | 2.4038 | 2.4272 | 2.4476 | 2.4662 | 2.4852 | 2.5003 | 2.5152 | 2.5286 | 2.5419 | 2.5531 | 2.5615 | 2.5717 | 2.5802 | 2.589 | 2.5992 | 2.6061 | 2.6201 | 2.625 | 2.6478 | 2.6645 | 2.686 | 2.708 | 2.7359 | 2.7645 | 2.7969
Time 5.03 h, -20 celcius | 0.084287 | 0.084287 | 0.084287 | 0.084287 | 0.084287 | 0.083003 | 0.084127 | 0.084469 | 0.085543 | 0.086743 | 0.087572 | 0.088047 | 0.088435 | 0.088561 | 0.088505 | 0.088289 | 0.08811 | 0.08822 | 0.08935 | 0.092855 | 0.10126 | 0.11875 | 0.15045 | 0.2025 | 0.28088 | 0.38519 | 0.51085 | 0.64628 | 0.77735 | 0.89303 | 0.98847 | 1.0619 | 1.1159 | 1.155 | 1.183 | 1.203 | 1.2167 | 1.2267 | 1.2373 | 1.2444 | 1.2521 | 1.2603 | 1.2702 | 1.2796 | 1.2944 | 1.3089 | 1.3285 | 1.3531 | 1.3812 | 1.4143 | 1.4539 | 1.499
Time 8.5 h, -40 celcius | 0.083695 | 0.083695 | 0.083695 | 0.083695 | 0.083695 | 0.08324 | 0.083622 | 0.083847 | 0.084834 | 0.086128 | 0.086714 | 0.087381 | 0.087886 | 0.088558 | 0.089009 | 0.089429 | 0.089952 | 0.090576 | 0.09152 | 0.092833 | 0.094685 | 0.097293 | 0.10083 | 0.10562 | 0.11226 | 0.12153 | 0.13502 | 0.15512 | 0.18502 | 0.22832 | 0.28819 | 0.36571 | 0.45912 | 0.56375 | 0.67242 | 0.77798 | 0.87404 | 0.95702 | 1.0284 | 1.0845 | 1.1268 | 1.1614 | 1.1873 | 1.2052 | 1.2167 | 1.2287 | 1.2359 | 1.2394 | 1.2468 | 1.2512 | 1.2552 | 1.2597
Chart: Log Frequency
| 6.1 | 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 0.23 h, 20.3 celcius | 0.017404 | 0.021874 | 0.027746 | 0.035469 | 0.045629 | 0.058992 | 0.069416 | 0.087012 | 0.11672 | 0.14702 | 0.18167 | 0.23092 | 0.28348 | 0.3471 | 0.41787 | 0.49639 | 0.57052 | 0.6315 | 0.66609 | 0.66868 | 0.64015 | 0.58783 | 0.52385 | 0.45806 | 0.3962 | 0.3406 | 0.2936 | 0.25428 | 0.22269 | 0.19622 | 0.17597 | 0.15804 | 0.14279 | 0.13021 | 0.12053 | 0.11142 | 0.10593 | 0.099251 | 0.10112 | 0.096409 | 0.099903 | 0.099379 | 0.11049 | 0.12227 | 0.13167 | 0.14829 | 0.16415 | 0.18118 | 0.20571 | 0.22768 | 0.25157 | 0.27443
Time 5.03 h, -20 celcius | 0.0046903 | 0.0046903 | 0.0060276 | 0.0048957 | 0.0052689 | 0.0080706 | 0.0070995 | 0.0085621 | 0.011385 | 0.014441 | 0.017939 | 0.023068 | 0.028754 | 0.036271 | 0.045614 | 0.058069 | 0.073874 | 0.094576 | 0.12111 | 0.15588 | 0.20027 | 0.2566 | 0.32289 | 0.39693 | 0.47149 | 0.53464 | 0.57592 | 0.58749 | 0.56901 | 0.52696 | 0.47094 | 0.40966 | 0.35018 | 0.29633 | 0.24993 | 0.21217 | 0.18272 | 0.16189 | 0.14492 | 0.13366 | 0.12856 | 0.12752 | 0.12988 | 0.13813 | 0.14441 | 0.16577 | 0.18416 | 0.20379 | 0.22754 | 0.25211 | 0.27626 | 0.29768
Time 8.5 h, -40 celcius | 0.0037257 | 0.0037257 | 0.0022099 | 0.0055149 | 0.0027352 | 0.0046061 | 0.0028107 | 0.0031095 | 0.0040131 | 0.0050574 | 0.0061923 | 0.007861 | 0.0096387 | 0.01196 | 0.014763 | 0.018394 | 0.022849 | 0.028488 | 0.035374 | 0.043956 | 0.054453 | 0.067485 | 0.083211 | 0.10248 | 0.12627 | 0.15518 | 0.19059 | 0.23316 | 0.28257 | 0.33712 | 0.39303 | 0.44437 | 0.48479 | 0.5087 | 0.51347 | 0.50011 | 0.47159 | 0.43276 | 0.38638 | 0.3379 | 0.29199 | 0.24717 | 0.21022 | 0.17522 | 0.14921 | 0.12645 | 0.10764 | 0.094062 | 0.085294 | 0.077832 | 0.076166 | 0.074374
20.3 °C
-20 °C
-40 °C
-40 °C
-20 °C
20.3 °C
(I)
18 GHz
Real part Capacitance
Imaginary part Capacitance
Notes: Full Shelf Water Study_281016_Ch4 (Fill factor 0.9) 3mm gap , 400 ubar

## Slide 12: Frozen Water and Dielectric Relaxation of Ice
: The polarization of the water dipole in liquid water at 20 ˚C, with a dielectric loss peak frequency of ~ 18 GHz
: The Maxwell-Wagner (MW) polarization of the glass wall of the TVIS vial at 20 ˚C, with a dielectric loss peak frequency of 17.8 kHz
Chart: Log Frequency
| 6.1 | 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 0.23 h, 20.3 celcius | 1.0162 | 1.0162 | 1.0162 | 1.0162 | 1.0162 | 1.0176 | 1.024 | 1.0293 | 1.0371 | 1.0465 | 1.0575 | 1.0742 | 1.0971 | 1.131 | 1.18 | 1.2538 | 1.3532 | 1.4803 | 1.6281 | 1.7827 | 1.9288 | 2.0567 | 2.1587 | 2.2374 | 2.2971 | 2.3414 | 2.3756 | 2.4038 | 2.4272 | 2.4476 | 2.4662 | 2.4852 | 2.5003 | 2.5152 | 2.5286 | 2.5419 | 2.5531 | 2.5615 | 2.5717 | 2.5802 | 2.589 | 2.5992 | 2.6061 | 2.6201 | 2.625 | 2.6478 | 2.6645 | 2.686 | 2.708 | 2.7359 | 2.7645 | 2.7969
Time 5.03 h, -20 celcius | 0.084287 | 0.084287 | 0.084287 | 0.084287 | 0.084287 | 0.083003 | 0.084127 | 0.084469 | 0.085543 | 0.086743 | 0.087572 | 0.088047 | 0.088435 | 0.088561 | 0.088505 | 0.088289 | 0.08811 | 0.08822 | 0.08935 | 0.092855 | 0.10126 | 0.11875 | 0.15045 | 0.2025 | 0.28088 | 0.38519 | 0.51085 | 0.64628 | 0.77735 | 0.89303 | 0.98847 | 1.0619 | 1.1159 | 1.155 | 1.183 | 1.203 | 1.2167 | 1.2267 | 1.2373 | 1.2444 | 1.2521 | 1.2603 | 1.2702 | 1.2796 | 1.2944 | 1.3089 | 1.3285 | 1.3531 | 1.3812 | 1.4143 | 1.4539 | 1.499
Time 8.5 h, -40 celcius | 0.083695 | 0.083695 | 0.083695 | 0.083695 | 0.083695 | 0.08324 | 0.083622 | 0.083847 | 0.084834 | 0.086128 | 0.086714 | 0.087381 | 0.087886 | 0.088558 | 0.089009 | 0.089429 | 0.089952 | 0.090576 | 0.09152 | 0.092833 | 0.094685 | 0.097293 | 0.10083 | 0.10562 | 0.11226 | 0.12153 | 0.13502 | 0.15512 | 0.18502 | 0.22832 | 0.28819 | 0.36571 | 0.45912 | 0.56375 | 0.67242 | 0.77798 | 0.87404 | 0.95702 | 1.0284 | 1.0845 | 1.1268 | 1.1614 | 1.1873 | 1.2052 | 1.2167 | 1.2287 | 1.2359 | 1.2394 | 1.2468 | 1.2512 | 1.2552 | 1.2597
Chart: Log Frequency
| 6.1 | 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 0.23 h, 20.3 celcius | 0.017404 | 0.021874 | 0.027746 | 0.035469 | 0.045629 | 0.058992 | 0.069416 | 0.087012 | 0.11672 | 0.14702 | 0.18167 | 0.23092 | 0.28348 | 0.3471 | 0.41787 | 0.49639 | 0.57052 | 0.6315 | 0.66609 | 0.66868 | 0.64015 | 0.58783 | 0.52385 | 0.45806 | 0.3962 | 0.3406 | 0.2936 | 0.25428 | 0.22269 | 0.19622 | 0.17597 | 0.15804 | 0.14279 | 0.13021 | 0.12053 | 0.11142 | 0.10593 | 0.099251 | 0.10112 | 0.096409 | 0.099903 | 0.099379 | 0.11049 | 0.12227 | 0.13167 | 0.14829 | 0.16415 | 0.18118 | 0.20571 | 0.22768 | 0.25157 | 0.27443
Time 5.03 h, -20 celcius | 0.0046903 | 0.0046903 | 0.0060276 | 0.0048957 | 0.0052689 | 0.0080706 | 0.0070995 | 0.0085621 | 0.011385 | 0.014441 | 0.017939 | 0.023068 | 0.028754 | 0.036271 | 0.045614 | 0.058069 | 0.073874 | 0.094576 | 0.12111 | 0.15588 | 0.20027 | 0.2566 | 0.32289 | 0.39693 | 0.47149 | 0.53464 | 0.57592 | 0.58749 | 0.56901 | 0.52696 | 0.47094 | 0.40966 | 0.35018 | 0.29633 | 0.24993 | 0.21217 | 0.18272 | 0.16189 | 0.14492 | 0.13366 | 0.12856 | 0.12752 | 0.12988 | 0.13813 | 0.14441 | 0.16577 | 0.18416 | 0.20379 | 0.22754 | 0.25211 | 0.27626 | 0.29768
Time 8.5 h, -40 celcius | 0.0037257 | 0.0037257 | 0.0022099 | 0.0055149 | 0.0027352 | 0.0046061 | 0.0028107 | 0.0031095 | 0.0040131 | 0.0050574 | 0.0061923 | 0.007861 | 0.0096387 | 0.01196 | 0.014763 | 0.018394 | 0.022849 | 0.028488 | 0.035374 | 0.043956 | 0.054453 | 0.067485 | 0.083211 | 0.10248 | 0.12627 | 0.15518 | 0.19059 | 0.23316 | 0.28257 | 0.33712 | 0.39303 | 0.44437 | 0.48479 | 0.5087 | 0.51347 | 0.50011 | 0.47159 | 0.43276 | 0.38638 | 0.3379 | 0.29199 | 0.24717 | 0.21022 | 0.17522 | 0.14921 | 0.12645 | 0.10764 | 0.094062 | 0.085294 | 0.077832 | 0.076166 | 0.074374
20.3 °C
-20 °C
-40 °C
-40 °C
-20 °C
20.3 °C
(II)
(I)
18 GHz
18 kHz
Real part Capacitance
Imaginary part Capacitance
Notes: Full Shelf Water Study_281016_Ch4 (Fill factor 0.9) 3mm gap , 400 ubar

## Slide 13: Frozen Water and Dielectric Relaxation of Ice
: The polarization of the water dipole in liquid water at 20 ˚C, with a dielectric loss peak frequency of ~ 18 GHz
: The Maxwell-Wagner (MW) polarization of the glass wall of the TVIS vial at 20 ˚C, with a dielectric loss peak frequency of 17.8 kHz
: The dielectric polarization of ice at −20 ˚C, with a dielectric loss peak frequencies of 2.57 kHz
Chart: Log Frequency
| 6.1 | 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 0.23 h, 20.3 celcius | 1.0162 | 1.0162 | 1.0162 | 1.0162 | 1.0162 | 1.0176 | 1.024 | 1.0293 | 1.0371 | 1.0465 | 1.0575 | 1.0742 | 1.0971 | 1.131 | 1.18 | 1.2538 | 1.3532 | 1.4803 | 1.6281 | 1.7827 | 1.9288 | 2.0567 | 2.1587 | 2.2374 | 2.2971 | 2.3414 | 2.3756 | 2.4038 | 2.4272 | 2.4476 | 2.4662 | 2.4852 | 2.5003 | 2.5152 | 2.5286 | 2.5419 | 2.5531 | 2.5615 | 2.5717 | 2.5802 | 2.589 | 2.5992 | 2.6061 | 2.6201 | 2.625 | 2.6478 | 2.6645 | 2.686 | 2.708 | 2.7359 | 2.7645 | 2.7969
Time 5.03 h, -20 celcius | 0.084287 | 0.084287 | 0.084287 | 0.084287 | 0.084287 | 0.083003 | 0.084127 | 0.084469 | 0.085543 | 0.086743 | 0.087572 | 0.088047 | 0.088435 | 0.088561 | 0.088505 | 0.088289 | 0.08811 | 0.08822 | 0.08935 | 0.092855 | 0.10126 | 0.11875 | 0.15045 | 0.2025 | 0.28088 | 0.38519 | 0.51085 | 0.64628 | 0.77735 | 0.89303 | 0.98847 | 1.0619 | 1.1159 | 1.155 | 1.183 | 1.203 | 1.2167 | 1.2267 | 1.2373 | 1.2444 | 1.2521 | 1.2603 | 1.2702 | 1.2796 | 1.2944 | 1.3089 | 1.3285 | 1.3531 | 1.3812 | 1.4143 | 1.4539 | 1.499
Time 8.5 h, -40 celcius | 0.083695 | 0.083695 | 0.083695 | 0.083695 | 0.083695 | 0.08324 | 0.083622 | 0.083847 | 0.084834 | 0.086128 | 0.086714 | 0.087381 | 0.087886 | 0.088558 | 0.089009 | 0.089429 | 0.089952 | 0.090576 | 0.09152 | 0.092833 | 0.094685 | 0.097293 | 0.10083 | 0.10562 | 0.11226 | 0.12153 | 0.13502 | 0.15512 | 0.18502 | 0.22832 | 0.28819 | 0.36571 | 0.45912 | 0.56375 | 0.67242 | 0.77798 | 0.87404 | 0.95702 | 1.0284 | 1.0845 | 1.1268 | 1.1614 | 1.1873 | 1.2052 | 1.2167 | 1.2287 | 1.2359 | 1.2394 | 1.2468 | 1.2512 | 1.2552 | 1.2597
Chart: Log Frequency
| 6.1 | 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 0.23 h, 20.3 celcius | 0.017404 | 0.021874 | 0.027746 | 0.035469 | 0.045629 | 0.058992 | 0.069416 | 0.087012 | 0.11672 | 0.14702 | 0.18167 | 0.23092 | 0.28348 | 0.3471 | 0.41787 | 0.49639 | 0.57052 | 0.6315 | 0.66609 | 0.66868 | 0.64015 | 0.58783 | 0.52385 | 0.45806 | 0.3962 | 0.3406 | 0.2936 | 0.25428 | 0.22269 | 0.19622 | 0.17597 | 0.15804 | 0.14279 | 0.13021 | 0.12053 | 0.11142 | 0.10593 | 0.099251 | 0.10112 | 0.096409 | 0.099903 | 0.099379 | 0.11049 | 0.12227 | 0.13167 | 0.14829 | 0.16415 | 0.18118 | 0.20571 | 0.22768 | 0.25157 | 0.27443
Time 5.03 h, -20 celcius | 0.0046903 | 0.0046903 | 0.0060276 | 0.0048957 | 0.0052689 | 0.0080706 | 0.0070995 | 0.0085621 | 0.011385 | 0.014441 | 0.017939 | 0.023068 | 0.028754 | 0.036271 | 0.045614 | 0.058069 | 0.073874 | 0.094576 | 0.12111 | 0.15588 | 0.20027 | 0.2566 | 0.32289 | 0.39693 | 0.47149 | 0.53464 | 0.57592 | 0.58749 | 0.56901 | 0.52696 | 0.47094 | 0.40966 | 0.35018 | 0.29633 | 0.24993 | 0.21217 | 0.18272 | 0.16189 | 0.14492 | 0.13366 | 0.12856 | 0.12752 | 0.12988 | 0.13813 | 0.14441 | 0.16577 | 0.18416 | 0.20379 | 0.22754 | 0.25211 | 0.27626 | 0.29768
Time 8.5 h, -40 celcius | 0.0037257 | 0.0037257 | 0.0022099 | 0.0055149 | 0.0027352 | 0.0046061 | 0.0028107 | 0.0031095 | 0.0040131 | 0.0050574 | 0.0061923 | 0.007861 | 0.0096387 | 0.01196 | 0.014763 | 0.018394 | 0.022849 | 0.028488 | 0.035374 | 0.043956 | 0.054453 | 0.067485 | 0.083211 | 0.10248 | 0.12627 | 0.15518 | 0.19059 | 0.23316 | 0.28257 | 0.33712 | 0.39303 | 0.44437 | 0.48479 | 0.5087 | 0.51347 | 0.50011 | 0.47159 | 0.43276 | 0.38638 | 0.3379 | 0.29199 | 0.24717 | 0.21022 | 0.17522 | 0.14921 | 0.12645 | 0.10764 | 0.094062 | 0.085294 | 0.077832 | 0.076166 | 0.074374
20.3 °C
-20 °C
-40 °C
-40 °C
-20 °C
20.3 °C
(II)
(I)
(III)
18 GHz
18 kHz
2.6 kHz
Real part Capacitance
Imaginary part Capacitance
Notes: Full Shelf Water Study_281016_Ch4 (Fill factor 0.9) 3mm gap , 400 ubar

## Slide 14: Frozen Water and Dielectric Relaxation of Ice
: The polarization of the water dipole in liquid water at 20 ˚C, with a dielectric loss peak frequency of ~ 18 GHz
: The Maxwell-Wagner (MW) polarization of the glass wall of the TVIS vial at 20 ˚C, with a dielectric loss peak frequency of 17.8 kHz
: The dielectric polarization of ice at −20 ˚C, with a dielectric loss peak frequencies of 2.57 kHz
: The dielectric polarization of ice at −40 ˚C with a dielectric loss peak frequencies of 537 Hz.
Chart: Log Frequency
| 6.1 | 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 0.23 h, 20.3 celcius | 1.0162 | 1.0162 | 1.0162 | 1.0162 | 1.0162 | 1.0176 | 1.024 | 1.0293 | 1.0371 | 1.0465 | 1.0575 | 1.0742 | 1.0971 | 1.131 | 1.18 | 1.2538 | 1.3532 | 1.4803 | 1.6281 | 1.7827 | 1.9288 | 2.0567 | 2.1587 | 2.2374 | 2.2971 | 2.3414 | 2.3756 | 2.4038 | 2.4272 | 2.4476 | 2.4662 | 2.4852 | 2.5003 | 2.5152 | 2.5286 | 2.5419 | 2.5531 | 2.5615 | 2.5717 | 2.5802 | 2.589 | 2.5992 | 2.6061 | 2.6201 | 2.625 | 2.6478 | 2.6645 | 2.686 | 2.708 | 2.7359 | 2.7645 | 2.7969
Time 5.03 h, -20 celcius | 0.084287 | 0.084287 | 0.084287 | 0.084287 | 0.084287 | 0.083003 | 0.084127 | 0.084469 | 0.085543 | 0.086743 | 0.087572 | 0.088047 | 0.088435 | 0.088561 | 0.088505 | 0.088289 | 0.08811 | 0.08822 | 0.08935 | 0.092855 | 0.10126 | 0.11875 | 0.15045 | 0.2025 | 0.28088 | 0.38519 | 0.51085 | 0.64628 | 0.77735 | 0.89303 | 0.98847 | 1.0619 | 1.1159 | 1.155 | 1.183 | 1.203 | 1.2167 | 1.2267 | 1.2373 | 1.2444 | 1.2521 | 1.2603 | 1.2702 | 1.2796 | 1.2944 | 1.3089 | 1.3285 | 1.3531 | 1.3812 | 1.4143 | 1.4539 | 1.499
Time 8.5 h, -40 celcius | 0.083695 | 0.083695 | 0.083695 | 0.083695 | 0.083695 | 0.08324 | 0.083622 | 0.083847 | 0.084834 | 0.086128 | 0.086714 | 0.087381 | 0.087886 | 0.088558 | 0.089009 | 0.089429 | 0.089952 | 0.090576 | 0.09152 | 0.092833 | 0.094685 | 0.097293 | 0.10083 | 0.10562 | 0.11226 | 0.12153 | 0.13502 | 0.15512 | 0.18502 | 0.22832 | 0.28819 | 0.36571 | 0.45912 | 0.56375 | 0.67242 | 0.77798 | 0.87404 | 0.95702 | 1.0284 | 1.0845 | 1.1268 | 1.1614 | 1.1873 | 1.2052 | 1.2167 | 1.2287 | 1.2359 | 1.2394 | 1.2468 | 1.2512 | 1.2552 | 1.2597
Chart: Log Frequency
| 6.1 | 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 0.23 h, 20.3 celcius | 0.017404 | 0.021874 | 0.027746 | 0.035469 | 0.045629 | 0.058992 | 0.069416 | 0.087012 | 0.11672 | 0.14702 | 0.18167 | 0.23092 | 0.28348 | 0.3471 | 0.41787 | 0.49639 | 0.57052 | 0.6315 | 0.66609 | 0.66868 | 0.64015 | 0.58783 | 0.52385 | 0.45806 | 0.3962 | 0.3406 | 0.2936 | 0.25428 | 0.22269 | 0.19622 | 0.17597 | 0.15804 | 0.14279 | 0.13021 | 0.12053 | 0.11142 | 0.10593 | 0.099251 | 0.10112 | 0.096409 | 0.099903 | 0.099379 | 0.11049 | 0.12227 | 0.13167 | 0.14829 | 0.16415 | 0.18118 | 0.20571 | 0.22768 | 0.25157 | 0.27443
Time 5.03 h, -20 celcius | 0.0046903 | 0.0046903 | 0.0060276 | 0.0048957 | 0.0052689 | 0.0080706 | 0.0070995 | 0.0085621 | 0.011385 | 0.014441 | 0.017939 | 0.023068 | 0.028754 | 0.036271 | 0.045614 | 0.058069 | 0.073874 | 0.094576 | 0.12111 | 0.15588 | 0.20027 | 0.2566 | 0.32289 | 0.39693 | 0.47149 | 0.53464 | 0.57592 | 0.58749 | 0.56901 | 0.52696 | 0.47094 | 0.40966 | 0.35018 | 0.29633 | 0.24993 | 0.21217 | 0.18272 | 0.16189 | 0.14492 | 0.13366 | 0.12856 | 0.12752 | 0.12988 | 0.13813 | 0.14441 | 0.16577 | 0.18416 | 0.20379 | 0.22754 | 0.25211 | 0.27626 | 0.29768
Time 8.5 h, -40 celcius | 0.0037257 | 0.0037257 | 0.0022099 | 0.0055149 | 0.0027352 | 0.0046061 | 0.0028107 | 0.0031095 | 0.0040131 | 0.0050574 | 0.0061923 | 0.007861 | 0.0096387 | 0.01196 | 0.014763 | 0.018394 | 0.022849 | 0.028488 | 0.035374 | 0.043956 | 0.054453 | 0.067485 | 0.083211 | 0.10248 | 0.12627 | 0.15518 | 0.19059 | 0.23316 | 0.28257 | 0.33712 | 0.39303 | 0.44437 | 0.48479 | 0.5087 | 0.51347 | 0.50011 | 0.47159 | 0.43276 | 0.38638 | 0.3379 | 0.29199 | 0.24717 | 0.21022 | 0.17522 | 0.14921 | 0.12645 | 0.10764 | 0.094062 | 0.085294 | 0.077832 | 0.076166 | 0.074374
20.3 °C
-20 °C
-40 °C
-40 °C
-20 °C
20.3 °C
(II)
(I)
(III)
(IV)
18 GHz
18 kHz
2.6 kHz
537 Hz
Real part Capacitance
Imaginary part Capacitance
Notes: Full Shelf Water Study_281016_Ch4 (Fill factor 0.9) 3mm gap , 400 ubar

## Slide 15
Through Vial Impedance Spectroscopy (TVIS)
Applications

## Slide 16: Through Vial Impedance Spectroscopy (TVIS)
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034
Time 0.3 h | 0.27396 | 0.24912 | 0.22663 | 0.2063 | 0.18797 | 0.17147 | 0.15668 | 0.14349 | 0.13183 | 0.12162 | 0.11284 | 0.10548 | 0.099574 | 0.095186 | 0.092431 | 0.091474 | 0.092551 | 0.095973 | 0.10215 | 0.11162 | 0.12501 | 0.14319 | 0.16715 | 0.198 | 0.23699 | 0.28544 | 0.34395 | 0.41214 | 0.48865 | 0.56683 | 0.63782 | 0.69048 | 0.71191 | 0.69725 | 0.65001 | 0.57934 | 0.49898 | 0.41726 | 0.34526 | 0.28116 | 0.22886 | 0.17991 | 0.14643 | 0.1181 | 0.089415 | 0.072054
Time 4 h | 0.099081 | 0.096153 | 0.095569 | 0.097591 | 0.10258 | 0.11101 | 0.12349 | 0.14075 | 0.16366 | 0.19315 | 0.23007 | 0.27492 | 0.32735 | 0.38533 | 0.44432 | 0.49685 | 0.53348 | 0.54588 | 0.53067 | 0.49125 | 0.43601 | 0.37424 | 0.31336 | 0.25794 | 0.20992 | 0.16948 | 0.13623 | 0.10926 | 0.087318 | 0.069838 | 0.055881 | 0.044566 | 0.035658 | 0.0285 | 0.02283 | 0.018246 | 0.014616 | 0.011669 | 0.0094101 | 0.0075633 | 0.0061328 | 0.0048349
Time 0.3 h | 0.26705 | 0.24478 | 0.21978 | 0.1979 | 0.17863 | 0.15615 | 0.14276 | 0.12639 | 0.11756 | 0.11024 | 0.10127 | 0.1004 | 0.098391 | 0.10089 | 0.10904 | 0.11009 | 0.11451 | 0.12496 | 0.13439 | 0.14683 | 0.16234 | 0.18119 | 0.20238 | 0.22944 | 0.26256 | 0.30331 | 0.35113 | 0.4089 | 0.47307 | 0.54005 | 0.60366 | 0.65431 | 0.67974 | 0.67269 | 0.63391 | 0.56955 | 0.49367 | 0.41413 | 0.34323 | 0.27998 | 0.22777
Time 4 h | 0.093429 | 0.092843 | 0.096342 | 0.10075 | 0.11232 | 0.12322 | 0.1435 | 0.1596 | 0.19684 | 0.23027 | 0.27214 | 0.31647 | 0.36246 | 0.41025 | 0.45421 | 0.48896 | 0.51147 | 0.51695 | 0.50285 | 0.46926 | 0.42098 | 0.36431 | 0.30625 | 0.25218 | 0.20511 | 0.16565 | 0.13355 | 0.10767 | 0.086605 | 0.069689 | 0.056026 | 0.045028 | 0.036377 | 0.029122 | 0.023205 | 0.018299 | 0.014688 | 0.011778 | 0.0095536 | 0.0077519 | 0.0063112
FPEAK
FPEAK
Liquid State
Solid State
Monitoring Phase Behaviour (ice nucleation temperature and solidification end points by using
Notes: Eg temperature cycling below to calibrate Fpeak for temperature (in order to predict product temperatures in primary drying) and the compensate cpeak for temperature so that it can used to monitor ice content and hence drying rate
Eg 2 left graph: monitoring phase behaviour (ice nucleation temperatures and solidufication end points)

## Slide 17: Through Vial Impedance Spectroscopy (TVIS)
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034
Time 0.3 h | 0.27396 | 0.24912 | 0.22663 | 0.2063 | 0.18797 | 0.17147 | 0.15668 | 0.14349 | 0.13183 | 0.12162 | 0.11284 | 0.10548 | 0.099574 | 0.095186 | 0.092431 | 0.091474 | 0.092551 | 0.095973 | 0.10215 | 0.11162 | 0.12501 | 0.14319 | 0.16715 | 0.198 | 0.23699 | 0.28544 | 0.34395 | 0.41214 | 0.48865 | 0.56683 | 0.63782 | 0.69048 | 0.71191 | 0.69725 | 0.65001 | 0.57934 | 0.49898 | 0.41726 | 0.34526 | 0.28116 | 0.22886 | 0.17991 | 0.14643 | 0.1181 | 0.089415 | 0.072054
Time 4 h | 0.099081 | 0.096153 | 0.095569 | 0.097591 | 0.10258 | 0.11101 | 0.12349 | 0.14075 | 0.16366 | 0.19315 | 0.23007 | 0.27492 | 0.32735 | 0.38533 | 0.44432 | 0.49685 | 0.53348 | 0.54588 | 0.53067 | 0.49125 | 0.43601 | 0.37424 | 0.31336 | 0.25794 | 0.20992 | 0.16948 | 0.13623 | 0.10926 | 0.087318 | 0.069838 | 0.055881 | 0.044566 | 0.035658 | 0.0285 | 0.02283 | 0.018246 | 0.014616 | 0.011669 | 0.0094101 | 0.0075633 | 0.0061328 | 0.0048349
Time 0.3 h | 0.26705 | 0.24478 | 0.21978 | 0.1979 | 0.17863 | 0.15615 | 0.14276 | 0.12639 | 0.11756 | 0.11024 | 0.10127 | 0.1004 | 0.098391 | 0.10089 | 0.10904 | 0.11009 | 0.11451 | 0.12496 | 0.13439 | 0.14683 | 0.16234 | 0.18119 | 0.20238 | 0.22944 | 0.26256 | 0.30331 | 0.35113 | 0.4089 | 0.47307 | 0.54005 | 0.60366 | 0.65431 | 0.67974 | 0.67269 | 0.63391 | 0.56955 | 0.49367 | 0.41413 | 0.34323 | 0.27998 | 0.22777
Time 4 h | 0.093429 | 0.092843 | 0.096342 | 0.10075 | 0.11232 | 0.12322 | 0.1435 | 0.1596 | 0.19684 | 0.23027 | 0.27214 | 0.31647 | 0.36246 | 0.41025 | 0.45421 | 0.48896 | 0.51147 | 0.51695 | 0.50285 | 0.46926 | 0.42098 | 0.36431 | 0.30625 | 0.25218 | 0.20511 | 0.16565 | 0.13355 | 0.10767 | 0.086605 | 0.069689 | 0.056026 | 0.045028 | 0.036377 | 0.029122 | 0.023205 | 0.018299 | 0.014688 | 0.011778 | 0.0095536 | 0.0077519 | 0.0063112
FPEAK
FPEAK
Liquid State
Solid State
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 4.4 h | 0.003361 | 0.0040465 | 0.0025946 | 0.0026214 | 0.0044553 | 0.0026323 | 0.0029087 | 0.0037366 | 0.0047149 | 0.0057564 | 0.0073092 | 0.0089323 | 0.011077 | 0.013703 | 0.017072 | 0.021236 | 0.026498 | 0.03301 | 0.041177 | 0.051293 | 0.064055 | 0.079683 | 0.099225 | 0.12384 | 0.15437 | 0.19231 | 0.23844 | 0.29224 | 0.35137 | 0.41125 | 0.46468 | 0.50437 | 0.52462 | 0.52356 | 0.50251 | 0.46708 | 0.42347 | 0.37253 | 0.32299 | 0.27537 | 0.23266 | 0.1978 | 0.16633 | 0.14424 | 0.12252 | 0.10889 | 0.097697 | 0.093226 | 0.09166 | 0.092155 | 0.093069
Time 4.5 h | 0.0034249 | 0.0041377 | 0.0027283 | 0.0027814 | 0.0046887 | 0.0029259 | 0.0032733 | 0.0042363 | 0.0053358 | 0.0065296 | 0.0082746 | 0.010141 | 0.012601 | 0.015587 | 0.01947 | 0.024233 | 0.030264 | 0.037737 | 0.047171 | 0.058901 | 0.073903 | 0.092482 | 0.11581 | 0.14538 | 0.18208 | 0.2274 | 0.28124 | 0.3416 | 0.40418 | 0.46243 | 0.50811 | 0.53414 | 0.53731 | 0.51809 | 0.48197 | 0.43577 | 0.38293 | 0.3311 | 0.28177 | 0.23748 | 0.20077 | 0.16802 | 0.14416 | 0.11848 | 0.11171 | 0.10187 | 0.09506 | 0.092791 | 0.094429 | 0.097904 | 0.10371
Time 4.6 h | 0.0036091 | 0.0043103 | 0.0029276 | 0.0030069 | 0.0049863 | 0.0032919 | 0.0037411 | 0.0048548 | 0.006126 | 0.0075084 | 0.0095269 | 0.011725 | 0.014579 | 0.018062 | 0.02259 | 0.028182 | 0.035264 | 0.044078 | 0.055262 | 0.069282 | 0.08729 | 0.10998 | 0.13862 | 0.17503 | 0.21981 | 0.27381 | 0.33549 | 0.40061 | 0.46249 | 0.51282 | 0.5432 | 0.54939 | 0.53148 | 0.49439 | 0.44517 | 0.39079 | 0.33577 | 0.28667 | 0.23998 | 0.20372 | 0.17181 | 0.14592 | 0.12721 | 0.11818 | 0.10489 | 0.098824 | 0.097908 | 0.099283 | 0.10394 | 0.11202 | 0.12205
Time 4.7 h | 0.0040254 | 0.0023197 | 0.0060162 | 0.0032852 | 0.0053707 | 0.003773 | 0.0043423 | 0.005672 | 0.0071637 | 0.0088006 | 0.011181 | 0.013771 | 0.017191 | 0.021329 | 0.026752 | 0.033478 | 0.04203 | 0.052722 | 0.066401 | 0.083738 | 0.10613 | 0.13436 | 0.17037 | 0.21573 | 0.27035 | 0.33347 | 0.40107 | 0.46603 | 0.51937 | 0.55253 | 0.55996 | 0.54171 | 0.50311 | 0.45097 | 0.39455 | 0.33876 | 0.28887 | 0.24104 | 0.20387 | 0.17339 | 0.14886 | 0.13225 | 0.11753 | 0.11 | 0.10482 | 0.10438 | 0.1074 | 0.1138 | 0.12408 | 0.13639 | 0.14993
Time 4.8 h | 0.0039527 | 0.0047899 | 0.0034912 | 0.0036734 | 0.0059298 | 0.004463 | 0.0052209 | 0.0068624 | 0.0086692 | 0.010687 | 0.013626 | 0.016821 | 0.021041 | 0.026213 | 0.03302 | 0.041432 | 0.052261 | 0.065911 | 0.083539 | 0.10606 | 0.13545 | 0.17252 | 0.21928 | 0.27665 | 0.34237 | 0.41271 | 0.47977 | 0.53353 | 0.56486 | 0.56874 | 0.5461 | 0.50317 | 0.44813 | 0.38872 | 0.33224 | 0.28045 | 0.23905 | 0.20172 | 0.17254 | 0.14998 | 0.13098 | 0.12232 | 0.11386 | 0.11155 | 0.11412 | 0.11812 | 0.12658 | 0.13848 | 0.15361 | 0.17099 | 0.18866
Time 4.9 h | 0.0043188 | 0.0052277 | 0.0039868 | 0.0042288 | 0.0066529 | 0.0053479 | 0.0063485 | 0.0083796 | 0.010604 | 0.013107 | 0.016765 | 0.020805 | 0.02609 | 0.032614 | 0.041224 | 0.052023 | 0.065957 | 0.083695 | 0.10683 | 0.13651 | 0.17523 | 0.22353 | 0.28261 | 0.35128 | 0.42376 | 0.49222 | 0.54583 | 0.5749 | 0.57491 | 0.54814 | 0.50162 | 0.4443 | 0.38384 | 0.32634 | 0.27564 | 0.23272 | 0.19908 | 0.1717 | 0.1495 | 0.13558 | 0.12502 | 0.11938 | 0.11838 | 0.12437 | 0.13062 | 0.13992 | 0.15492 | 0.1723 | 0.19073 | 0.21306 | 0.23377
Time 5 h | 0.0049133 | 0.0058042 | 0.0046338 | 0.0049712 | 0.0076664 | 0.0066031 | 0.0079479 | 0.010533 | 0.013358 | 0.01657 | 0.021298 | 0.026516 | 0.033397 | 0.041936 | 0.053268 | 0.067599 | 0.086277 | 0.11024 | 0.14161 | 0.18183 | 0.2334 | 0.29549 | 0.36662 | 0.44167 | 0.50969 | 0.5605 | 0.58424 | 0.57776 | 0.54458 | 0.49349 | 0.43374 | 0.3731 | 0.31662 | 0.26729 | 0.22557 | 0.19336 | 0.16657 | 0.14984 | 0.13717 | 0.1291 | 0.12498 | 0.12759 | 0.13258 | 0.1444 | 0.15601 | 0.17194 | 0.19151 | 0.21398 | 0.23437 | 0.26066 | 0.28203
Time 5.1 h | 0.0051615 | 0.0065394 | 0.0054847 | 0.0059393 | 0.0090001 | 0.0082333 | 0.010001 | 0.013324 | 0.016914 | 0.021051 | 0.027147 | 0.03396 | 0.042946 | 0.054219 | 0.069319 | 0.088593 | 0.11394 | 0.14657 | 0.18909 | 0.24251 | 0.30812 | 0.38167 | 0.45767 | 0.52576 | 0.57299 | 0.59132 | 0.57824 | 0.5397 | 0.48481 | 0.42344 | 0.36292 | 0.3071 | 0.25963 | 0.22024 | 0.18846 | 0.16573 | 0.14947 | 0.13846 | 0.13032 | 0.13113 | 0.1334 | 0.1422 | 0.15348 | 0.16641 | 0.18887 | 0.2096 | 0.23238 | 0.25699 | 0.28353 | 0.3063 | 0.32649
-38 oC
-18 oC
Monitoring Phase Behaviour (ice nucleation temperature and solidification end points by using
temperature calibration for predicting temperature of the product in primary drying
Notes: Eg temperature cycling below to calibrate Fpeak for temperature (in order to predict product temperatures in primary drying) and the compensate cpeak for temperature so that it can used to monitor ice content and hence drying rate
Eg 2 left graph: monitoring phase behaviour (ice nucleation temperatures and solidufication end points)

## Slide 18: Through Vial Impedance Spectroscopy (TVIS)
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034
Time 0.3 h | 0.27396 | 0.24912 | 0.22663 | 0.2063 | 0.18797 | 0.17147 | 0.15668 | 0.14349 | 0.13183 | 0.12162 | 0.11284 | 0.10548 | 0.099574 | 0.095186 | 0.092431 | 0.091474 | 0.092551 | 0.095973 | 0.10215 | 0.11162 | 0.12501 | 0.14319 | 0.16715 | 0.198 | 0.23699 | 0.28544 | 0.34395 | 0.41214 | 0.48865 | 0.56683 | 0.63782 | 0.69048 | 0.71191 | 0.69725 | 0.65001 | 0.57934 | 0.49898 | 0.41726 | 0.34526 | 0.28116 | 0.22886 | 0.17991 | 0.14643 | 0.1181 | 0.089415 | 0.072054
Time 4 h | 0.099081 | 0.096153 | 0.095569 | 0.097591 | 0.10258 | 0.11101 | 0.12349 | 0.14075 | 0.16366 | 0.19315 | 0.23007 | 0.27492 | 0.32735 | 0.38533 | 0.44432 | 0.49685 | 0.53348 | 0.54588 | 0.53067 | 0.49125 | 0.43601 | 0.37424 | 0.31336 | 0.25794 | 0.20992 | 0.16948 | 0.13623 | 0.10926 | 0.087318 | 0.069838 | 0.055881 | 0.044566 | 0.035658 | 0.0285 | 0.02283 | 0.018246 | 0.014616 | 0.011669 | 0.0094101 | 0.0075633 | 0.0061328 | 0.0048349
Time 0.3 h | 0.26705 | 0.24478 | 0.21978 | 0.1979 | 0.17863 | 0.15615 | 0.14276 | 0.12639 | 0.11756 | 0.11024 | 0.10127 | 0.1004 | 0.098391 | 0.10089 | 0.10904 | 0.11009 | 0.11451 | 0.12496 | 0.13439 | 0.14683 | 0.16234 | 0.18119 | 0.20238 | 0.22944 | 0.26256 | 0.30331 | 0.35113 | 0.4089 | 0.47307 | 0.54005 | 0.60366 | 0.65431 | 0.67974 | 0.67269 | 0.63391 | 0.56955 | 0.49367 | 0.41413 | 0.34323 | 0.27998 | 0.22777
Time 4 h | 0.093429 | 0.092843 | 0.096342 | 0.10075 | 0.11232 | 0.12322 | 0.1435 | 0.1596 | 0.19684 | 0.23027 | 0.27214 | 0.31647 | 0.36246 | 0.41025 | 0.45421 | 0.48896 | 0.51147 | 0.51695 | 0.50285 | 0.46926 | 0.42098 | 0.36431 | 0.30625 | 0.25218 | 0.20511 | 0.16565 | 0.13355 | 0.10767 | 0.086605 | 0.069689 | 0.056026 | 0.045028 | 0.036377 | 0.029122 | 0.023205 | 0.018299 | 0.014688 | 0.011778 | 0.0095536 | 0.0077519 | 0.0063112
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034
Time 0.3 h | 0.27396 | 0.24912 | 0.22663 | 0.2063 | 0.18797 | 0.17147 | 0.15668 | 0.14349 | 0.13183 | 0.12162 | 0.11284 | 0.10548 | 0.099574 | 0.095186 | 0.092431 | 0.091474 | 0.092551 | 0.095973 | 0.10215 | 0.11162 | 0.12501 | 0.14319 | 0.16715 | 0.198 | 0.23699 | 0.28544 | 0.34395 | 0.41214 | 0.48865 | 0.56683 | 0.63782 | 0.69048 | 0.71191 | 0.69725 | 0.65001 | 0.57934 | 0.49898 | 0.41726 | 0.34526 | 0.28116 | 0.22886 | 0.17991 | 0.14643 | 0.1181 | 0.089415 | 0.072054
Time 4 h | 0.099081 | 0.096153 | 0.095569 | 0.097591 | 0.10258 | 0.11101 | 0.12349 | 0.14075 | 0.16366 | 0.19315 | 0.23007 | 0.27492 | 0.32735 | 0.38533 | 0.44432 | 0.49685 | 0.53348 | 0.54588 | 0.53067 | 0.49125 | 0.43601 | 0.37424 | 0.31336 | 0.25794 | 0.20992 | 0.16948 | 0.13623 | 0.10926 | 0.087318 | 0.069838 | 0.055881 | 0.044566 | 0.035658 | 0.0285 | 0.02283 | 0.018246 | 0.014616 | 0.011669 | 0.0094101 | 0.0075633 | 0.0061328 | 0.0048349
Time 0.3 h | 0.26705 | 0.24478 | 0.21978 | 0.1979 | 0.17863 | 0.15615 | 0.14276 | 0.12639 | 0.11756 | 0.11024 | 0.10127 | 0.1004 | 0.098391 | 0.10089 | 0.10904 | 0.11009 | 0.11451 | 0.12496 | 0.13439 | 0.14683 | 0.16234 | 0.18119 | 0.20238 | 0.22944 | 0.26256 | 0.30331 | 0.35113 | 0.4089 | 0.47307 | 0.54005 | 0.60366 | 0.65431 | 0.67974 | 0.67269 | 0.63391 | 0.56955 | 0.49367 | 0.41413 | 0.34323 | 0.27998 | 0.22777
Time 4 h | 0.093429 | 0.092843 | 0.096342 | 0.10075 | 0.11232 | 0.12322 | 0.1435 | 0.1596 | 0.19684 | 0.23027 | 0.27214 | 0.31647 | 0.36246 | 0.41025 | 0.45421 | 0.48896 | 0.51147 | 0.51695 | 0.50285 | 0.46926 | 0.42098 | 0.36431 | 0.30625 | 0.25218 | 0.20511 | 0.16565 | 0.13355 | 0.10767 | 0.086605 | 0.069689 | 0.056026 | 0.045028 | 0.036377 | 0.029122 | 0.023205 | 0.018299 | 0.014688 | 0.011778 | 0.0095536 | 0.0077519 | 0.0063112
FPEAK
FPEAK
Liquid State
Solid State
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625 | 5.7569630370485534 | 5.8239109124113 | 5.9030899869919438 | 6
Time 13.1 h | 0.10929 | 0.1004 | 0.092502 | 0.087998 | 0.086049 | 0.088115 | 0.0938 | 0.10112 | 0.11429 | 0.13106 | 0.15364 | 0.18166 | 0.21583 | 0.25512 | 0.29985 | 0.34774 | 0.3951 | 0.43795 | 0.46981 | 0.48496 | 0.48002 | 0.45469 | 0.41262 | 0.36037 | 0.30465 | 0.25101 | 0.20331 | 0.16311 | 0.13007 | 0.10386 | 0.083113 | 0.06641 | 0.053298 | 0.042748 | 0.034381 | 0.027604 | 0.022219 | 0.017839 | 0.014449 | 0.011634 | 0.0093969 | 0.0073773 | 0.0059706 | 0.0046834 | 0.003604 | 0.003119 | 0.0043457 | 0.0022557 | 0.0032184 | 0.0030577 | 0.0023234
Time 16.1 h | 0.11007 | 0.10093 | 0.090991 | 0.082047 | 0.079543 | 0.078788 | 0.081723 | 0.087173 | 0.096992 | 0.11005 | 0.12829 | 0.15003 | 0.17689 | 0.20803 | 0.24102 | 0.27482 | 0.30518 | 0.32877 | 0.34107 | 0.33971 | 0.32434 | 0.2969 | 0.26113 | 0.22177 | 0.1829 | 0.1475 | 0.11733 | 0.092752 | 0.073068 | 0.057802 | 0.045912 | 0.036469 | 0.029123 | 0.02329 | 0.018706 | 0.015032 | 0.012128 | 0.0097734 | 0.0079623 | 0.0064562 | 0.0052363 | 0.0041244 | 0.0033647 | 0.0027806 | 0.0022635 | 0.0029162 | 0.0030408 | 0.0014411 | 0.0021872 | 0.0031647 | 0.0015537
Time 19.1 h | 0.096631 | 0.084665 | 0.074113 | 0.066591 | 0.062691 | 0.058864 | 0.058511 | 0.060476 | 0.065071 | 0.072096 | 0.081916 | 0.094578 | 0.10999 | 0.12716 | 0.1449 | 0.16219 | 0.17646 | 0.18582 | 0.1881 | 0.18242 | 0.16923 | 0.15032 | 0.12827 | 0.10587 | 0.085142 | 0.067241 | 0.052563 | 0.040958 | 0.031878 | 0.024951 | 0.019669 | 0.015531 | 0.012366 | 0.0098704 | 0.0079292 | 0.0063816 | 0.0051741 | 0.0042062 | 0.0034663 | 0.0028605 | 0.0023538 | 0.0018919 | 0.0016185 | 0.0014417 | 0.0019173 | 0.0019251 | 0.0021233 | 0.0015948 | 0.0015646 | 0.0023139 | 0.00072884
 | 0.11394 | 0.093481 | 0.077657 | 0.065696 | 0.057013 | 0.051186 | 0.047934 | 0.047103 | 0.048651 | 0.05264 | 0.059221 | 0.068608 | 0.08103 | 0.096621 | 0.11525 | 0.13624 | 0.15804 | 0.17801 | 0.19272 | 0.1989 | 0.19496 | 0.18175 | 0.16215 | 0.13967 | 0.1172 | 0.096485 | 0.078464 | 0.063316 | 0.050716 | 0.040549 | 0.032372 | 0.025721 | 0.020479 | 0.01627 | 0.012943 | 0.010261 | 0.008146 | 0.0064366 | 0.0051337 | 0.0040747 | 0.00326 | 0.0025268 | 0.0020378 | 0.0016303
 | 0.12912 | 0.10754 | 0.091277 | 0.079538 | 0.071764 | 0.067588 | 0.066822 | 0.06944 | 0.075575 | 0.085504 | 0.099635 | 0.11846 | 0.1425 | 0.17204 | 0.20691 | 0.24592 | 0.28627 | 0.32317 | 0.35038 | 0.36195 | 0.35498 | 0.33106 | 0.29542 | 0.25449 | 0.21357 | 0.17583 | 0.143 | 0.11539 | 0.092429 | 0.0739 | 0.058998 | 0.046874 | 0.037321 | 0.02965 | 0.023587 | 0.018699 | 0.014845 | 0.01173 | 0.0093551 | 0.0074254 | 0.0059407 | 0.0046044 | 0.0037133 | 0.0029707
 | 0.1177 | 0.10593 | 0.097045 | 0.090946 | 0.087643 | 0.087239 | 0.089948 | 0.096102 | 0.10617 | 0.12074 | 0.14056 | 0.16644 | 0.1992 | 0.2394 | 0.28694 | 0.3404 | 0.39622 | 0.44803 | 0.48732 | 0.50564 | 0.49835 | 0.46702 | 0.41851 | 0.36179 | 0.30443 | 0.25118 | 0.20462 | 0.16535 | 0.13261 | 0.10615 | 0.084841 | 0.067487 | 0.053798 | 0.042795 | 0.034092 | 0.027068 | 0.021523 | 0.017036 | 0.013612 | 0.010826 | 0.00868 | 0.0067451 | 0.0054527 | 0.0043739
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625 | 5.7569630370485534 | 5.8239109124113 | 5.9030899869919438 | 6
Time 13.1 h | 0.10929 | 0.1004 | 0.092502 | 0.087998 | 0.086049 | 0.088115 | 0.0938 | 0.10112 | 0.11429 | 0.13106 | 0.15364 | 0.18166 | 0.21583 | 0.25512 | 0.29985 | 0.34774 | 0.3951 | 0.43795 | 0.46981 | 0.48496 | 0.48002 | 0.45469 | 0.41262 | 0.36037 | 0.30465 | 0.25101 | 0.20331 | 0.16311 | 0.13007 | 0.10386 | 0.083113 | 0.06641 | 0.053298 | 0.042748 | 0.034381 | 0.027604 | 0.022219 | 0.017839 | 0.014449 | 0.011634 | 0.0093969 | 0.0073773 | 0.0059706 | 0.0046834 | 0.003604 | 0.003119 | 0.0043457 | 0.0022557 | 0.0032184 | 0.0030577 | 0.0023234
Time 16.1 h | 0.11007 | 0.10093 | 0.090991 | 0.082047 | 0.079543 | 0.078788 | 0.081723 | 0.087173 | 0.096992 | 0.11005 | 0.12829 | 0.15003 | 0.17689 | 0.20803 | 0.24102 | 0.27482 | 0.30518 | 0.32877 | 0.34107 | 0.33971 | 0.32434 | 0.2969 | 0.26113 | 0.22177 | 0.1829 | 0.1475 | 0.11733 | 0.092752 | 0.073068 | 0.057802 | 0.045912 | 0.036469 | 0.029123 | 0.02329 | 0.018706 | 0.015032 | 0.012128 | 0.0097734 | 0.0079623 | 0.0064562 | 0.0052363 | 0.0041244 | 0.0033647 | 0.0027806 | 0.0022635 | 0.0029162 | 0.0030408 | 0.0014411 | 0.0021872 | 0.0031647 | 0.0015537
Time 19.1 h | 0.096631 | 0.084665 | 0.074113 | 0.066591 | 0.062691 | 0.058864 | 0.058511 | 0.060476 | 0.065071 | 0.072096 | 0.081916 | 0.094578 | 0.10999 | 0.12716 | 0.1449 | 0.16219 | 0.17646 | 0.18582 | 0.1881 | 0.18242 | 0.16923 | 0.15032 | 0.12827 | 0.10587 | 0.085142 | 0.067241 | 0.052563 | 0.040958 | 0.031878 | 0.024951 | 0.019669 | 0.015531 | 0.012366 | 0.0098704 | 0.0079292 | 0.0063816 | 0.0051741 | 0.0042062 | 0.0034663 | 0.0028605 | 0.0023538 | 0.0018919 | 0.0016185 | 0.0014417 | 0.0019173 | 0.0019251 | 0.0021233 | 0.0015948 | 0.0015646 | 0.0023139 | 0.00072884
 | 0.11394 | 0.093481 | 0.077657 | 0.065696 | 0.057013 | 0.051186 | 0.047934 | 0.047103 | 0.048651 | 0.05264 | 0.059221 | 0.068608 | 0.08103 | 0.096621 | 0.11525 | 0.13624 | 0.15804 | 0.17801 | 0.19272 | 0.1989 | 0.19496 | 0.18175 | 0.16215 | 0.13967 | 0.1172 | 0.096485 | 0.078464 | 0.063316 | 0.050716 | 0.040549 | 0.032372 | 0.025721 | 0.020479 | 0.01627 | 0.012943 | 0.010261 | 0.008146 | 0.0064366 | 0.0051337 | 0.0040747 | 0.00326 | 0.0025268 | 0.0020378 | 0.0016303
 | 0.12912 | 0.10754 | 0.091277 | 0.079538 | 0.071764 | 0.067588 | 0.066822 | 0.06944 | 0.075575 | 0.085504 | 0.099635 | 0.11846 | 0.1425 | 0.17204 | 0.20691 | 0.24592 | 0.28627 | 0.32317 | 0.35038 | 0.36195 | 0.35498 | 0.33106 | 0.29542 | 0.25449 | 0.21357 | 0.17583 | 0.143 | 0.11539 | 0.092429 | 0.0739 | 0.058998 | 0.046874 | 0.037321 | 0.02965 | 0.023587 | 0.018699 | 0.014845 | 0.01173 | 0.0093551 | 0.0074254 | 0.0059407 | 0.0046044 | 0.0037133 | 0.0029707
 | 0.1177 | 0.10593 | 0.097045 | 0.090946 | 0.087643 | 0.087239 | 0.089948 | 0.096102 | 0.10617 | 0.12074 | 0.14056 | 0.16644 | 0.1992 | 0.2394 | 0.28694 | 0.3404 | 0.39622 | 0.44803 | 0.48732 | 0.50564 | 0.49835 | 0.46702 | 0.41851 | 0.36179 | 0.30443 | 0.25118 | 0.20462 | 0.16535 | 0.13261 | 0.10615 | 0.084841 | 0.067487 | 0.053798 | 0.042795 | 0.034092 | 0.027068 | 0.021523 | 0.017036 | 0.013612 | 0.010826 | 0.00868 | 0.0067451 | 0.0054527 | 0.0043739
increase
Drying time
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 4.4 h | 0.003361 | 0.0040465 | 0.0025946 | 0.0026214 | 0.0044553 | 0.0026323 | 0.0029087 | 0.0037366 | 0.0047149 | 0.0057564 | 0.0073092 | 0.0089323 | 0.011077 | 0.013703 | 0.017072 | 0.021236 | 0.026498 | 0.03301 | 0.041177 | 0.051293 | 0.064055 | 0.079683 | 0.099225 | 0.12384 | 0.15437 | 0.19231 | 0.23844 | 0.29224 | 0.35137 | 0.41125 | 0.46468 | 0.50437 | 0.52462 | 0.52356 | 0.50251 | 0.46708 | 0.42347 | 0.37253 | 0.32299 | 0.27537 | 0.23266 | 0.1978 | 0.16633 | 0.14424 | 0.12252 | 0.10889 | 0.097697 | 0.093226 | 0.09166 | 0.092155 | 0.093069
Time 4.5 h | 0.0034249 | 0.0041377 | 0.0027283 | 0.0027814 | 0.0046887 | 0.0029259 | 0.0032733 | 0.0042363 | 0.0053358 | 0.0065296 | 0.0082746 | 0.010141 | 0.012601 | 0.015587 | 0.01947 | 0.024233 | 0.030264 | 0.037737 | 0.047171 | 0.058901 | 0.073903 | 0.092482 | 0.11581 | 0.14538 | 0.18208 | 0.2274 | 0.28124 | 0.3416 | 0.40418 | 0.46243 | 0.50811 | 0.53414 | 0.53731 | 0.51809 | 0.48197 | 0.43577 | 0.38293 | 0.3311 | 0.28177 | 0.23748 | 0.20077 | 0.16802 | 0.14416 | 0.11848 | 0.11171 | 0.10187 | 0.09506 | 0.092791 | 0.094429 | 0.097904 | 0.10371
Time 4.6 h | 0.0036091 | 0.0043103 | 0.0029276 | 0.0030069 | 0.0049863 | 0.0032919 | 0.0037411 | 0.0048548 | 0.006126 | 0.0075084 | 0.0095269 | 0.011725 | 0.014579 | 0.018062 | 0.02259 | 0.028182 | 0.035264 | 0.044078 | 0.055262 | 0.069282 | 0.08729 | 0.10998 | 0.13862 | 0.17503 | 0.21981 | 0.27381 | 0.33549 | 0.40061 | 0.46249 | 0.51282 | 0.5432 | 0.54939 | 0.53148 | 0.49439 | 0.44517 | 0.39079 | 0.33577 | 0.28667 | 0.23998 | 0.20372 | 0.17181 | 0.14592 | 0.12721 | 0.11818 | 0.10489 | 0.098824 | 0.097908 | 0.099283 | 0.10394 | 0.11202 | 0.12205
Time 4.7 h | 0.0040254 | 0.0023197 | 0.0060162 | 0.0032852 | 0.0053707 | 0.003773 | 0.0043423 | 0.005672 | 0.0071637 | 0.0088006 | 0.011181 | 0.013771 | 0.017191 | 0.021329 | 0.026752 | 0.033478 | 0.04203 | 0.052722 | 0.066401 | 0.083738 | 0.10613 | 0.13436 | 0.17037 | 0.21573 | 0.27035 | 0.33347 | 0.40107 | 0.46603 | 0.51937 | 0.55253 | 0.55996 | 0.54171 | 0.50311 | 0.45097 | 0.39455 | 0.33876 | 0.28887 | 0.24104 | 0.20387 | 0.17339 | 0.14886 | 0.13225 | 0.11753 | 0.11 | 0.10482 | 0.10438 | 0.1074 | 0.1138 | 0.12408 | 0.13639 | 0.14993
Time 4.8 h | 0.0039527 | 0.0047899 | 0.0034912 | 0.0036734 | 0.0059298 | 0.004463 | 0.0052209 | 0.0068624 | 0.0086692 | 0.010687 | 0.013626 | 0.016821 | 0.021041 | 0.026213 | 0.03302 | 0.041432 | 0.052261 | 0.065911 | 0.083539 | 0.10606 | 0.13545 | 0.17252 | 0.21928 | 0.27665 | 0.34237 | 0.41271 | 0.47977 | 0.53353 | 0.56486 | 0.56874 | 0.5461 | 0.50317 | 0.44813 | 0.38872 | 0.33224 | 0.28045 | 0.23905 | 0.20172 | 0.17254 | 0.14998 | 0.13098 | 0.12232 | 0.11386 | 0.11155 | 0.11412 | 0.11812 | 0.12658 | 0.13848 | 0.15361 | 0.17099 | 0.18866
Time 4.9 h | 0.0043188 | 0.0052277 | 0.0039868 | 0.0042288 | 0.0066529 | 0.0053479 | 0.0063485 | 0.0083796 | 0.010604 | 0.013107 | 0.016765 | 0.020805 | 0.02609 | 0.032614 | 0.041224 | 0.052023 | 0.065957 | 0.083695 | 0.10683 | 0.13651 | 0.17523 | 0.22353 | 0.28261 | 0.35128 | 0.42376 | 0.49222 | 0.54583 | 0.5749 | 0.57491 | 0.54814 | 0.50162 | 0.4443 | 0.38384 | 0.32634 | 0.27564 | 0.23272 | 0.19908 | 0.1717 | 0.1495 | 0.13558 | 0.12502 | 0.11938 | 0.11838 | 0.12437 | 0.13062 | 0.13992 | 0.15492 | 0.1723 | 0.19073 | 0.21306 | 0.23377
Time 5 h | 0.0049133 | 0.0058042 | 0.0046338 | 0.0049712 | 0.0076664 | 0.0066031 | 0.0079479 | 0.010533 | 0.013358 | 0.01657 | 0.021298 | 0.026516 | 0.033397 | 0.041936 | 0.053268 | 0.067599 | 0.086277 | 0.11024 | 0.14161 | 0.18183 | 0.2334 | 0.29549 | 0.36662 | 0.44167 | 0.50969 | 0.5605 | 0.58424 | 0.57776 | 0.54458 | 0.49349 | 0.43374 | 0.3731 | 0.31662 | 0.26729 | 0.22557 | 0.19336 | 0.16657 | 0.14984 | 0.13717 | 0.1291 | 0.12498 | 0.12759 | 0.13258 | 0.1444 | 0.15601 | 0.17194 | 0.19151 | 0.21398 | 0.23437 | 0.26066 | 0.28203
Time 5.1 h | 0.0051615 | 0.0065394 | 0.0054847 | 0.0059393 | 0.0090001 | 0.0082333 | 0.010001 | 0.013324 | 0.016914 | 0.021051 | 0.027147 | 0.03396 | 0.042946 | 0.054219 | 0.069319 | 0.088593 | 0.11394 | 0.14657 | 0.18909 | 0.24251 | 0.30812 | 0.38167 | 0.45767 | 0.52576 | 0.57299 | 0.59132 | 0.57824 | 0.5397 | 0.48481 | 0.42344 | 0.36292 | 0.3071 | 0.25963 | 0.22024 | 0.18846 | 0.16573 | 0.14947 | 0.13846 | 0.13032 | 0.13113 | 0.1334 | 0.1422 | 0.15348 | 0.16641 | 0.18887 | 0.2096 | 0.23238 | 0.25699 | 0.28353 | 0.3063 | 0.32649
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 4.4 h | 0.003361 | 0.0040465 | 0.0025946 | 0.0026214 | 0.0044553 | 0.0026323 | 0.0029087 | 0.0037366 | 0.0047149 | 0.0057564 | 0.0073092 | 0.0089323 | 0.011077 | 0.013703 | 0.017072 | 0.021236 | 0.026498 | 0.03301 | 0.041177 | 0.051293 | 0.064055 | 0.079683 | 0.099225 | 0.12384 | 0.15437 | 0.19231 | 0.23844 | 0.29224 | 0.35137 | 0.41125 | 0.46468 | 0.50437 | 0.52462 | 0.52356 | 0.50251 | 0.46708 | 0.42347 | 0.37253 | 0.32299 | 0.27537 | 0.23266 | 0.1978 | 0.16633 | 0.14424 | 0.12252 | 0.10889 | 0.097697 | 0.093226 | 0.09166 | 0.092155 | 0.093069
Time 4.5 h | 0.0034249 | 0.0041377 | 0.0027283 | 0.0027814 | 0.0046887 | 0.0029259 | 0.0032733 | 0.0042363 | 0.0053358 | 0.0065296 | 0.0082746 | 0.010141 | 0.012601 | 0.015587 | 0.01947 | 0.024233 | 0.030264 | 0.037737 | 0.047171 | 0.058901 | 0.073903 | 0.092482 | 0.11581 | 0.14538 | 0.18208 | 0.2274 | 0.28124 | 0.3416 | 0.40418 | 0.46243 | 0.50811 | 0.53414 | 0.53731 | 0.51809 | 0.48197 | 0.43577 | 0.38293 | 0.3311 | 0.28177 | 0.23748 | 0.20077 | 0.16802 | 0.14416 | 0.11848 | 0.11171 | 0.10187 | 0.09506 | 0.092791 | 0.094429 | 0.097904 | 0.10371
Time 4.6 h | 0.0036091 | 0.0043103 | 0.0029276 | 0.0030069 | 0.0049863 | 0.0032919 | 0.0037411 | 0.0048548 | 0.006126 | 0.0075084 | 0.0095269 | 0.011725 | 0.014579 | 0.018062 | 0.02259 | 0.028182 | 0.035264 | 0.044078 | 0.055262 | 0.069282 | 0.08729 | 0.10998 | 0.13862 | 0.17503 | 0.21981 | 0.27381 | 0.33549 | 0.40061 | 0.46249 | 0.51282 | 0.5432 | 0.54939 | 0.53148 | 0.49439 | 0.44517 | 0.39079 | 0.33577 | 0.28667 | 0.23998 | 0.20372 | 0.17181 | 0.14592 | 0.12721 | 0.11818 | 0.10489 | 0.098824 | 0.097908 | 0.099283 | 0.10394 | 0.11202 | 0.12205
Time 4.7 h | 0.0040254 | 0.0023197 | 0.0060162 | 0.0032852 | 0.0053707 | 0.003773 | 0.0043423 | 0.005672 | 0.0071637 | 0.0088006 | 0.011181 | 0.013771 | 0.017191 | 0.021329 | 0.026752 | 0.033478 | 0.04203 | 0.052722 | 0.066401 | 0.083738 | 0.10613 | 0.13436 | 0.17037 | 0.21573 | 0.27035 | 0.33347 | 0.40107 | 0.46603 | 0.51937 | 0.55253 | 0.55996 | 0.54171 | 0.50311 | 0.45097 | 0.39455 | 0.33876 | 0.28887 | 0.24104 | 0.20387 | 0.17339 | 0.14886 | 0.13225 | 0.11753 | 0.11 | 0.10482 | 0.10438 | 0.1074 | 0.1138 | 0.12408 | 0.13639 | 0.14993
Time 4.8 h | 0.0039527 | 0.0047899 | 0.0034912 | 0.0036734 | 0.0059298 | 0.004463 | 0.0052209 | 0.0068624 | 0.0086692 | 0.010687 | 0.013626 | 0.016821 | 0.021041 | 0.026213 | 0.03302 | 0.041432 | 0.052261 | 0.065911 | 0.083539 | 0.10606 | 0.13545 | 0.17252 | 0.21928 | 0.27665 | 0.34237 | 0.41271 | 0.47977 | 0.53353 | 0.56486 | 0.56874 | 0.5461 | 0.50317 | 0.44813 | 0.38872 | 0.33224 | 0.28045 | 0.23905 | 0.20172 | 0.17254 | 0.14998 | 0.13098 | 0.12232 | 0.11386 | 0.11155 | 0.11412 | 0.11812 | 0.12658 | 0.13848 | 0.15361 | 0.17099 | 0.18866
Time 4.9 h | 0.0043188 | 0.0052277 | 0.0039868 | 0.0042288 | 0.0066529 | 0.0053479 | 0.0063485 | 0.0083796 | 0.010604 | 0.013107 | 0.016765 | 0.020805 | 0.02609 | 0.032614 | 0.041224 | 0.052023 | 0.065957 | 0.083695 | 0.10683 | 0.13651 | 0.17523 | 0.22353 | 0.28261 | 0.35128 | 0.42376 | 0.49222 | 0.54583 | 0.5749 | 0.57491 | 0.54814 | 0.50162 | 0.4443 | 0.38384 | 0.32634 | 0.27564 | 0.23272 | 0.19908 | 0.1717 | 0.1495 | 0.13558 | 0.12502 | 0.11938 | 0.11838 | 0.12437 | 0.13062 | 0.13992 | 0.15492 | 0.1723 | 0.19073 | 0.21306 | 0.23377
Time 5 h | 0.0049133 | 0.0058042 | 0.0046338 | 0.0049712 | 0.0076664 | 0.0066031 | 0.0079479 | 0.010533 | 0.013358 | 0.01657 | 0.021298 | 0.026516 | 0.033397 | 0.041936 | 0.053268 | 0.067599 | 0.086277 | 0.11024 | 0.14161 | 0.18183 | 0.2334 | 0.29549 | 0.36662 | 0.44167 | 0.50969 | 0.5605 | 0.58424 | 0.57776 | 0.54458 | 0.49349 | 0.43374 | 0.3731 | 0.31662 | 0.26729 | 0.22557 | 0.19336 | 0.16657 | 0.14984 | 0.13717 | 0.1291 | 0.12498 | 0.12759 | 0.13258 | 0.1444 | 0.15601 | 0.17194 | 0.19151 | 0.21398 | 0.23437 | 0.26066 | 0.28203
Time 5.1 h | 0.0051615 | 0.0065394 | 0.0054847 | 0.0059393 | 0.0090001 | 0.0082333 | 0.010001 | 0.013324 | 0.016914 | 0.021051 | 0.027147 | 0.03396 | 0.042946 | 0.054219 | 0.069319 | 0.088593 | 0.11394 | 0.14657 | 0.18909 | 0.24251 | 0.30812 | 0.38167 | 0.45767 | 0.52576 | 0.57299 | 0.59132 | 0.57824 | 0.5397 | 0.48481 | 0.42344 | 0.36292 | 0.3071 | 0.25963 | 0.22024 | 0.18846 | 0.16573 | 0.14947 | 0.13846 | 0.13032 | 0.13113 | 0.1334 | 0.1422 | 0.15348 | 0.16641 | 0.18887 | 0.2096 | 0.23238 | 0.25699 | 0.28353 | 0.3063 | 0.32649
-38 oC
-18 oC
Monitoring Phase Behaviour (ice nucleation temperature and solidification end points by using
temperature calibration for predicting temperature of the product in primary drying
Surrogate drying rate (from )
Notes: Eg temperature cycling below to calibrate Fpeak for temperature (in order to predict product temperatures in primary drying) and the compensate cpeak for temperature so that it can used to monitor ice content and hence drying rate
Eg 2 left graph: monitoring phase behaviour (ice nucleation temperatures and solidufication end points)

## Slide 19: Through Vial Impedance Spectroscopy (TVIS)
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034
Time 0.3 h | 0.27396 | 0.24912 | 0.22663 | 0.2063 | 0.18797 | 0.17147 | 0.15668 | 0.14349 | 0.13183 | 0.12162 | 0.11284 | 0.10548 | 0.099574 | 0.095186 | 0.092431 | 0.091474 | 0.092551 | 0.095973 | 0.10215 | 0.11162 | 0.12501 | 0.14319 | 0.16715 | 0.198 | 0.23699 | 0.28544 | 0.34395 | 0.41214 | 0.48865 | 0.56683 | 0.63782 | 0.69048 | 0.71191 | 0.69725 | 0.65001 | 0.57934 | 0.49898 | 0.41726 | 0.34526 | 0.28116 | 0.22886 | 0.17991 | 0.14643 | 0.1181 | 0.089415 | 0.072054
Time 4 h | 0.099081 | 0.096153 | 0.095569 | 0.097591 | 0.10258 | 0.11101 | 0.12349 | 0.14075 | 0.16366 | 0.19315 | 0.23007 | 0.27492 | 0.32735 | 0.38533 | 0.44432 | 0.49685 | 0.53348 | 0.54588 | 0.53067 | 0.49125 | 0.43601 | 0.37424 | 0.31336 | 0.25794 | 0.20992 | 0.16948 | 0.13623 | 0.10926 | 0.087318 | 0.069838 | 0.055881 | 0.044566 | 0.035658 | 0.0285 | 0.02283 | 0.018246 | 0.014616 | 0.011669 | 0.0094101 | 0.0075633 | 0.0061328 | 0.0048349
Time 0.3 h | 0.26705 | 0.24478 | 0.21978 | 0.1979 | 0.17863 | 0.15615 | 0.14276 | 0.12639 | 0.11756 | 0.11024 | 0.10127 | 0.1004 | 0.098391 | 0.10089 | 0.10904 | 0.11009 | 0.11451 | 0.12496 | 0.13439 | 0.14683 | 0.16234 | 0.18119 | 0.20238 | 0.22944 | 0.26256 | 0.30331 | 0.35113 | 0.4089 | 0.47307 | 0.54005 | 0.60366 | 0.65431 | 0.67974 | 0.67269 | 0.63391 | 0.56955 | 0.49367 | 0.41413 | 0.34323 | 0.27998 | 0.22777
Time 4 h | 0.093429 | 0.092843 | 0.096342 | 0.10075 | 0.11232 | 0.12322 | 0.1435 | 0.1596 | 0.19684 | 0.23027 | 0.27214 | 0.31647 | 0.36246 | 0.41025 | 0.45421 | 0.48896 | 0.51147 | 0.51695 | 0.50285 | 0.46926 | 0.42098 | 0.36431 | 0.30625 | 0.25218 | 0.20511 | 0.16565 | 0.13355 | 0.10767 | 0.086605 | 0.069689 | 0.056026 | 0.045028 | 0.036377 | 0.029122 | 0.023205 | 0.018299 | 0.014688 | 0.011778 | 0.0095536 | 0.0077519 | 0.0063112
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034
Time 0.3 h | 0.27396 | 0.24912 | 0.22663 | 0.2063 | 0.18797 | 0.17147 | 0.15668 | 0.14349 | 0.13183 | 0.12162 | 0.11284 | 0.10548 | 0.099574 | 0.095186 | 0.092431 | 0.091474 | 0.092551 | 0.095973 | 0.10215 | 0.11162 | 0.12501 | 0.14319 | 0.16715 | 0.198 | 0.23699 | 0.28544 | 0.34395 | 0.41214 | 0.48865 | 0.56683 | 0.63782 | 0.69048 | 0.71191 | 0.69725 | 0.65001 | 0.57934 | 0.49898 | 0.41726 | 0.34526 | 0.28116 | 0.22886 | 0.17991 | 0.14643 | 0.1181 | 0.089415 | 0.072054
Time 4 h | 0.099081 | 0.096153 | 0.095569 | 0.097591 | 0.10258 | 0.11101 | 0.12349 | 0.14075 | 0.16366 | 0.19315 | 0.23007 | 0.27492 | 0.32735 | 0.38533 | 0.44432 | 0.49685 | 0.53348 | 0.54588 | 0.53067 | 0.49125 | 0.43601 | 0.37424 | 0.31336 | 0.25794 | 0.20992 | 0.16948 | 0.13623 | 0.10926 | 0.087318 | 0.069838 | 0.055881 | 0.044566 | 0.035658 | 0.0285 | 0.02283 | 0.018246 | 0.014616 | 0.011669 | 0.0094101 | 0.0075633 | 0.0061328 | 0.0048349
Time 0.3 h | 0.26705 | 0.24478 | 0.21978 | 0.1979 | 0.17863 | 0.15615 | 0.14276 | 0.12639 | 0.11756 | 0.11024 | 0.10127 | 0.1004 | 0.098391 | 0.10089 | 0.10904 | 0.11009 | 0.11451 | 0.12496 | 0.13439 | 0.14683 | 0.16234 | 0.18119 | 0.20238 | 0.22944 | 0.26256 | 0.30331 | 0.35113 | 0.4089 | 0.47307 | 0.54005 | 0.60366 | 0.65431 | 0.67974 | 0.67269 | 0.63391 | 0.56955 | 0.49367 | 0.41413 | 0.34323 | 0.27998 | 0.22777
Time 4 h | 0.093429 | 0.092843 | 0.096342 | 0.10075 | 0.11232 | 0.12322 | 0.1435 | 0.1596 | 0.19684 | 0.23027 | 0.27214 | 0.31647 | 0.36246 | 0.41025 | 0.45421 | 0.48896 | 0.51147 | 0.51695 | 0.50285 | 0.46926 | 0.42098 | 0.36431 | 0.30625 | 0.25218 | 0.20511 | 0.16565 | 0.13355 | 0.10767 | 0.086605 | 0.069689 | 0.056026 | 0.045028 | 0.036377 | 0.029122 | 0.023205 | 0.018299 | 0.014688 | 0.011778 | 0.0095536 | 0.0077519 | 0.0063112
FPEAK
FPEAK
Liquid State
Solid State
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625 | 5.7569630370485534 | 5.8239109124113 | 5.9030899869919438 | 6
Time 13.1 h | 0.10929 | 0.1004 | 0.092502 | 0.087998 | 0.086049 | 0.088115 | 0.0938 | 0.10112 | 0.11429 | 0.13106 | 0.15364 | 0.18166 | 0.21583 | 0.25512 | 0.29985 | 0.34774 | 0.3951 | 0.43795 | 0.46981 | 0.48496 | 0.48002 | 0.45469 | 0.41262 | 0.36037 | 0.30465 | 0.25101 | 0.20331 | 0.16311 | 0.13007 | 0.10386 | 0.083113 | 0.06641 | 0.053298 | 0.042748 | 0.034381 | 0.027604 | 0.022219 | 0.017839 | 0.014449 | 0.011634 | 0.0093969 | 0.0073773 | 0.0059706 | 0.0046834 | 0.003604 | 0.003119 | 0.0043457 | 0.0022557 | 0.0032184 | 0.0030577 | 0.0023234
Time 16.1 h | 0.11007 | 0.10093 | 0.090991 | 0.082047 | 0.079543 | 0.078788 | 0.081723 | 0.087173 | 0.096992 | 0.11005 | 0.12829 | 0.15003 | 0.17689 | 0.20803 | 0.24102 | 0.27482 | 0.30518 | 0.32877 | 0.34107 | 0.33971 | 0.32434 | 0.2969 | 0.26113 | 0.22177 | 0.1829 | 0.1475 | 0.11733 | 0.092752 | 0.073068 | 0.057802 | 0.045912 | 0.036469 | 0.029123 | 0.02329 | 0.018706 | 0.015032 | 0.012128 | 0.0097734 | 0.0079623 | 0.0064562 | 0.0052363 | 0.0041244 | 0.0033647 | 0.0027806 | 0.0022635 | 0.0029162 | 0.0030408 | 0.0014411 | 0.0021872 | 0.0031647 | 0.0015537
Time 19.1 h | 0.096631 | 0.084665 | 0.074113 | 0.066591 | 0.062691 | 0.058864 | 0.058511 | 0.060476 | 0.065071 | 0.072096 | 0.081916 | 0.094578 | 0.10999 | 0.12716 | 0.1449 | 0.16219 | 0.17646 | 0.18582 | 0.1881 | 0.18242 | 0.16923 | 0.15032 | 0.12827 | 0.10587 | 0.085142 | 0.067241 | 0.052563 | 0.040958 | 0.031878 | 0.024951 | 0.019669 | 0.015531 | 0.012366 | 0.0098704 | 0.0079292 | 0.0063816 | 0.0051741 | 0.0042062 | 0.0034663 | 0.0028605 | 0.0023538 | 0.0018919 | 0.0016185 | 0.0014417 | 0.0019173 | 0.0019251 | 0.0021233 | 0.0015948 | 0.0015646 | 0.0023139 | 0.00072884
 | 0.11394 | 0.093481 | 0.077657 | 0.065696 | 0.057013 | 0.051186 | 0.047934 | 0.047103 | 0.048651 | 0.05264 | 0.059221 | 0.068608 | 0.08103 | 0.096621 | 0.11525 | 0.13624 | 0.15804 | 0.17801 | 0.19272 | 0.1989 | 0.19496 | 0.18175 | 0.16215 | 0.13967 | 0.1172 | 0.096485 | 0.078464 | 0.063316 | 0.050716 | 0.040549 | 0.032372 | 0.025721 | 0.020479 | 0.01627 | 0.012943 | 0.010261 | 0.008146 | 0.0064366 | 0.0051337 | 0.0040747 | 0.00326 | 0.0025268 | 0.0020378 | 0.0016303
 | 0.12912 | 0.10754 | 0.091277 | 0.079538 | 0.071764 | 0.067588 | 0.066822 | 0.06944 | 0.075575 | 0.085504 | 0.099635 | 0.11846 | 0.1425 | 0.17204 | 0.20691 | 0.24592 | 0.28627 | 0.32317 | 0.35038 | 0.36195 | 0.35498 | 0.33106 | 0.29542 | 0.25449 | 0.21357 | 0.17583 | 0.143 | 0.11539 | 0.092429 | 0.0739 | 0.058998 | 0.046874 | 0.037321 | 0.02965 | 0.023587 | 0.018699 | 0.014845 | 0.01173 | 0.0093551 | 0.0074254 | 0.0059407 | 0.0046044 | 0.0037133 | 0.0029707
 | 0.1177 | 0.10593 | 0.097045 | 0.090946 | 0.087643 | 0.087239 | 0.089948 | 0.096102 | 0.10617 | 0.12074 | 0.14056 | 0.16644 | 0.1992 | 0.2394 | 0.28694 | 0.3404 | 0.39622 | 0.44803 | 0.48732 | 0.50564 | 0.49835 | 0.46702 | 0.41851 | 0.36179 | 0.30443 | 0.25118 | 0.20462 | 0.16535 | 0.13261 | 0.10615 | 0.084841 | 0.067487 | 0.053798 | 0.042795 | 0.034092 | 0.027068 | 0.021523 | 0.017036 | 0.013612 | 0.010826 | 0.00868 | 0.0067451 | 0.0054527 | 0.0043739
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625 | 5.7569630370485534 | 5.8239109124113 | 5.9030899869919438 | 6
Time 13.1 h | 0.10929 | 0.1004 | 0.092502 | 0.087998 | 0.086049 | 0.088115 | 0.0938 | 0.10112 | 0.11429 | 0.13106 | 0.15364 | 0.18166 | 0.21583 | 0.25512 | 0.29985 | 0.34774 | 0.3951 | 0.43795 | 0.46981 | 0.48496 | 0.48002 | 0.45469 | 0.41262 | 0.36037 | 0.30465 | 0.25101 | 0.20331 | 0.16311 | 0.13007 | 0.10386 | 0.083113 | 0.06641 | 0.053298 | 0.042748 | 0.034381 | 0.027604 | 0.022219 | 0.017839 | 0.014449 | 0.011634 | 0.0093969 | 0.0073773 | 0.0059706 | 0.0046834 | 0.003604 | 0.003119 | 0.0043457 | 0.0022557 | 0.0032184 | 0.0030577 | 0.0023234
Time 16.1 h | 0.11007 | 0.10093 | 0.090991 | 0.082047 | 0.079543 | 0.078788 | 0.081723 | 0.087173 | 0.096992 | 0.11005 | 0.12829 | 0.15003 | 0.17689 | 0.20803 | 0.24102 | 0.27482 | 0.30518 | 0.32877 | 0.34107 | 0.33971 | 0.32434 | 0.2969 | 0.26113 | 0.22177 | 0.1829 | 0.1475 | 0.11733 | 0.092752 | 0.073068 | 0.057802 | 0.045912 | 0.036469 | 0.029123 | 0.02329 | 0.018706 | 0.015032 | 0.012128 | 0.0097734 | 0.0079623 | 0.0064562 | 0.0052363 | 0.0041244 | 0.0033647 | 0.0027806 | 0.0022635 | 0.0029162 | 0.0030408 | 0.0014411 | 0.0021872 | 0.0031647 | 0.0015537
Time 19.1 h | 0.096631 | 0.084665 | 0.074113 | 0.066591 | 0.062691 | 0.058864 | 0.058511 | 0.060476 | 0.065071 | 0.072096 | 0.081916 | 0.094578 | 0.10999 | 0.12716 | 0.1449 | 0.16219 | 0.17646 | 0.18582 | 0.1881 | 0.18242 | 0.16923 | 0.15032 | 0.12827 | 0.10587 | 0.085142 | 0.067241 | 0.052563 | 0.040958 | 0.031878 | 0.024951 | 0.019669 | 0.015531 | 0.012366 | 0.0098704 | 0.0079292 | 0.0063816 | 0.0051741 | 0.0042062 | 0.0034663 | 0.0028605 | 0.0023538 | 0.0018919 | 0.0016185 | 0.0014417 | 0.0019173 | 0.0019251 | 0.0021233 | 0.0015948 | 0.0015646 | 0.0023139 | 0.00072884
 | 0.11394 | 0.093481 | 0.077657 | 0.065696 | 0.057013 | 0.051186 | 0.047934 | 0.047103 | 0.048651 | 0.05264 | 0.059221 | 0.068608 | 0.08103 | 0.096621 | 0.11525 | 0.13624 | 0.15804 | 0.17801 | 0.19272 | 0.1989 | 0.19496 | 0.18175 | 0.16215 | 0.13967 | 0.1172 | 0.096485 | 0.078464 | 0.063316 | 0.050716 | 0.040549 | 0.032372 | 0.025721 | 0.020479 | 0.01627 | 0.012943 | 0.010261 | 0.008146 | 0.0064366 | 0.0051337 | 0.0040747 | 0.00326 | 0.0025268 | 0.0020378 | 0.0016303
 | 0.12912 | 0.10754 | 0.091277 | 0.079538 | 0.071764 | 0.067588 | 0.066822 | 0.06944 | 0.075575 | 0.085504 | 0.099635 | 0.11846 | 0.1425 | 0.17204 | 0.20691 | 0.24592 | 0.28627 | 0.32317 | 0.35038 | 0.36195 | 0.35498 | 0.33106 | 0.29542 | 0.25449 | 0.21357 | 0.17583 | 0.143 | 0.11539 | 0.092429 | 0.0739 | 0.058998 | 0.046874 | 0.037321 | 0.02965 | 0.023587 | 0.018699 | 0.014845 | 0.01173 | 0.0093551 | 0.0074254 | 0.0059407 | 0.0046044 | 0.0037133 | 0.0029707
 | 0.1177 | 0.10593 | 0.097045 | 0.090946 | 0.087643 | 0.087239 | 0.089948 | 0.096102 | 0.10617 | 0.12074 | 0.14056 | 0.16644 | 0.1992 | 0.2394 | 0.28694 | 0.3404 | 0.39622 | 0.44803 | 0.48732 | 0.50564 | 0.49835 | 0.46702 | 0.41851 | 0.36179 | 0.30443 | 0.25118 | 0.20462 | 0.16535 | 0.13261 | 0.10615 | 0.084841 | 0.067487 | 0.053798 | 0.042795 | 0.034092 | 0.027068 | 0.021523 | 0.017036 | 0.013612 | 0.010826 | 0.00868 | 0.0067451 | 0.0054527 | 0.0043739
increase
Drying time
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 4.4 h | 0.003361 | 0.0040465 | 0.0025946 | 0.0026214 | 0.0044553 | 0.0026323 | 0.0029087 | 0.0037366 | 0.0047149 | 0.0057564 | 0.0073092 | 0.0089323 | 0.011077 | 0.013703 | 0.017072 | 0.021236 | 0.026498 | 0.03301 | 0.041177 | 0.051293 | 0.064055 | 0.079683 | 0.099225 | 0.12384 | 0.15437 | 0.19231 | 0.23844 | 0.29224 | 0.35137 | 0.41125 | 0.46468 | 0.50437 | 0.52462 | 0.52356 | 0.50251 | 0.46708 | 0.42347 | 0.37253 | 0.32299 | 0.27537 | 0.23266 | 0.1978 | 0.16633 | 0.14424 | 0.12252 | 0.10889 | 0.097697 | 0.093226 | 0.09166 | 0.092155 | 0.093069
Time 4.5 h | 0.0034249 | 0.0041377 | 0.0027283 | 0.0027814 | 0.0046887 | 0.0029259 | 0.0032733 | 0.0042363 | 0.0053358 | 0.0065296 | 0.0082746 | 0.010141 | 0.012601 | 0.015587 | 0.01947 | 0.024233 | 0.030264 | 0.037737 | 0.047171 | 0.058901 | 0.073903 | 0.092482 | 0.11581 | 0.14538 | 0.18208 | 0.2274 | 0.28124 | 0.3416 | 0.40418 | 0.46243 | 0.50811 | 0.53414 | 0.53731 | 0.51809 | 0.48197 | 0.43577 | 0.38293 | 0.3311 | 0.28177 | 0.23748 | 0.20077 | 0.16802 | 0.14416 | 0.11848 | 0.11171 | 0.10187 | 0.09506 | 0.092791 | 0.094429 | 0.097904 | 0.10371
Time 4.6 h | 0.0036091 | 0.0043103 | 0.0029276 | 0.0030069 | 0.0049863 | 0.0032919 | 0.0037411 | 0.0048548 | 0.006126 | 0.0075084 | 0.0095269 | 0.011725 | 0.014579 | 0.018062 | 0.02259 | 0.028182 | 0.035264 | 0.044078 | 0.055262 | 0.069282 | 0.08729 | 0.10998 | 0.13862 | 0.17503 | 0.21981 | 0.27381 | 0.33549 | 0.40061 | 0.46249 | 0.51282 | 0.5432 | 0.54939 | 0.53148 | 0.49439 | 0.44517 | 0.39079 | 0.33577 | 0.28667 | 0.23998 | 0.20372 | 0.17181 | 0.14592 | 0.12721 | 0.11818 | 0.10489 | 0.098824 | 0.097908 | 0.099283 | 0.10394 | 0.11202 | 0.12205
Time 4.7 h | 0.0040254 | 0.0023197 | 0.0060162 | 0.0032852 | 0.0053707 | 0.003773 | 0.0043423 | 0.005672 | 0.0071637 | 0.0088006 | 0.011181 | 0.013771 | 0.017191 | 0.021329 | 0.026752 | 0.033478 | 0.04203 | 0.052722 | 0.066401 | 0.083738 | 0.10613 | 0.13436 | 0.17037 | 0.21573 | 0.27035 | 0.33347 | 0.40107 | 0.46603 | 0.51937 | 0.55253 | 0.55996 | 0.54171 | 0.50311 | 0.45097 | 0.39455 | 0.33876 | 0.28887 | 0.24104 | 0.20387 | 0.17339 | 0.14886 | 0.13225 | 0.11753 | 0.11 | 0.10482 | 0.10438 | 0.1074 | 0.1138 | 0.12408 | 0.13639 | 0.14993
Time 4.8 h | 0.0039527 | 0.0047899 | 0.0034912 | 0.0036734 | 0.0059298 | 0.004463 | 0.0052209 | 0.0068624 | 0.0086692 | 0.010687 | 0.013626 | 0.016821 | 0.021041 | 0.026213 | 0.03302 | 0.041432 | 0.052261 | 0.065911 | 0.083539 | 0.10606 | 0.13545 | 0.17252 | 0.21928 | 0.27665 | 0.34237 | 0.41271 | 0.47977 | 0.53353 | 0.56486 | 0.56874 | 0.5461 | 0.50317 | 0.44813 | 0.38872 | 0.33224 | 0.28045 | 0.23905 | 0.20172 | 0.17254 | 0.14998 | 0.13098 | 0.12232 | 0.11386 | 0.11155 | 0.11412 | 0.11812 | 0.12658 | 0.13848 | 0.15361 | 0.17099 | 0.18866
Time 4.9 h | 0.0043188 | 0.0052277 | 0.0039868 | 0.0042288 | 0.0066529 | 0.0053479 | 0.0063485 | 0.0083796 | 0.010604 | 0.013107 | 0.016765 | 0.020805 | 0.02609 | 0.032614 | 0.041224 | 0.052023 | 0.065957 | 0.083695 | 0.10683 | 0.13651 | 0.17523 | 0.22353 | 0.28261 | 0.35128 | 0.42376 | 0.49222 | 0.54583 | 0.5749 | 0.57491 | 0.54814 | 0.50162 | 0.4443 | 0.38384 | 0.32634 | 0.27564 | 0.23272 | 0.19908 | 0.1717 | 0.1495 | 0.13558 | 0.12502 | 0.11938 | 0.11838 | 0.12437 | 0.13062 | 0.13992 | 0.15492 | 0.1723 | 0.19073 | 0.21306 | 0.23377
Time 5 h | 0.0049133 | 0.0058042 | 0.0046338 | 0.0049712 | 0.0076664 | 0.0066031 | 0.0079479 | 0.010533 | 0.013358 | 0.01657 | 0.021298 | 0.026516 | 0.033397 | 0.041936 | 0.053268 | 0.067599 | 0.086277 | 0.11024 | 0.14161 | 0.18183 | 0.2334 | 0.29549 | 0.36662 | 0.44167 | 0.50969 | 0.5605 | 0.58424 | 0.57776 | 0.54458 | 0.49349 | 0.43374 | 0.3731 | 0.31662 | 0.26729 | 0.22557 | 0.19336 | 0.16657 | 0.14984 | 0.13717 | 0.1291 | 0.12498 | 0.12759 | 0.13258 | 0.1444 | 0.15601 | 0.17194 | 0.19151 | 0.21398 | 0.23437 | 0.26066 | 0.28203
Time 5.1 h | 0.0051615 | 0.0065394 | 0.0054847 | 0.0059393 | 0.0090001 | 0.0082333 | 0.010001 | 0.013324 | 0.016914 | 0.021051 | 0.027147 | 0.03396 | 0.042946 | 0.054219 | 0.069319 | 0.088593 | 0.11394 | 0.14657 | 0.18909 | 0.24251 | 0.30812 | 0.38167 | 0.45767 | 0.52576 | 0.57299 | 0.59132 | 0.57824 | 0.5397 | 0.48481 | 0.42344 | 0.36292 | 0.3071 | 0.25963 | 0.22024 | 0.18846 | 0.16573 | 0.14947 | 0.13846 | 0.13032 | 0.13113 | 0.1334 | 0.1422 | 0.15348 | 0.16641 | 0.18887 | 0.2096 | 0.23238 | 0.25699 | 0.28353 | 0.3063 | 0.32649
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 4.4 h | 0.003361 | 0.0040465 | 0.0025946 | 0.0026214 | 0.0044553 | 0.0026323 | 0.0029087 | 0.0037366 | 0.0047149 | 0.0057564 | 0.0073092 | 0.0089323 | 0.011077 | 0.013703 | 0.017072 | 0.021236 | 0.026498 | 0.03301 | 0.041177 | 0.051293 | 0.064055 | 0.079683 | 0.099225 | 0.12384 | 0.15437 | 0.19231 | 0.23844 | 0.29224 | 0.35137 | 0.41125 | 0.46468 | 0.50437 | 0.52462 | 0.52356 | 0.50251 | 0.46708 | 0.42347 | 0.37253 | 0.32299 | 0.27537 | 0.23266 | 0.1978 | 0.16633 | 0.14424 | 0.12252 | 0.10889 | 0.097697 | 0.093226 | 0.09166 | 0.092155 | 0.093069
Time 4.5 h | 0.0034249 | 0.0041377 | 0.0027283 | 0.0027814 | 0.0046887 | 0.0029259 | 0.0032733 | 0.0042363 | 0.0053358 | 0.0065296 | 0.0082746 | 0.010141 | 0.012601 | 0.015587 | 0.01947 | 0.024233 | 0.030264 | 0.037737 | 0.047171 | 0.058901 | 0.073903 | 0.092482 | 0.11581 | 0.14538 | 0.18208 | 0.2274 | 0.28124 | 0.3416 | 0.40418 | 0.46243 | 0.50811 | 0.53414 | 0.53731 | 0.51809 | 0.48197 | 0.43577 | 0.38293 | 0.3311 | 0.28177 | 0.23748 | 0.20077 | 0.16802 | 0.14416 | 0.11848 | 0.11171 | 0.10187 | 0.09506 | 0.092791 | 0.094429 | 0.097904 | 0.10371
Time 4.6 h | 0.0036091 | 0.0043103 | 0.0029276 | 0.0030069 | 0.0049863 | 0.0032919 | 0.0037411 | 0.0048548 | 0.006126 | 0.0075084 | 0.0095269 | 0.011725 | 0.014579 | 0.018062 | 0.02259 | 0.028182 | 0.035264 | 0.044078 | 0.055262 | 0.069282 | 0.08729 | 0.10998 | 0.13862 | 0.17503 | 0.21981 | 0.27381 | 0.33549 | 0.40061 | 0.46249 | 0.51282 | 0.5432 | 0.54939 | 0.53148 | 0.49439 | 0.44517 | 0.39079 | 0.33577 | 0.28667 | 0.23998 | 0.20372 | 0.17181 | 0.14592 | 0.12721 | 0.11818 | 0.10489 | 0.098824 | 0.097908 | 0.099283 | 0.10394 | 0.11202 | 0.12205
Time 4.7 h | 0.0040254 | 0.0023197 | 0.0060162 | 0.0032852 | 0.0053707 | 0.003773 | 0.0043423 | 0.005672 | 0.0071637 | 0.0088006 | 0.011181 | 0.013771 | 0.017191 | 0.021329 | 0.026752 | 0.033478 | 0.04203 | 0.052722 | 0.066401 | 0.083738 | 0.10613 | 0.13436 | 0.17037 | 0.21573 | 0.27035 | 0.33347 | 0.40107 | 0.46603 | 0.51937 | 0.55253 | 0.55996 | 0.54171 | 0.50311 | 0.45097 | 0.39455 | 0.33876 | 0.28887 | 0.24104 | 0.20387 | 0.17339 | 0.14886 | 0.13225 | 0.11753 | 0.11 | 0.10482 | 0.10438 | 0.1074 | 0.1138 | 0.12408 | 0.13639 | 0.14993
Time 4.8 h | 0.0039527 | 0.0047899 | 0.0034912 | 0.0036734 | 0.0059298 | 0.004463 | 0.0052209 | 0.0068624 | 0.0086692 | 0.010687 | 0.013626 | 0.016821 | 0.021041 | 0.026213 | 0.03302 | 0.041432 | 0.052261 | 0.065911 | 0.083539 | 0.10606 | 0.13545 | 0.17252 | 0.21928 | 0.27665 | 0.34237 | 0.41271 | 0.47977 | 0.53353 | 0.56486 | 0.56874 | 0.5461 | 0.50317 | 0.44813 | 0.38872 | 0.33224 | 0.28045 | 0.23905 | 0.20172 | 0.17254 | 0.14998 | 0.13098 | 0.12232 | 0.11386 | 0.11155 | 0.11412 | 0.11812 | 0.12658 | 0.13848 | 0.15361 | 0.17099 | 0.18866
Time 4.9 h | 0.0043188 | 0.0052277 | 0.0039868 | 0.0042288 | 0.0066529 | 0.0053479 | 0.0063485 | 0.0083796 | 0.010604 | 0.013107 | 0.016765 | 0.020805 | 0.02609 | 0.032614 | 0.041224 | 0.052023 | 0.065957 | 0.083695 | 0.10683 | 0.13651 | 0.17523 | 0.22353 | 0.28261 | 0.35128 | 0.42376 | 0.49222 | 0.54583 | 0.5749 | 0.57491 | 0.54814 | 0.50162 | 0.4443 | 0.38384 | 0.32634 | 0.27564 | 0.23272 | 0.19908 | 0.1717 | 0.1495 | 0.13558 | 0.12502 | 0.11938 | 0.11838 | 0.12437 | 0.13062 | 0.13992 | 0.15492 | 0.1723 | 0.19073 | 0.21306 | 0.23377
Time 5 h | 0.0049133 | 0.0058042 | 0.0046338 | 0.0049712 | 0.0076664 | 0.0066031 | 0.0079479 | 0.010533 | 0.013358 | 0.01657 | 0.021298 | 0.026516 | 0.033397 | 0.041936 | 0.053268 | 0.067599 | 0.086277 | 0.11024 | 0.14161 | 0.18183 | 0.2334 | 0.29549 | 0.36662 | 0.44167 | 0.50969 | 0.5605 | 0.58424 | 0.57776 | 0.54458 | 0.49349 | 0.43374 | 0.3731 | 0.31662 | 0.26729 | 0.22557 | 0.19336 | 0.16657 | 0.14984 | 0.13717 | 0.1291 | 0.12498 | 0.12759 | 0.13258 | 0.1444 | 0.15601 | 0.17194 | 0.19151 | 0.21398 | 0.23437 | 0.26066 | 0.28203
Time 5.1 h | 0.0051615 | 0.0065394 | 0.0054847 | 0.0059393 | 0.0090001 | 0.0082333 | 0.010001 | 0.013324 | 0.016914 | 0.021051 | 0.027147 | 0.03396 | 0.042946 | 0.054219 | 0.069319 | 0.088593 | 0.11394 | 0.14657 | 0.18909 | 0.24251 | 0.30812 | 0.38167 | 0.45767 | 0.52576 | 0.57299 | 0.59132 | 0.57824 | 0.5397 | 0.48481 | 0.42344 | 0.36292 | 0.3071 | 0.25963 | 0.22024 | 0.18846 | 0.16573 | 0.14947 | 0.13846 | 0.13032 | 0.13113 | 0.1334 | 0.1422 | 0.15348 | 0.16641 | 0.18887 | 0.2096 | 0.23238 | 0.25699 | 0.28353 | 0.3063 | 0.32649
-38 oC
-18 oC
Monitoring Phase Behaviour (ice nucleation temperature and solidification end points by using
temperature calibration for predicting temperature of the product in primary drying
Surrogate drying rate (from )
  is highly sensitive to low ice volumes; therefore it could be used for determination end point of primary drying
Notes: Eg temperature cycling below to calibrate Fpeak for temperature (in order to predict product temperatures in primary drying) and the compensate cpeak for temperature so that it can used to monitor ice content and hence drying rate
Eg 2 left graph: monitoring phase behaviour (ice nucleation temperatures and solidufication end points)

## Slide 20: TVIS Response Surface (3D-Plot)
Imaginary Part of Capacitance
Real Part of Capacitance
High frequency
Liquid state
Liquid state
Frozen solid
Frozen solid
low frequency
Annealing = Re-heating and Re-cooling
Re-heating
Re-heating
Intermediate frequency
Re-cooling
Re-cooling
low frequency
Primary drying
Primary drying
low frequency
Notes: Full Shelf Water Study_010716_Ch1 (Filling volume = 2 mL) touch base electrode

## Slide 21: Dielectric loss spectrum
Data analysing software (LyoView ™) identifies the peak frequency () and peak amplitude () in the imaginary part of the capacitance spectrum
Chart: Log Frequency
Chart: Log Frequency

## Slide 22
Through Vial Impedance Spectroscopy (TVIS)
Dual-electrode system and its applications
(Ice temperature, Drying rate and Heat transfer coefficient )

## Slide 23: Dual-electrode system
Electrode height ~10 mm
Electrode distance from base ~3 mm
Standard TVIS vial
(Single electrode system)
Electrode Dimension
10 x 19 mm
New feature of TVIS vial
(Dual electrode system)
Electrode Dimension
Top electrode (TE): 10 x 19 mm
Bottom Electrode (BE): 5 x 19 mm
Bottom electrode height ~10 mm
Electrode distance from base ~3 mm
Top electrode height ~15 mm
Gap between electrode ~3 mm
A dual electrode system comprises two pairs of copper electrode glued to the external surface of a Type I tubular glass vial.
This option is suitable for large volume samples, including those used for determination.
~3.4 g of water
~8 g of water

## Slide 24: Temperature Determination
: TVIS predicted temperature
from top electrode (TE)
: TVIS predicted temperature
from bottom electrode (BE)
Both and can be estimated by extrapolating from the temperatures predicted from the centers of top electrode () and bottom electrode ().
Dual electrode
3 mm gap
3 mm gap
5 x 19 mm
15 x 19 mm

## Slide 25: Aims & Objectives
Aims
To determine the heat transfer coefficient () by using a novel dual electrode TVIS approach
Temperature calibration of log of top electrode () and bottom electrode ()
I
Prediction ice temperatures for both electrodes during primary drying
II
Temperature calibration of
III
Compensation of during primary drying
IV
Calibration of for ice layer height
V
Estimation of ice layer height during primary drying
VI
Prediction ice temperatures at (i) sublimation interface () and (ii) vial’s base () including qualification TVIS technique ()
VII
Comparison of TVIS drying rate () with gravimetric method (weight loss)
VIII
Determination (i) the drying rate () and (ii) ice base temperature () during the steady state period
IX
Heat transfer coefficient ( calculation
X

## Slide 26: Objective
Temperature calibration of log of top electrode () and bottom electrode ()
I
Annealing the sample
In-line TVIS measurement
Identifying peak frequency () using LyoView ™ software
Identifying peak frequency (𝐹_𝑃𝐸𝐴𝐾) using LyoView ™ software
Predicting product temperature using calibration plot
Calibration plot (temperature vs Log )
Calibration plot (temperature vs Log 𝐹_𝑃𝐸𝐴𝐾)
Annealing the sample
In-line TVIS measurement
Predicting product temperature using calibration plot

## Slide 27: Objective
Chart: Time /h
Shelf Temp.
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 6.5 h | -0.011026 | -0.012215 | -0.027689 | -0.023597 | -0.0069787 | -0.0053348 | -0.0030174 | 0.0013325 | 0.0065751 | 0.0097044 | 0.013438 | 0.017118 | 0.021618 | 0.026952 | 0.033746 | 0.041929 | 0.052016 | 0.064186 | 0.079127 | 0.097201 | 0.11944 | 0.146 | 0.1783 | 0.2178 | 0.26517 | 0.32227 | 0.39079 | 0.47092 | 0.56228 | 0.66169 | 0.76175 | 0.85177 | 0.91802 | 0.94907 | 0.93932 | 0.8934 | 0.82052 | 0.72661 | 0.63632 | 0.54819 | 0.46021 | 0.40033 | 0.33842 | 0.3044 | 0.2558 | 0.23181 | 0.21255 | 0.20425 | 0.20259 | 0.20868 | 0.22018
Time 6.6 h | -0.010844 | -0.013094 | -0.028215 | -0.023415 | -0.0063289 | -0.0045493 | -0.0018557 | 0.0026925 | 0.0080769 | 0.011546 | 0.015785 | 0.020036 | 0.025265 | 0.031534 | 0.039471 | 0.049058 | 0.06099 | 0.075285 | 0.092893 | 0.11431 | 0.14093 | 0.1727 | 0.21171 | 0.25948 | 0.31676 | 0.38554 | 0.46659 | 0.55908 | 0.65975 | 0.76252 | 0.85551 | 0.92559 | 0.96052 | 0.95286 | 0.90669 | 0.83273 | 0.74011 | 0.65262 | 0.55298 | 0.4715 | 0.40258 | 0.34125 | 0.29102 | 0.26019 | 0.22652 | 0.21129 | 0.1945 | 0.20035 | 0.20414 | 0.21607 | 0.23162
Time 6.7 h | -0.010233 | -0.013175 | -0.028274 | -0.022743 | -0.0054069 | -0.0034808 | -0.00042223 | 0.0044144 | 0.01012 | 0.014066 | 0.018998 | 0.024035 | 0.030298 | 0.037789 | 0.047292 | 0.058826 | 0.073223 | 0.090722 | 0.11222 | 0.13839 | 0.17098 | 0.21029 | 0.25838 | 0.3172 | 0.3871 | 0.46963 | 0.56416 | 0.66707 | 0.7712 | 0.86582 | 0.93655 | 0.97094 | 0.96234 | 0.91434 | 0.83635 | 0.74343 | 0.64608 | 0.55833 | 0.47266 | 0.39948 | 0.33451 | 0.29245 | 0.25353 | 0.23378 | 0.20871 | 0.19984 | 0.1997 | 0.2015 | 0.21198 | 0.22733 | 0.24895
Time 6.8 h | -0.0099355 | -0.010961 | -0.025677 | -0.02165 | -0.0043092 | -0.0020771 | 0.001197 | 0.0065657 | 0.012802 | 0.017476 | 0.023499 | 0.029632 | 0.037339 | 0.046572 | 0.058415 | 0.072736 | 0.09063 | 0.11238 | 0.13935 | 0.17222 | 0.21322 | 0.26261 | 0.32283 | 0.39563 | 0.48042 | 0.57735 | 0.68277 | 0.7887 | 0.88235 | 0.9508 | 0.98058 | 0.96664 | 0.91424 | 0.83246 | 0.73667 | 0.63856 | 0.5471 | 0.46601 | 0.38946 | 0.33579 | 0.28239 | 0.25179 | 0.22417 | 0.21528 | 0.19822 | 0.19607 | 0.20091 | 0.21064 | 0.22634 | 0.24719 | 0.27274
Time 6.9 h | -0.0089346 | -0.010265 | -0.024797 | -0.020617 | -0.0027907 | -0.00025292 | 0.0035689 | 0.0097638 | 0.016718 | 0.022291 | 0.029711 | 0.037389 | 0.047034 | 0.05871 | 0.073704 | 0.091917 | 0.11467 | 0.14237 | 0.17678 | 0.21881 | 0.27107 | 0.33354 | 0.40861 | 0.49749 | 0.59727 | 0.70497 | 0.81179 | 0.90383 | 0.96632 | 0.98864 | 0.96657 | 0.90704 | 0.82145 | 0.72363 | 0.62477 | 0.53332 | 0.45336 | 0.38378 | 0.32751 | 0.28107 | 0.24243 | 0.21966 | 0.20661 | 0.19414 | 0.1959 | 0.20098 | 0.21192 | 0.22772 | 0.24915 | 0.2732 | 0.30165
Time 7 h | -0.0079524 | -0.0096422 | -0.024013 | -0.019274 | -0.00065527 | 0.0023502 | 0.0069457 | 0.014294 | 0.022332 | 0.029183 | 0.0386 | 0.048527 | 0.061028 | 0.076199 | 0.095709 | 0.11952 | 0.14917 | 0.18538 | 0.23021 | 0.28467 | 0.35158 | 0.43028 | 0.52255 | 0.62693 | 0.73646 | 0.8422 | 0.92939 | 0.98297 | 0.99229 | 0.95868 | 0.8898 | 0.79968 | 0.70081 | 0.60339 | 0.51292 | 0.43557 | 0.37221 | 0.3158 | 0.27694 | 0.24492 | 0.21667 | 0.20627 | 0.19841 | 0.20582 | 0.20492 | 0.21652 | 0.23273 | 0.2534 | 0.27813 | 0.30477 | 0.33401
Time 7.1 h | -0.0072225 | -0.0078733 | -0.021826 | -0.017104 | 0.0022508 | 0.0058654 | 0.011358 | 0.020219 | 0.029818 | 0.038452 | 0.050584 | 0.063481 | 0.079875 | 0.09973 | 0.12537 | 0.15659 | 0.19542 | 0.24264 | 0.30066 | 0.37027 | 0.45418 | 0.54985 | 0.65682 | 0.7684 | 0.8716 | 0.95211 | 0.9947 | 0.99196 | 0.94565 | 0.86934 | 0.77503 | 0.67549 | 0.5806 | 0.49486 | 0.41965 | 0.35895 | 0.30957 | 0.26685 | 0.23671 | 0.2202 | 0.21086 | 0.19933 | 0.20391 | 0.21158 | 0.22344 | 0.23974 | 0.26069 | 0.28443 | 0.31084 | 0.33736 | 0.36523
Time 7.2 h | -0.0058046 | -0.0061413 | -0.019458 | -0.014478 | 0.0061524 | 0.010622 | 0.01737 | 0.028317 | 0.039983 | 0.051176 | 0.06699 | 0.083978 | 0.10566 | 0.13187 | 0.16569 | 0.20672 | 0.25752 | 0.3187 | 0.39276 | 0.47969 | 0.58074 | 0.68993 | 0.80068 | 0.9005 | 0.97199 | 1.0024 | 0.98623 | 0.92908 | 0.84405 | 0.74734 | 0.64904 | 0.55798 | 0.47566 | 0.40517 | 0.34692 | 0.30148 | 0.26468 | 0.23973 | 0.21784 | 0.21142 | 0.20723 | 0.20722 | 0.21754 | 0.22399 | 0.24796 | 0.26887 | 0.29214 | 0.31742 | 0.3425 | 0.36748 | 0.39208
-15.5 °C
Chart: Log FPEAK
| 2.7097859999999998 | 2.7409289999999999 | 2.7741730000000002 | 2.8071109999999999 | 2.8420100000000001 | 2.8772090000000001 | 2.9142049999999999 | 2.9513449999999999 | 2.9901089999999999 | 3.0288729999999999 | 3.0676369999999999 | 3.1087199999999999 | 3.1495929999999999 | 3.1926139999999998 | 3.2352349999999999 | 3.2774969999999999 | 3.3203369999999999 | 3.3627950000000002 | 3.4058160000000002 | 3.4484370000000002 | 3.4906990000000002 | 3.5335390000000002 | 3.57369 | 3.6122730000000001 | 3.6463709999999998 | 3.6786080000000001
Top electrode | -40.6 | -40 | -39.1 | -38.3 | -37.4 | -36.6 | -35.5 | -34.8 | -33.9 | -33.1 | -32 | -31 | -30.2 | -29 | -28 | -27 | -26 | -25 | -24.1 | -23.1 | -22.2 | -21.1 | -20.2 | -19.3 | -18.5 | -17.9
Bottom electrode | -40.1 | -39.4 | -38.6 | -37.8 | -36.8 | -35.9 | -35 | -34.1 | -33.2 | -32.3 | -31.2 | -30.3 | -29.3 | -28.4 | -27.3 | -26.2 | -25.3 | -24.2 | -23.4 | -22.3 | -21.3 | -20.4 | -19.4 | -18.5 | -17.8 | -17.2
Bottom electrode
Top electrode
Top electrode
TE | -1.02 | 30.1 | -114
BE | -5.43 x 10-1 | 26.7 | -109
TE | -1.02 | 30.1 | -114
BE | -5.43 x 10-1 | 26.7 | -109
Polynomial coefficient from calibration
Temperature calibration of log of top electrode () and bottom electrode ()
I

## Slide 28: Objective
Prediction ice temperatures for both electrodes during primary drying
II

## Slide 29: Objective
Chart: Time / h
Shelf Temperature
Chart: Time / h
Temp. constant
Temp. constant
Prediction ice temperatures for both electrodes during primary drying
II
The product temperature predicted by TVIS can demonstrate the temperature gradient across ice cylinder height

## Slide 30: Objective
Temperature calibration of
III
Annealing the sample
In-line TVIS measurement
Identifying peak amplitude () using LyoView ™ software
Identifying peak amplitude (𝐶_𝑃𝐸𝐴𝐾^″) using LyoView ™ software
Temperature compensation of using calibration plot
Temperature compensation of 𝐶_𝑃𝐸𝐴𝐾^″ using calibration plot
Calibration plot ( vs temperature)
Calibration plot (𝐶_𝑃𝐸𝐴𝐾^″ vs temperature)
Annealing the sample
In-line TVIS measurement

## Slide 31: Objective
Chart: Time /h
Shelf Temp.
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 6.5 h | -0.011026 | -0.012215 | -0.027689 | -0.023597 | -0.0069787 | -0.0053348 | -0.0030174 | 0.0013325 | 0.0065751 | 0.0097044 | 0.013438 | 0.017118 | 0.021618 | 0.026952 | 0.033746 | 0.041929 | 0.052016 | 0.064186 | 0.079127 | 0.097201 | 0.11944 | 0.146 | 0.1783 | 0.2178 | 0.26517 | 0.32227 | 0.39079 | 0.47092 | 0.56228 | 0.66169 | 0.76175 | 0.85177 | 0.91802 | 0.94907 | 0.93932 | 0.8934 | 0.82052 | 0.72661 | 0.63632 | 0.54819 | 0.46021 | 0.40033 | 0.33842 | 0.3044 | 0.2558 | 0.23181 | 0.21255 | 0.20425 | 0.20259 | 0.20868 | 0.22018
Time 6.6 h | -0.010844 | -0.013094 | -0.028215 | -0.023415 | -0.0063289 | -0.0045493 | -0.0018557 | 0.0026925 | 0.0080769 | 0.011546 | 0.015785 | 0.020036 | 0.025265 | 0.031534 | 0.039471 | 0.049058 | 0.06099 | 0.075285 | 0.092893 | 0.11431 | 0.14093 | 0.1727 | 0.21171 | 0.25948 | 0.31676 | 0.38554 | 0.46659 | 0.55908 | 0.65975 | 0.76252 | 0.85551 | 0.92559 | 0.96052 | 0.95286 | 0.90669 | 0.83273 | 0.74011 | 0.65262 | 0.55298 | 0.4715 | 0.40258 | 0.34125 | 0.29102 | 0.26019 | 0.22652 | 0.21129 | 0.1945 | 0.20035 | 0.20414 | 0.21607 | 0.23162
Time 6.7 h | -0.010233 | -0.013175 | -0.028274 | -0.022743 | -0.0054069 | -0.0034808 | -0.00042223 | 0.0044144 | 0.01012 | 0.014066 | 0.018998 | 0.024035 | 0.030298 | 0.037789 | 0.047292 | 0.058826 | 0.073223 | 0.090722 | 0.11222 | 0.13839 | 0.17098 | 0.21029 | 0.25838 | 0.3172 | 0.3871 | 0.46963 | 0.56416 | 0.66707 | 0.7712 | 0.86582 | 0.93655 | 0.97094 | 0.96234 | 0.91434 | 0.83635 | 0.74343 | 0.64608 | 0.55833 | 0.47266 | 0.39948 | 0.33451 | 0.29245 | 0.25353 | 0.23378 | 0.20871 | 0.19984 | 0.1997 | 0.2015 | 0.21198 | 0.22733 | 0.24895
Time 6.8 h | -0.0099355 | -0.010961 | -0.025677 | -0.02165 | -0.0043092 | -0.0020771 | 0.001197 | 0.0065657 | 0.012802 | 0.017476 | 0.023499 | 0.029632 | 0.037339 | 0.046572 | 0.058415 | 0.072736 | 0.09063 | 0.11238 | 0.13935 | 0.17222 | 0.21322 | 0.26261 | 0.32283 | 0.39563 | 0.48042 | 0.57735 | 0.68277 | 0.7887 | 0.88235 | 0.9508 | 0.98058 | 0.96664 | 0.91424 | 0.83246 | 0.73667 | 0.63856 | 0.5471 | 0.46601 | 0.38946 | 0.33579 | 0.28239 | 0.25179 | 0.22417 | 0.21528 | 0.19822 | 0.19607 | 0.20091 | 0.21064 | 0.22634 | 0.24719 | 0.27274
Time 6.9 h | -0.0089346 | -0.010265 | -0.024797 | -0.020617 | -0.0027907 | -0.00025292 | 0.0035689 | 0.0097638 | 0.016718 | 0.022291 | 0.029711 | 0.037389 | 0.047034 | 0.05871 | 0.073704 | 0.091917 | 0.11467 | 0.14237 | 0.17678 | 0.21881 | 0.27107 | 0.33354 | 0.40861 | 0.49749 | 0.59727 | 0.70497 | 0.81179 | 0.90383 | 0.96632 | 0.98864 | 0.96657 | 0.90704 | 0.82145 | 0.72363 | 0.62477 | 0.53332 | 0.45336 | 0.38378 | 0.32751 | 0.28107 | 0.24243 | 0.21966 | 0.20661 | 0.19414 | 0.1959 | 0.20098 | 0.21192 | 0.22772 | 0.24915 | 0.2732 | 0.30165
Time 7 h | -0.0079524 | -0.0096422 | -0.024013 | -0.019274 | -0.00065527 | 0.0023502 | 0.0069457 | 0.014294 | 0.022332 | 0.029183 | 0.0386 | 0.048527 | 0.061028 | 0.076199 | 0.095709 | 0.11952 | 0.14917 | 0.18538 | 0.23021 | 0.28467 | 0.35158 | 0.43028 | 0.52255 | 0.62693 | 0.73646 | 0.8422 | 0.92939 | 0.98297 | 0.99229 | 0.95868 | 0.8898 | 0.79968 | 0.70081 | 0.60339 | 0.51292 | 0.43557 | 0.37221 | 0.3158 | 0.27694 | 0.24492 | 0.21667 | 0.20627 | 0.19841 | 0.20582 | 0.20492 | 0.21652 | 0.23273 | 0.2534 | 0.27813 | 0.30477 | 0.33401
Time 7.1 h | -0.0072225 | -0.0078733 | -0.021826 | -0.017104 | 0.0022508 | 0.0058654 | 0.011358 | 0.020219 | 0.029818 | 0.038452 | 0.050584 | 0.063481 | 0.079875 | 0.09973 | 0.12537 | 0.15659 | 0.19542 | 0.24264 | 0.30066 | 0.37027 | 0.45418 | 0.54985 | 0.65682 | 0.7684 | 0.8716 | 0.95211 | 0.9947 | 0.99196 | 0.94565 | 0.86934 | 0.77503 | 0.67549 | 0.5806 | 0.49486 | 0.41965 | 0.35895 | 0.30957 | 0.26685 | 0.23671 | 0.2202 | 0.21086 | 0.19933 | 0.20391 | 0.21158 | 0.22344 | 0.23974 | 0.26069 | 0.28443 | 0.31084 | 0.33736 | 0.36523
Time 7.2 h | -0.0058046 | -0.0061413 | -0.019458 | -0.014478 | 0.0061524 | 0.010622 | 0.01737 | 0.028317 | 0.039983 | 0.051176 | 0.06699 | 0.083978 | 0.10566 | 0.13187 | 0.16569 | 0.20672 | 0.25752 | 0.3187 | 0.39276 | 0.47969 | 0.58074 | 0.68993 | 0.80068 | 0.9005 | 0.97199 | 1.0024 | 0.98623 | 0.92908 | 0.84405 | 0.74734 | 0.64904 | 0.55798 | 0.47566 | 0.40517 | 0.34692 | 0.30148 | 0.26468 | 0.23973 | 0.21784 | 0.21142 | 0.20723 | 0.20722 | 0.21754 | 0.22399 | 0.24796 | 0.26887 | 0.29214 | 0.31742 | 0.3425 | 0.36748 | 0.39208
-15.5 °C
Top electrode
Chart: Temperature /oC
| -40 | -39.1 | -38.299999999999997 | -37.4 | -36.6 | -35.5 | -34.799999999999997 | -33.9 | -33.1 | -32 | -31 | -30.2 | -29 | -28 | -27 | -26 | -25 | -24.1 | -23.1 | -22.2 | -21.1 | -20.2 | -19.3 | -18.5 | -17.899999999999999 | -17.100000000000001 | -16.600000000000001 | -16 | -15.5 | -15.2 | -14.7 | -14.5
Re-heating | 0.94718 | 0.95071 | 0.95411 | 0.95727 | 0.96001 | 0.96248 | 0.96474 | 0.96681 | 0.9687 | 0.97069 | 0.97255 | 0.97418 | 0.97558 | 0.97677 | 0.97768 | 0.97839 | 0.97899 | 0.97945 | 0.97971 | 0.97976 | 0.97957 | 0.97914 | 0.97846 | 0.97766 | 0.97681 | 0.97596 | 0.97513 | 0.97408 | 0.97303 | 0.97213 | 0.97127 | 0.97048
-1.00 x 10-4 | -5.20 x 10-3 | 9.19 x 10-1
-1.00 x 10-4 | -5.20 x 10-3 | 9.19 x 10-1
Polynomial coefficient from calibration
-14.5 °C
-40 °C
Top electrode
Temperature calibration of
III

## Slide 32: Objective
Compensation of during primary drying
IV

## Slide 33: Objective
During primary drying, is attributed to both the loss of ice and product temperature; therefore, it requires a standardization factor () for temperature compensation:
The expression for can be re-written in terms of the polynomial coefficients (slide 22):
Values of during primary drying are then standardized to the reference temperature by dividing by to give a standardized peak amplitude of
and are the peak amplitudes at temperatures () and reference temperature () during the re-heating ramp. In this presentation, a temperature of -20 °C is used as the reference temperature value
Chart: Time / h
The standardized is defined as
Chart: Time / h
Compensation of during primary drying
IV

## Slide 34: Objective
Calibration of for ice layer height
V
Filling water into TVIS vial
Freezing the sample
In-line TVIS measurement
Calibration plot ( vs temperature)
Calibration plot (𝐶_𝑃𝐸𝐴𝐾^″ vs temperature)
Thawing the sample
Identifying using simple peak finding
Identifying 𝐶_𝑃𝐸𝐴𝐾^″ using simple peak finding
Filling water into TVIS vial
Freezing the sample
In-line TVIS measurement
Thawing the sample

## Slide 35: Objective
Calibration of for ice layer height
V
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
0 mm | 0.0013746 | 0.00016925 | -0.010913 | -0.01186 | -0.0023631 | -0.0014374 | -0.0010379 | 7.0431e-05 | 0.00087291 | 0.00194 | 0.0032132 | 0.0043643 | 0.00576 | 0.0072677 | 0.0090276 | 0.010941 | 0.013034 | 0.015075 | 0.017002 | 0.018574 | 0.019679 | 0.020215 | 0.020242 | 0.019765 | 0.018906 | 0.017718 | 0.016447 | 0.015119 | 0.013913 | 0.012933 | 0.011994 | 0.011229 | 0.01049 | 0.0096809 | 0.0090548 | 0.0087375 | 0.00867 | 0.0083255 | 0.008589 | 0.0084126 | 0.0081989 | 0.0089905 | 0.0096046 | 0.0087964 | 0.011076 | 0.012465 | 0.013829 | 0.015592 | 0.017693 | 0.020004 | 0.022882
1 mm | 0.0027725 | 0.0018515 | -0.0092505 | -0.0099074 | 0.00073702 | 0.0022699 | 0.0034919 | 0.004273 | 0.0077695 | 0.0099906 | 0.01276 | 0.015324 | 0.018246 | 0.021403 | 0.025051 | 0.028978 | 0.033133 | 0.037091 | 0.04047 | 0.04265 | 0.043373 | 0.042565 | 0.040531 | 0.037669 | 0.034467 | 0.031215 | 0.028188 | 0.025539 | 0.023299 | 0.021591 | 0.020147 | 0.019038 | 0.018056 | 0.01722 | 0.016663 | 0.01663 | 0.016575 | 0.016995 | 0.017358 | 0.018559 | 0.019496 | 0.021177 | 0.022859 | 0.027278 | 0.028264 | 0.03148 | 0.03562 | 0.040533 | 0.046256 | 0.05258 | 0.060127
2 mm | 0.0025556 | 0.0042062 | -0.0069189 | -0.0070531 | 0.0050612 | 0.0074392 | 0.0097845 | 0.012275 | 0.017396 | 0.021327 | 0.026469 | 0.031628 | 0.037888 | 0.045042 | 0.053699 | 0.063205 | 0.073305 | 0.082842 | 0.090906 | 0.096188 | 0.098255 | 0.096899 | 0.0928 | 0.086718 | 0.079692 | 0.07254 | 0.065856 | 0.059996 | 0.055189 | 0.051295 | 0.048301 | 0.045995 | 0.04407 | 0.042496 | 0.041423 | 0.04147 | 0.042299 | 0.043522 | 0.045149 | 0.048416 | 0.051491 | 0.056861 | 0.062123 | 0.069787 | 0.075482 | 0.083507 | 0.092027 | 0.1012 | 0.11189 | 0.12216 | 0.13382
4 mm | 0.0098217 | 0.012857 | 0.0029162 | 0.0044657 | 0.021923 | 0.027475 | 0.03396 | 0.04256 | 0.053595 | 0.063826 | 0.077317 | 0.090942 | 0.1071 | 0.12594 | 0.14869 | 0.17399 | 0.20175 | 0.22964 | 0.25581 | 0.27689 | 0.29031 | 0.29404 | 0.28863 | 0.27556 | 0.25765 | 0.23771 | 0.21783 | 0.19948 | 0.18328 | 0.16931 | 0.15713 | 0.14631 | 0.13552 | 0.12507 | 0.11537 | 0.10625 | 0.097516 | 0.09077 | 0.085436 | 0.080247 | 0.074581 | 0.074997 | 0.072017 | 0.074699 | 0.073925 | 0.076692 | 0.080183 | 0.084679 | 0.090239 | 0.0972 | 0.1056
5 mm | 0.012924 | 0.016644 | 0.0043205 | 0.0094563 | 0.029609 | 0.036832 | 0.043438 | 0.058294 | 0.073283 | 0.087891 | 0.10754 | 0.12734 | 0.15165 | 0.17954 | 0.21314 | 0.25045 | 0.29123 | 0.33235 | 0.37121 | 0.40313 | 0.42452 | 0.43239 | 0.42664 | 0.40904 | 0.38343 | 0.35365 | 0.3231 | 0.29438 | 0.26853 | 0.24642 | 0.22701 | 0.20991 | 0.19402 | 0.17831 | 0.16442 | 0.15249 | 0.14204 | 0.13185 | 0.12652 | 0.1191 | 0.11251 | 0.11106 | 0.10999 | 0.10922 | 0.11194 | 0.11425 | 0.11948 | 0.12542 | 0.13331 | 0.14182 | 0.15289
6 mm | 0.016613 | 0.020933 | 0.0091737 | 0.015165 | 0.038068 | 0.04686 | 0.055308 | 0.072766 | 0.089943 | 0.10644 | 0.12816 | 0.14963 | 0.17566 | 0.20549 | 0.24169 | 0.28233 | 0.32749 | 0.37416 | 0.41944 | 0.45806 | 0.48528 | 0.49695 | 0.49191 | 0.47172 | 0.44084 | 0.40403 | 0.36577 | 0.32961 | 0.2979 | 0.27153 | 0.24965 | 0.23216 | 0.21743 | 0.20358 | 0.19134 | 0.18108 | 0.16982 | 0.15981 | 0.15095 | 0.14309 | 0.13722 | 0.12983 | 0.12612 | 0.12392 | 0.12523 | 0.12723 | 0.13111 | 0.13663 | 0.14387 | 0.15196 | 0.16262
7 mm | 0.019664 | 0.024611 | 0.012967 | 0.019785 | 0.045059 | 0.055353 | 0.065722 | 0.086411 | 0.1067 | 0.1266 | 0.15232 | 0.17875 | 0.21031 | 0.24651 | 0.29062 | 0.34037 | 0.39602 | 0.45396 | 0.51107 | 0.56056 | 0.59714 | 0.61478 | 0.61172 | 0.58909 | 0.55206 | 0.50663 | 0.45873 | 0.41313 | 0.37261 | 0.33899 | 0.31092 | 0.28831 | 0.26853 | 0.25064 | 0.23364 | 0.21847 | 0.20334 | 0.18982 | 0.17493 | 0.16472 | 0.15562 | 0.1442 | 0.14214 | 0.13187 | 0.13682 | 0.13729 | 0.14081 | 0.14603 | 0.15301 | 0.16257 | 0.17436
8 mm | 0.022191 | 0.024302 | 0.015888 | 0.023391 | 0.050987 | 0.062633 | 0.0751 | 0.099163 | 0.1228 | 0.14637 | 0.17727 | 0.20917 | 0.24729 | 0.29108 | 0.34437 | 0.40439 | 0.47128 | 0.5404 | 0.60767 | 0.66462 | 0.70504 | 0.72217 | 0.71465 | 0.68435 | 0.63803 | 0.58285 | 0.52571 | 0.47208 | 0.42452 | 0.38537 | 0.35336 | 0.32731 | 0.30445 | 0.28392 | 0.26444 | 0.24623 | 0.23159 | 0.21324 | 0.198 | 0.18347 | 0.16657 | 0.16468 | 0.15801 | 0.16092 | 0.15041 | 0.15193 | 0.15353 | 0.1576 | 0.16308 | 0.17249 | 0.18304
9 mm | 0.025178 | 0.02843 | 0.020221 | 0.028243 | 0.058241 | 0.071385 | 0.085782 | 0.11285 | 0.13953 | 0.1659 | 0.20185 | 0.2382 | 0.28153 | 0.331 | 0.3909 | 0.45815 | 0.53273 | 0.60911 | 0.68318 | 0.74564 | 0.78956 | 0.80825 | 0.80044 | 0.76808 | 0.71858 | 0.65905 | 0.59628 | 0.53734 | 0.48422 | 0.44063 | 0.40484 | 0.37572 | 0.35241 | 0.33355 | 0.316 | 0.3018 | 0.287 | 0.27231 | 0.25696 | 0.23941 | 0.22479 | 0.21074 | 0.20058 | 0.19055 | 0.18084 | 0.17458 | 0.17198 | 0.17099 | 0.17232 | 0.17616 | 0.18292
10 mm | 0.02931 | 0.032954 | 0.025296 | 0.034295 | 0.06725 | 0.078945 | 0.09936 | 0.13039 | 0.16103 | 0.19132 | 0.23267 | 0.27445 | 0.32415 | 0.38106 | 0.44994 | 0.52754 | 0.61349 | 0.70185 | 0.78726 | 0.85846 | 0.90721 | 0.92497 | 0.91033 | 0.86657 | 0.80291 | 0.72859 | 0.65233 | 0.5809 | 0.51827 | 0.46644 | 0.42508 | 0.39319 | 0.36784 | 0.34779 | 0.33105 | 0.31867 | 0.30683 | 0.29229 | 0.27976 | 0.26514 | 0.24949 | 0.23653 | 0.23008 | 0.23378 | 0.21063 | 0.2058 | 0.20258 | 0.20255 | 0.20419 | 0.20957 | 0.21751
11 mm | 0.030885 | 0.033883 | 0.025794 | 0.035193 | 0.069015 | 0.081061 | 0.10236 | 0.13477 | 0.16686 | 0.19878 | 0.24247 | 0.28687 | 0.3397 | 0.40029 | 0.47401 | 0.55745 | 0.65108 | 0.74884 | 0.84577 | 0.92979 | 0.99174 | 1.0211 | 1.015 | 0.97487 | 0.9107 | 0.83062 | 0.74636 | 0.66785 | 0.59703 | 0.53979 | 0.48998 | 0.45281 | 0.42298 | 0.39929 | 0.37893 | 0.36513 | 0.35333 | 0.34194 | 0.3318 | 0.31863 | 0.3062 | 0.29458 | 0.28223 | 0.28972 | 0.26381 | 0.25593 | 0.24876 | 0.2448 | 0.24301 | 0.2428 | 0.24608
12 mm | 0.032482 | 0.036695 | 0.028707 | 0.038445 | 0.073987 | 0.087194 | 0.10992 | 0.14456 | 0.17825 | 0.21335 | 0.26054 | 0.30844 | 0.36552 | 0.43109 | 0.51129 | 0.60208 | 0.70471 | 0.81264 | 0.92018 | 1.0145 | 1.0847 | 1.1186 | 1.1125 | 1.0687 | 0.99632 | 0.90626 | 0.81153 | 0.72126 | 0.64 | 0.57268 | 0.51614 | 0.47171 | 0.43528 | 0.40514 | 0.38388 | 0.36905 | 0.35669 | 0.3488 | 0.34184 | 0.33695 | 0.32609 | 0.32647 | 0.31616 | 0.32781 | 0.30771 | 0.30488 | 0.29988 | 0.29784 | 0.29674 | 0.29804 | 0.30238
Chart: -C″PEAK / pF
| 2.0282321658632974E-2 | 4.3255845536608106E-2 | 9.8028305551333794E-2 | #N/A | 0.29349694651821778 | 0.43151334859090174 | 0.49610799625102509 | 0.61477274370953339 | 0.72094798390835635 | 0.80705883860768424 | 0.92264736847743656 | 1.0206764847121299 | 1.118562503327631 | 1.1888173213507329 | 1.3470508093994198 | 1.4443671859293801 | #N/A | #N/A
 | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | #N/A | #N/A
Trendline | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15
0 mm
15 mm

## Slide 36: Objective
Estimation of ice layer height during primary drying
VI

## Slide 37: Objective
At 2.4 h into primary drying
= 0.863 pF
Ice height = 9.7485 x 0.863 + 1.0462
= 9.459 mm (from the bottom edge of TE)
Ice front height = 9.459+(2+5+3)
= 19.46 mm
Bottom electrode (BE)
5 x 19 mm
3 mm gap
Freeze Dryer Shelf
Chart: -C″PEAK / pF
| 2.0282321658632974E-2 | 4.3255845536608106E-2 | 9.8028305551333794E-2 | #N/A | 0.29349694651821778 | 0.43151334859090174 | 0.49610799625102509 | 0.61477274370953339 | 0.72094798390835635 | 0.80705883860768424 | 0.92264736847743656 | 1.0206764847121299 | 1.118562503327631 | 1.1888173213507329 | 1.3470508093994198 | 1.4443671859293801 | #N/A | #N/A
 | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | #N/A | #N/A
Trendline | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15
3 mm gap
Top electrode (TE)
15 X 19 mm
At -20 °C
Freeze drying shelf
2 mm
5 mm
3 mm
0.863 pF
19.46 mm
10 mm
9.46 mm
Gradient of the line ()
Estimation of ice layer height during primary drying
VI

## Slide 38: Objective
Surrogate drying rate can be estimated in terms of decreasing ice height
The dependency of on the ice cylinder height in linear region
Chart: Time / h
20.50 mm
Estimation of ice layer height during primary drying
VI

## Slide 39: Objective
Prediction ice temperatures at (i) sublimation interface () and (ii) vial’s base () including qualification TVIS technique ()
VII

## Slide 40: Objective
Prediction ice temperatures at (i) sublimation interface () and (ii) vial’s base () including qualification TVIS technique ()
VII
Freeze drying shelf
Chart: Temperature (𝑇) / oC
| -33.119466634243032 | -32.321208025650492 | -30.594741732648018 | -29.835299081767161
TE_BE_predict | 19.46 | 14.73 | 4.5 | 0
Chart: Time / h
-32.3 ˚C
Chart: Time / h
Ice height 19.46 mm at 2.4 h
-30.6 ˚C
At 2.4 h
2 mm
5 mm
3 mm
9.46 mm
0 mm
19.46 mm
-30.6 ˚C
-32.3 ˚C

## Slide 41: Objective
Prediction ice temperatures at (i) sublimation interface () and (ii) vial’s base () including qualification TVIS technique ()
VII
Freeze drying shelf
Chart: Temperature (𝑇) / oC
| -33.119466634243032 | -32.321208025650492 | -30.594741732648018 | -29.835299081767161
TE_BE_predict | 19.46 | 14.73 | 4.5 | 0
Chart: Time / h
4.5 mm
-32.3 ˚C
14.73 mm
Chart: Time / h
Ice height 19.46 mm at 2.4 h
-30.6 ˚C
At 2.4 h
Ice height for = () = 14.73 mm
Ice height for = () = 4.50 mm
2 mm
5 mm
3 mm
9.46 mm
0 mm
-30.6 ˚C
-32.3 ˚C

## Slide 42: Objective
Prediction ice temperatures at (i) sublimation interface () and (ii) vial’s base () including qualification TVIS technique ()
VII
Freeze drying shelf
Chart: Temperature (𝑇) / oC
| -33.119466634243032 | -32.321208025650492 | -30.594741732648018 | -29.835299081767161
TE_BE_predict | 19.46 | 14.73 | 4.5 | 0
Chart: Time / h
4.5 mm
-32.3 ˚C
14.73 mm
Chart: Time / h
Ice height 19.46 mm at 2.4 h
-30.6 ˚C
At 2.4 h
Ice height for = () = 14.73 mm
Ice height for = () = 4.50 mm
2 mm
5 mm
3 mm
9.46 mm
0 mm
19.46 mm
-30.6 ˚C
-32.3 ˚C

## Slide 43: Objective
Prediction ice temperatures at (i) sublimation interface () and (ii) vial’s base () including qualification TVIS technique ()
VII
Freeze drying shelf
Chart: Temperature (𝑇) / oC
| -33.119466634243032 | -32.321208025650492 | -30.594741732648018 | -29.835299081767161
TE_BE_predict | 19.46 | 14.73 | 4.5 | 0
Chart: Time / h
4.5 mm
-32.3 ˚C
14.73 mm
Chart: Time / h
Ice height 19.46 mm at 2.4 h
-30.6 ˚C
At 2.4 h
Ice height for = () = 14.73 mm
Ice height for = () = 4.50 mm
2 mm
5 mm
3 mm
9.46 mm
-33.1 ˚C
-29.8 ˚C
0 mm
19.46 mm
-30.6 ˚C
-32.3 ˚C
Ice Temperature
At interface (, 19.46 mm) = = -33.1 ˚C
At vial’s base (, 0 mm) = = -29.8 ˚C

## Slide 44: Objective
The product temperature at ice interface predicted by using a 2-points temperature extrapolation close to the temperature of ice vapour at chamber pressure of 270 µbar ()
Chart: Time / h
Shelf Temperature
Chart: Time / h
Temp. constant
(= -33.1± 0.05˚C )
Temp. constant
(= -33.1± 0.05˚C )
Prediction ice temperatures at (i) sublimation interface () and (ii) vial’s base () including qualification TVIS technique ()
VII

## Slide 45: Objective
Comparison of TVIS drying rate () with gravimetric method (weight loss)
VIII

## Slide 46: Objective
Drying rate is based on the assumption of a planar sublimation front
The change in ice cylinder height () can be equated to the change in ice volume ()
Where is internal radius of vial and is internal cross section area of vial ()
Ice volume can be converted to ice mass () by multiplying with ice density ()
Hence; drying rate () can be expressed by
Comparison of TVIS drying rate () with gravimetric method (weight loss)
VIII

## Slide 47: Objective
An average surrogate drying rate calculation
  Ice density at -32 ˚C = 0.920 gcm-3
  Internal vial diameter (VC010-20C) = 2.21 cm Cross-section area () = 3.80 cm2
  Ice height at 0 h () = 20.50 mm
  Ice height at 4.9 h ( = 16.68 mm
 | 0.27 g/h
 | 0.25 g/h
 | 0.27 g/h
 | 0.25 g/h
Chart: Time / h
20.50 mm
Comparison of TVIS drying rate () with gravimetric method (weight loss)
VIII

## Slide 48: Objective
Determination (i) the drying rate () and (ii) ice base temperature () during the steady state period
IX

## Slide 49: Objective
Chart: Time / h
Chart: Time / h
Determination (i) the drying rate () and (ii) ice base temperature () during the steady state period for heat transfer coefficient ( calculation
IX

## Slide 50: Objective
Chart: Time / h
Chart: Time / h
Temp. constant = -33.1± 0.05˚C
= -29.8± 0.03˚C
~32 ˚C
I
II
Determination (i) the drying rate () and (ii) ice base temperature () during the steady state period for heat transfer coefficient ( calculation
IX

## Slide 51: Objective
Chart: Time / h
Chart: Time / h
Temp. constant = -33.1± 0.05˚C
= -29.8± 0.03˚C
~32 ˚C
Drying rate during the steady state
Determination (i) the drying rate () and (ii) ice base temperature () during the steady state period for heat transfer coefficient ( calculation
IX

## Slide 52: Objective
Chart: Time / h
Chart: Time / h
Temp. constant = -33.1± 0.05˚C
= -29.8± 0.03˚C
~32 ˚C
Drying rate during the steady state
  Ice density at -32˚C = 0.920 gcm-3
  (Calculated ice temperature between
Determination (i) the drying rate () and (ii) ice base temperature () during the steady state period for heat transfer coefficient ( calculation
IX

## Slide 53: Objective
Chart: Time / h
Chart: Time / h
Temp. constant = -33.1± 0.05˚C
= -29.8± 0.03˚C
~31.5 ˚C
Drying rate during the steady state
  Ice density at -32˚C = 0.920 gcm-3
  (Calculated ice temperature between
  Internal vial diameter (VC010-20C) = 2.21 cm
Determination (i) the drying rate () and (ii) ice base temperature () during the steady state period for heat transfer coefficient ( calculation
IX

## Slide 54: Objective
Chart: Time / h
Chart: Time / h
Temp. constant = -33.1± 0.05˚C
= -29.8± 0.03˚C
~31.5 ˚C
Drying rate during the steady state
  Ice density at -32˚C = 0.920 gcm-3
  (Calculated ice temperature between
  Internal vial diameter (VC010-20C) = 2.21 cm Cross-section area () = 3.80 cm2
Determination (i) the drying rate () and (ii) ice base temperature () during the steady state period for heat transfer coefficient ( calculation
IX

## Slide 55: Objective
Chart: Time / h
Chart: Time / h
Temp. constant = -33.1± 0.05˚C
= -29.8± 0.03˚C
19.94 mm
~31.5 ˚C
Drying rate during the steady state
  Ice density at -32˚C = 0.920 gcm-3
  (Calculated ice temperature between
  Internal vial diameter (VC010-20C) = 2.21 cm Cross-section area () = 3.80 cm2
  Ice height at 2 h () = 19.94 mm
Determination (i) the drying rate () and (ii) ice base temperature () during the steady state period for heat transfer coefficient ( calculation
IX

## Slide 56: Objective
Chart: Time / h
Chart: Time / h
Temp. constant = -33.1± 0.05˚C
= -29.8± 0.03˚C
19.94 mm
~31.5 ˚C
Drying rate during the steady state
  Ice density at -32˚C = 0.920 gcm-3
  (Calculated ice temperature between
  Internal vial diameter (VC010-20C) = 2.21 cm Cross-section area () = 3.80 cm2
  Ice height at 2 h () = 19.94 mm
  Ice height at 2.8 h ( = 18.98 mm
Determination (i) the drying rate () and (ii) ice base temperature () during the steady state period for heat transfer coefficient ( calculation
IX

## Slide 57: Objective
Chart: Time / h
Chart: Time / h
Temp. constant = -33.1± 0.05˚C
= -29.8± 0.03˚C
19.94 mm
~31.5 ˚C
Drying rate during the steady state
  Ice density at -32˚C = 0.920 gcm-3
  (Calculated ice temperature between
  Internal vial diameter (VC010-20C) = 2.21 cm Cross-section area () = 3.80 cm2
  Ice height at 2 h () = 19.94 mm
  Ice height at 2.8 h ( = 18.98 mm
Determination (i) the drying rate () and (ii) ice base temperature () during the steady state period for heat transfer coefficient ( calculation
IX

## Slide 58: Objective
Chart: Time / h
Chart: Time / h
Temp. constant = -33.1± 0.05˚C
= -29.8± 0.03˚C
19.94 mm
Drying rate during the steady state
  Ice density at -32˚C = 0.920 gcm-3
  (Calculated ice temperature between
  Internal vial diameter (VC010-20C) = 2.21 cm Cross-section area () = 3.80 cm2
  Ice height at 2 h () = 19.94 mm
  Ice height at 2.8 h ( = 18.98 mm
~32 ˚C
TVIS parameters used for determination:
= 0.42 g·h-1
= -29.8˚C
Determination (i) the drying rate () and (ii) ice base temperature () during the steady state period for heat transfer coefficient ( calculation
IX

## Slide 59: Objective
Heat transfer coefficient ( calculation
X

## Slide 60: Objective
Parameters | TVIS
Drying rate at steady state (g/h) (2-2.8 h into primary drying) | 0.42
Shelf Temperature, (K) | 273.3
Vial’s base Temperature, (K) | 243.3
Parameters | TVIS
Drying rate at steady state (g/h) (2-2.8 h into primary drying) | 0.42
 | 273.3
 | 243.3
Heat transfer coefficient ( calculation
X
Notes: Freezing -40 C
Annealing -10 C
Drying -20 C

## Slide 61: Objective
Parameters | TVIS
Drying rate at steady state (g/h) (2-2.8 h into primary drying) | 0.42
Shelf Temperature, (K) | 273.3
Vial’s base Temperature, (K) | 243.3
Parameters | TVIS
Drying rate at steady state (g/h) (2-2.8 h into primary drying) | 0.42
 | 273.3
 | 243.3
Heat transfer coefficient ( calculation
X
is the latent heat of sublimation of ice (2844 Jg-1 or 679.7 cal g-1) and is external cross-sectional area of the base of the TVIS vial (4.62 cm2)
Notes: Freezing -40 C
Annealing -10 C
Drying -20 C

## Slide 62: Objective
Parameters | TVIS
Drying rate at steady state (g/h) (2-2.8 h into primary drying) | 0.42
Shelf Temperature, (K) | 273.3
Vial’s base Temperature, (K) | 243.3
Parameters | TVIS
Drying rate at steady state (g/h) (2-2.8 h into primary drying) | 0.42
 | 273.3
 | 243.3
Heat transfer coefficient ( calculation
X
is the latent heat of sublimation of ice (2844 Jg-1 or 679.7 cal g-1) and is external cross-sectional area of the base of the TVIS vial (4.62 cm2)
Notes: Freezing -40 C
Annealing -10 C
Drying -20 C

## Slide 63: Objective
Chart: Pressure (mTorr)
| 203.25
TVIS | 0.00057283
Tchessalov S (2017) | 0.00075778 | 0.00057449 | 0.00057449 | 0.0002011 | 0.00030003 | 0.00038954 | 0.0004709 | 0.00054519 | 0.00061327 | 0.00016575 | 0.0002352 | 0.00036064 | 0.00056849
Kuu W (2009) | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.000284 | 0.000372 | 0.000508 | 0.000726 | 0.000899
Brülls M (2002) | 0.00019353 | 0.00044029 | 0.0012289
TVIS @ 270 bar)
5.73 x 10-4
is the latent heat of sublimation of ice (2844 Jg-1 or 679.7 cal g-1) and is external cross-sectional area of the base of the TVIS vial (4.62 cm2)
Parameters | TVIS
Drying rate at steady state (g/h) (2-2.8 h into primary drying) | 0.42
Shelf Temperature, (K) | 273.3
Vial’s base Temperature, (K) | 243.3
Parameters | TVIS
Drying rate at steady state (g/h) (2-2.8 h into primary drying) | 0.42
 | 273.3
 | 243.3
Heat transfer coefficient ( calculation
X
Notes: Freezing -40 C
Annealing -10 C
Drying -20 C

## Slide 64: Additional comments
1) Calculation the relative heat transfer from the vial’s base
2) Qualification of steady state heat transfer mechanisms

## Slide 65
Chart: Time / h
Chart: Time / h
Parameters | TVIS
Ice interfacial temperature, (K) | 240.1
Vial’s base Temperature, (K) | 243.3
Ice height or thickness of material, (m) | 0.0195
Average temperature (between & , (˚C) | -31.5
(between & (K) | 3.3
Parameters | TVIS
 | 240.1
 | 243.3
 | 0.0195
 | -31.5
 | 3.3
Ice height 19.46 mm at 2.4 h
--29.8 ˚C
--33.1 ˚C

## Slide 66: Objective
Chart: Temperature / oC
| 0 | -5 | -10 | -15 | -20 | -25 | -30 | -35 | -40 | -50 | -60 | -70 | -80 | -90 | -100
Thermal conductivity value (W/mK) | 2.22 | 2.25 | 2.3 | 2.34 | 2.39 | 2.45 | 2.5 | 2.57 | 2.63 | 2.76 | 2.9 | 3.05 | 3.19 | 3.34 | 3.48
y = 4E-07x3 + 0.0001x2 - 0.0069x + 2.2174
Thermal conductivity of ice
of -31.5 ˚C = 2.52 W/mK
Or 2.52 J/smK
k is the thermal conductivity constant for the material
A is the cross sectional area of the material transferring heat (≡ internal cross-sectional area of vial) 0.00038 m2
ΔT is the difference in temperature between one side of the material and the other
is Latent heat of sublimation at 240 K (2844 J/g)
d is the thickness of the material (≡ ice height)
This drying rate is 50% of drying rate calculated by TVIS
Parameters | TVIS
Ice interfacial temperature, (K) | 240.1
Vial’s base Temperature, (K) | 243.3
Ice height or thickness of material, (m) | 0.0195
Average temperature (between & , (˚C) | -31.5
(between & (K) | 3.3
Parameters | TVIS
 | 240.1
 | 243.3
 | 0.0195
 | -31.5
 | 3.3
Notes: 1 Watt = 1 Joule/sec
Watts/(meter∗Kelvin)

## Slide 67: A single vial technique Pikal, et al. (1984)
The mean sublimation rate was calculated from the mass of ice sublimed and the time required for sublimation.
Notes: [1] M. Pikal, M. Roy, S. Shah, Mass and heat transfer in vial freeze‐drying of pharmaceuticals: Role of the vial, J. Pharm. Sci. 73 (1984) 1224-1237.
[2] S. Hibler, C. Wagner, H. Gieseler, Vial freeze-drying, part 1: New insights into heat transfer characteristics of tubing and molded vials, J. Pharm. Sci. 101 (2012) 1189-1201.
[3] B. Scutella, S. Passot, E. Bourles, F. Fonseca, I.C. Trelea, How Vial Geometry Variability Influences Heat Transfer and Product Temperature During Freeze-Drying, J. Pharm. Sci. 106 (2017) 770-778.

## Slide 68: A single vial technique Scutella, et al. 2017
Notes: [1] M. Pikal, M. Roy, S. Shah, Mass and heat transfer in vial freeze‐drying of pharmaceuticals: Role of the vial, J. Pharm. Sci. 73 (1984) 1224-1237.
[2] S. Hibler, C. Wagner, H. Gieseler, Vial freeze-drying, part 1: New insights into heat transfer characteristics of tubing and molded vials, J. Pharm. Sci. 101 (2012) 1189-1201.
[3] B. Scutella, S. Passot, E. Bourles, F. Fonseca, I.C. Trelea, How Vial Geometry Variability Influences Heat Transfer and Product Temperature During Freeze-Drying, J. Pharm. Sci. 106 (2017) 770-778.

## Slide 69: Assumption for determination
How do we know the heat transfer mechanisms constant up to 25% loss of ice mass?
If the heat transfer mechanisms change because of ice- glass interface contact or because of the change of ice shape (surface area), then surely heat transfer coefficient will change.
It requires the technique to qualify when the heat transfer mechanisms change

## Slide 70: Limitation of TVIS Systems
Chart: Log Frequency
| 1 | 1.0999998579423673 | 1.1999999472615537 | 1.2999999314429707 | 1.3999999253938222 | 1.5000000466711041 | 1.600000032123142 | 1.6999999708609113 | 1.7999999693838207 | 1.8999999810146946 | 2 | 2.0999998579423673 | 2.1999999472615537 | 2.2999999314429704 | 2.3999999253938222 | 2.5000000466711039 | 2.600000032123142 | 2.6999999708609113 | 2.7999999693838209 | 2.8999999810146946 | 3 | 3.0999998579423673 | 3.1999999472615537 | 3.2999999314429704 | 3.3999999253938222 | 3.5000000466711039 | 3.600000032123142 | 3.6999999708609113 | 3.7999999693838209 | 3.8999999810146946 | 4 | 4.0999998579423673 | 4.1999999472615537 | 4.2999999314429704 | 4.3999999253938222 | 4.5000000466711043 | 4.600000032123142 | 4.6999999708609108 | 4.7999999693838209 | 4.8999999810146946 | 5 | 5.0999998579423673 | 5.1999999472615537 | 5.2999999314429704 | 5.3999999253938222 | 5.5000000466711043 | 5.600000032123142 | 5.6999999708609108 | 5.7999999693838209 | 5.8999999810146946 | 6
Data | 0.14608 | 0.14492 | 0.14322 | 0.13877 | 0.13274 | 0.12435 | 0.11389 | 0.10356 | 0.093681 | 0.083851 | 0.074906 | 0.06705 | 0.060224 | 0.054763 | 0.050679 | 0.048268 | 0.047644 | 0.049094 | 0.052239 | 0.057957 | 0.066043 | 0.076623 | 0.08925 | 0.10432 | 0.1206 | 0.13645 | 0.14992 | 0.15793 | 0.15821 | 0.14965 | 0.13423 | 0.11423 | 0.093279 | 0.073862 | 0.057521 | 0.044451 | 0.034421 | 0.026766 | 0.02089 | 0.016448 | 0.013021 | 0.010225 | 0.0084856 | 0.0069896 | 0.0058967 | 0.0050832 | 0.0045142 | 0.0041497 | 0.0039574 | 0.0038715 | 0.0038344
Fit results CG=CPE-Cs=DE | 0.13513 | 0.14157 | 0.14388 | 0.14177 | 0.1356 | 0.12628 | 0.11498 | 0.10288 | 0.090935 | 0.079866 | 0.07012 | 0.061949 | 0.055477 | 0.050757 | 0.047819 | 0.046704 | 0.047484 | 0.05027 | 0.055219 | 0.06251 | 0.072302 | 0.084639 | 0.099291 | 0.11552 | 0.13183 | 0.14591 | 0.15502 | 0.15693 | 0.151 | 0.13852 | 0.12198 | 0.10405 | 0.086694 | 0.071061 | 0.057606 | 0.04636 | 0.037134 | 0.029653 | 0.023633 | 0.018813 | 0.014964 | 0.011896 | 0.0094544 | 0.0075125 | 0.0059687 | 0.0047417 | 0.0037668 | 0.0029922 | 0.0023769 | 0.0018881 | 0.0014998
simulation results from fitting CG=CPE-Cs=DE and remove CG=CPE | 0.0027685 | 0.0034854 | 0.0043877 | 0.0055237 | 0.0069536 | 0.0087536 | 0.011019 | 0.01387 | 0.017457 | 0.021968 | 0.027639 | 0.03476 | 0.043691 | 0.054866 | 0.068799 | 0.086074 | 0.1073 | 0.13303 | 0.16352 | 0.19844 | 0.23632 | 0.27403 | 0.30665 | 0.32823 | 0.33395 | 0.3224 | 0.29636 | 0.26134 | 0.2231 | 0.18598 | 0.15249 | 0.12364 | 0.099512 | 0.079713 | 0.063657 | 0.050736 | 0.040387 | 0.032124 | 0.025539 | 0.020297 | 0.016128 | 0.012814 | 0.01018 | 0.0080868 | 0.0064239 | 0.0051029 | 0.0040534 | 0.0032198 | 0.0025576 | 0.0020316 | 0.0016138
simulation results from fitting CG=CPE-Cs=DE but reduce CG 10% | 0.13971 | 0.14743 | 0.15098 | 0.14989 | 0.14437 | 0.13526 | 0.12373 | 0.11104 | 0.098278 | 0.086253 | 0.075494 | 0.066308 | 0.058841 | 0.053152 | 0.049263 | 0.047194 | 0.046989 | 0.048731 | 0.052545 | 0.058586 | 0.067001 | 0.077859 | 0.09101 | 0.10588 | 0.12124 | 0.13507 | 0.14481 | 0.14818 | 0.14416 | 0.13358 | 0.11862 | 0.10182 | 0.085229 | 0.07008 | 0.056931 | 0.04588 | 0.036781 | 0.029389 | 0.023431 | 0.018656 | 0.014841 | 0.0118 | 0.0093785 | 0.0074524 | 0.0059211 | 0.004704 | 0.0037369 | 0.0029685 | 0.002358 | 0.0018731 | 0.0014879
Chart: Log Frequency
| 1 | 1.0999998579423673 | 1.1999999472615537 | 1.2999999314429707 | 1.3999999253938222 | 1.5000000466711041 | 1.600000032123142 | 1.6999999708609113 | 1.7999999693838207 | 1.8999999810146946 | 2 | 2.0999998579423673 | 2.1999999472615537 | 2.2999999314429704 | 2.3999999253938222 | 2.5000000466711039 | 2.600000032123142 | 2.6999999708609113 | 2.7999999693838209 | 2.8999999810146946 | 3 | 3.0999998579423673 | 3.1999999472615537 | 3.2999999314429704 | 3.3999999253938222 | 3.5000000466711039 | 3.600000032123142 | 3.6999999708609113 | 3.7999999693838209 | 3.8999999810146946 | 4 | 4.0999998579423673 | 4.1999999472615537 | 4.2999999314429704 | 4.3999999253938222 | 4.5000000466711043 | 4.600000032123142 | 4.6999999708609108 | 4.7999999693838209 | 4.8999999810146946 | 5 | 5.0999998579423673 | 5.1999999472615537 | 5.2999999314429704 | 5.3999999253938222 | 5.5000000466711043 | 5.600000032123142 | 5.6999999708609108 | 5.7999999693838209 | 5.8999999810146946 | 6
Data | 0.14608 | 0.14492 | 0.14322 | 0.13877 | 0.13274 | 0.12435 | 0.11389 | 0.10356 | 0.093681 | 0.083851 | 0.074906 | 0.06705 | 0.060224 | 0.054763 | 0.050679 | 0.048268 | 0.047644 | 0.049094 | 0.052239 | 0.057957 | 0.066043 | 0.076623 | 0.08925 | 0.10432 | 0.1206 | 0.13645 | 0.14992 | 0.15793 | 0.15821 | 0.14965 | 0.13423 | 0.11423 | 0.093279 | 0.073862 | 0.057521 | 0.044451 | 0.034421 | 0.026766 | 0.02089 | 0.016448 | 0.013021 | 0.010225 | 0.0084856 | 0.0069896 | 0.0058967 | 0.0050832 | 0.0045142 | 0.0041497 | 0.0039574 | 0.0038715 | 0.0038344
Fit results CG=CPE-Cs=DE | 0.13513 | 0.14157 | 0.14388 | 0.14177 | 0.1356 | 0.12628 | 0.11498 | 0.10288 | 0.090935 | 0.079866 | 0.07012 | 0.061949 | 0.055477 | 0.050757 | 0.047819 | 0.046704 | 0.047484 | 0.05027 | 0.055219 | 0.06251 | 0.072302 | 0.084639 | 0.099291 | 0.11552 | 0.13183 | 0.14591 | 0.15502 | 0.15693 | 0.151 | 0.13852 | 0.12198 | 0.10405 | 0.086694 | 0.071061 | 0.057606 | 0.04636 | 0.037134 | 0.029653 | 0.023633 | 0.018813 | 0.014964 | 0.011896 | 0.0094544 | 0.0075125 | 0.0059687 | 0.0047417 | 0.0037668 | 0.0029922 | 0.0023769 | 0.0018881 | 0.0014998
simulation results from fitting CG=CPE-Cs=DE and remove CG=CPE | 0.0027685 | 0.0034854 | 0.0043877 | 0.0055237 | 0.0069536 | 0.0087536 | 0.011019 | 0.01387 | 0.017457 | 0.021968 | 0.027639 | 0.03476 | 0.043691 | 0.054866 | 0.068799 | 0.086074 | 0.1073 | 0.13303 | 0.16352 | 0.19844 | 0.23632 | 0.27403 | 0.30665 | 0.32823 | 0.33395 | 0.3224 | 0.29636 | 0.26134 | 0.2231 | 0.18598 | 0.15249 | 0.12364 | 0.099512 | 0.079713 | 0.063657 | 0.050736 | 0.040387 | 0.032124 | 0.025539 | 0.020297 | 0.016128 | 0.012814 | 0.01018 | 0.0080868 | 0.0064239 | 0.0051029 | 0.0040534 | 0.0032198 | 0.0025576 | 0.0020316 | 0.0016138
simulation results from fitting CG=CPE-Cs=DE but reduce CG 10% | 0.13971 | 0.14743 | 0.15098 | 0.14989 | 0.14437 | 0.13526 | 0.12373 | 0.11104 | 0.098278 | 0.086253 | 0.075494 | 0.066308 | 0.058841 | 0.053152 | 0.049263 | 0.047194 | 0.046989 | 0.048731 | 0.052545 | 0.058586 | 0.067001 | 0.077859 | 0.09101 | 0.10588 | 0.12124 | 0.13507 | 0.14481 | 0.14818 | 0.14416 | 0.13358 | 0.11862 | 0.10182 | 0.085229 | 0.07008 | 0.056931 | 0.04588 | 0.036781 | 0.029389 | 0.023431 | 0.018656 | 0.014841 | 0.0118 | 0.0093785 | 0.0074524 | 0.0059211 | 0.004704 | 0.0037369 | 0.0029685 | 0.002358 | 0.0018731 | 0.0014879
Chart: Log Frequency
| 1 | 1.0999998579423673 | 1.1999999472615537 | 1.2999999314429707 | 1.3999999253938222 | 1.5000000466711041 | 1.600000032123142 | 1.6999999708609113 | 1.7999999693838207 | 1.8999999810146946 | 2 | 2.0999998579423673 | 2.1999999472615537 | 2.2999999314429704 | 2.3999999253938222 | 2.5000000466711039 | 2.600000032123142 | 2.6999999708609113 | 2.7999999693838209 | 2.8999999810146946 | 3 | 3.0999998579423673 | 3.1999999472615537 | 3.2999999314429704 | 3.3999999253938222 | 3.5000000466711039 | 3.600000032123142 | 3.6999999708609113 | 3.7999999693838209 | 3.8999999810146946 | 4 | 4.0999998579423673 | 4.1999999472615537 | 4.2999999314429704 | 4.3999999253938222 | 4.5000000466711043 | 4.600000032123142 | 4.6999999708609108 | 4.7999999693838209 | 4.8999999810146946 | 5 | 5.0999998579423673 | 5.1999999472615537 | 5.2999999314429704 | 5.3999999253938222 | 5.5000000466711043 | 5.600000032123142 | 5.6999999708609108 | 5.7999999693838209 | 5.8999999810146946 | 6
Data | 0.14608 | 0.14492 | 0.14322 | 0.13877 | 0.13274 | 0.12435 | 0.11389 | 0.10356 | 0.093681 | 0.083851 | 0.074906 | 0.06705 | 0.060224 | 0.054763 | 0.050679 | 0.048268 | 0.047644 | 0.049094 | 0.052239 | 0.057957 | 0.066043 | 0.076623 | 0.08925 | 0.10432 | 0.1206 | 0.13645 | 0.14992 | 0.15793 | 0.15821 | 0.14965 | 0.13423 | 0.11423 | 0.093279 | 0.073862 | 0.057521 | 0.044451 | 0.034421 | 0.026766 | 0.02089 | 0.016448 | 0.013021 | 0.010225 | 0.0084856 | 0.0069896 | 0.0058967 | 0.0050832 | 0.0045142 | 0.0041497 | 0.0039574 | 0.0038715 | 0.0038344
Fit results CG=CPE-Cs=DE | 0.13513 | 0.14157 | 0.14388 | 0.14177 | 0.1356 | 0.12628 | 0.11498 | 0.10288 | 0.090935 | 0.079866 | 0.07012 | 0.061949 | 0.055477 | 0.050757 | 0.047819 | 0.046704 | 0.047484 | 0.05027 | 0.055219 | 0.06251 | 0.072302 | 0.084639 | 0.099291 | 0.11552 | 0.13183 | 0.14591 | 0.15502 | 0.15693 | 0.151 | 0.13852 | 0.12198 | 0.10405 | 0.086694 | 0.071061 | 0.057606 | 0.04636 | 0.037134 | 0.029653 | 0.023633 | 0.018813 | 0.014964 | 0.011896 | 0.0094544 | 0.0075125 | 0.0059687 | 0.0047417 | 0.0037668 | 0.0029922 | 0.0023769 | 0.0018881 | 0.0014998
simulation results from fitting CG=CPE-Cs=DE and remove CG=CPE | 0.0027685 | 0.0034854 | 0.0043877 | 0.0055237 | 0.0069536 | 0.0087536 | 0.011019 | 0.01387 | 0.017457 | 0.021968 | 0.027639 | 0.03476 | 0.043691 | 0.054866 | 0.068799 | 0.086074 | 0.1073 | 0.13303 | 0.16352 | 0.19844 | 0.23632 | 0.27403 | 0.30665 | 0.32823 | 0.33395 | 0.3224 | 0.29636 | 0.26134 | 0.2231 | 0.18598 | 0.15249 | 0.12364 | 0.099512 | 0.079713 | 0.063657 | 0.050736 | 0.040387 | 0.032124 | 0.025539 | 0.020297 | 0.016128 | 0.012814 | 0.01018 | 0.0080868 | 0.0064239 | 0.0051029 | 0.0040534 | 0.0032198 | 0.0025576 | 0.0020316 | 0.0016138
simulation results from fitting CG=CPE-Cs=DE but reduce CG 10% | 0.13971 | 0.14743 | 0.15098 | 0.14989 | 0.14437 | 0.13526 | 0.12373 | 0.11104 | 0.098278 | 0.086253 | 0.075494 | 0.066308 | 0.058841 | 0.053152 | 0.049263 | 0.047194 | 0.046989 | 0.048731 | 0.052545 | 0.058586 | 0.067001 | 0.077859 | 0.09101 | 0.10588 | 0.12124 | 0.13507 | 0.14481 | 0.14818 | 0.14416 | 0.13358 | 0.11862 | 0.10182 | 0.085229 | 0.07008 | 0.056931 | 0.04588 | 0.036781 | 0.029389 | 0.023431 | 0.018656 | 0.014841 | 0.0118 | 0.0093785 | 0.0074524 | 0.0059211 | 0.004704 | 0.0037369 | 0.0029685 | 0.002358 | 0.0018731 | 0.0014879
Glass wall impedance can increase the peak frequency () but reduce the peak amplitude ()
Decrease in during primary drying is due to loss of contact of ice with the side wall
10% decrease
Notes: Could you have a look at book chapter version 50 that we reviewed on friday and recreate the fig (that was in Zview) which shows the fit result to frozen water (Cg=CPE-DE=Cs). Then remove Cg=CPE and simulate the model keeping the fit result estimates for DE and Cs
That will show how the glass wall impedance increases the peak frequency but reduces the peak height
Then go back to the full model and make a 10% reduction in Cg (from 6.561E-13 to 5.9049E-13 )to and stimulate the full model to see of the peak shifts to lower frequency
Hopefully that will then allow us to qualify that the decrease in Fpeak during primary drying is due to loss of contact of ice with the side wall
When the ice loses contact then the heat transfer mechanism will change and so will Kv

## Slide 71
Chart: Time / h
Chart: Time / h
Temp. constant = -33.1± 0.05˚C
= -29.8± 0.03˚C
~32 ˚C
I
II

## Slide 72: Discussion
Decrease in suggests that the temperature may be decreasing after the steady state period, contrary to accepted knowledge that the temperature starts to increase owing to a reduction in drying rate and hence the degree of self cooling
Decrease in is more likely to be due to a change in the ice-glass contact associated with a change in the shape of the ice cylinder.
Conclusion
The period for determining the drying rate should be decreased from 25% ice loss to 10% suggested by our experiment.

## Slide 73: Limitations
and parameters rely on intimate contact of ice cylinder with glass wall
(100 kHz) parameter does not dependent on contact and can be used for end point but relationship between (100 kHz) ice constant is non-linear
Cable length limited to 1m at present
C-TVIS not compatible with front loading system
Incompatible with TCs in same TVIS vial (use fibre optic sensors – INFAP)

## Slide 74: Future Work
Development mapping a drying characteristics
  heat transfer coefficients ()
  dry layer resistance ()
Instrument Development
  Commercial C-TVIS (2018)
  Non-contact TVIS (2018-19)
    Micro-well screening
    Vial clusters in batch FD
  TVIS - Shuttle (2019-20)
Non-invasive real time information for characterising the freeze drying

## Slide 75: Acknowledgements, Recent Projects & Collaborators
De Montfort University, School of Pharmacy
  Evgeny Polygalov: co-inventor of TVIS instrument
  Yowwares Jeeraruangrattana. PhD student
  Bhaskar Pandya. PhD student
  Irina Ermolina. Senior Lecturer
Biopharmaceutical Stability at
Room Temperature
Analytical Technologies for the
Stabilization of Biopharmaceuticals
Government Support for industry
LyoDEA
Lyophilization process analytics
By dielectric analysis

## Slide 76: Acknowledgements, Recent Projects & Collaborators
De Montfort University
  Evgeny Polygalov. Senior Research Fellow
  Irina Ermolina. Senior Lecturer
  Yowwares Jeerarunangrattana. PhD student
  Bhaskar Pandya. PhD student
GEA Process Engineering
  Trevor Page & Julian Taylor
  Daniela Buchmeyer & Thomas Beutler
BlueFrog : Chris Samwell Ben Irvin
NIBSC : Paul Matejtschuk
Sanofi : Tim McCoy

## Slide 77: Thank you

## Slide 78: Objective
TVIS temperature calibration of log of top electrode () and bottom electrode ()
I

## Slide 79: Objective
Calibration by accounting for the temperature dependency of
II

## Slide 80: Objective
Development method for temperature compensation of during primary drying
III

## Slide 81: Objective
Prediction of ice temperatures at (i) interface () and (ii) base () including qualification TVIS technique ()
IV

## Slide 82: Objective
Estimation the surrogate drying rate () and evaluation TVIS determination with gravimetric method (weight loss)
V

## Slide 83: Objective
Determination (i) the drying rate () and (ii) ice base temperature () during the steady state period for heat transfer coefficient ( calculation
VI

## Slide 84
Freeze drying shelf
0 mm
Filling water into TVIS vial up to bottom edge of top electrode (≡ 0 mm)
Freezing water at -20°C for 3 hours
Freeze drying shelf
0 mm
Freeze drying shelf
0 mm
Thawing ice at 5°C for 3 hours
Freeze drying shelf
~0.92 mm
An aliquot part (0.35 g) of water corresponding to ice height 1 mm is added into TVIS vial
Repeat this process until the ice height reaches top edge of top electrode (15 mm)
Freeze drying shelf
~13.8 mm
Notes: *An aliquot part for Ice is calculated based on ice density of 0.9194 g/mL at -20 oC

## Slide 86: Objective
Prediction ice temperatures at (i) sublimation interface () and (ii) vial’s base () including qualification TVIS technique ()
VII
Freeze drying shelf
Chart: Temperature (𝑇) / oC
| -33.119466634243032 | -32.321208025650492 | -30.594741732648018 | -29.835299081767161
TE_BE_predict | 19.46 | 14.73 | 4.5 | 0
Chart: Time / h
-30.6 ˚C
4.5 mm
-32.3 ˚C
14.73 mm
Chart: Time / h
Ice height 19.46 mm at 2.4 h
-32.3 ˚C
-30.6 ˚C
At 2.4 h
Ice height for = () = 14.73 mm
Ice height for = () = 4.50 mm
2 mm
5 mm
3 mm
9.46 mm
-33.1 ˚C
-29.8 ˚C
0 mm
19.46 mm
Ice Temperature
At interface (, 19.46 mm) = = -33.1 ˚C
At vial’s base (, 0 mm) = = -29.8 ˚C

## Slide 87: Objective
Prediction ice temperatures at (i) sublimation interface () and (ii) vial’s base () including qualification TVIS technique ()
VII
Freeze drying shelf
Chart: Temperature (𝑇) / oC
| -33.119466634243032 | -32.321208025650492 | -30.594741732648018 | -29.835299081767161
TE_BE_predict | 19.46 | 14.73 | 4.5 | 0
Chart: Time / h
4.5 mm
-32.3 ˚C
14.73 mm
Chart: Time / h
Ice height 19.46 mm at 2.4 h
-30.6 ˚C
At 2.4 h
Ice height for = () = 14.73 mm
Ice height for = () = 4.50 mm
2 mm
5 mm
3 mm
9.46 mm
-33.1 ˚C
-29.8 ˚C
0 mm
19.46 mm
-30.6 ˚C
-32.3 ˚C
Ice Temperature
At interface (, 19.46 mm) = = -33.1 ˚C
At vial’s base (, 0 mm) = = -29.8 ˚C

## Slide 88: Ice Cylinder in Contact with the Inside of the Glass Wall
10.5 mm. from bottom edge of electrode (
25% of ice has been removed (
1
2
2 mm. from bottom edge of electrode (
Notes: Edge vial: Reduction in the drying rate is due to the reduction in the vertical surface area that received heat from radiation wall
Core vial: Any change in surface area might impact heat received from the radiation of top surface
Radiation is negligible

## Slide 89: Ice Cylinder in Contact with the Inside of the Glass Wall
10.5 mm. from bottom edge of electrode (
25% of ice has been removed (
3
4
2 mm. from bottom edge of electrode (
Notes: Edge vial: Reduction in the drying rate is due to the reduction in the vertical surface area that received heat from radiation wall
Core vial: Any change in surface area might impact heat received from the radiation of top surface
Radiation is negligible

## Slide 90: Ice Cylinder in Contact with the Inside of the Glass Wall
I
II
III
IV

## Slide 91: Ice Cylinder in Contact with the Inside of the Glass Wall
I
II
A
B

## Slide 92
Notes: Edited in version 14